## FDA Data
event_type | date_of_event | product_problems | complaint_txt | follow_up | manufacturer_narrative | device_generic_name | device_manufacturer_name | device_model_number | device_lot_number | device_report_product_code | expiration_date_of_device | date_returned_to_manufacturer | device_availability | device_operator | device_name
Malfunction | 20200418 | ['Battery Problem', 'Power Problem'] | INFORMATION RECEIVED BY MEDTRONIC INDICATED THAT THE INSULIN PUMP ALARMED WITH ALARM GENERATED WHEN A POWER SYSTEM ERROR WAS DETECTED AND THIS ERROR DOES NOT PREVENT THE PUMP FROM RUNNING. NO HARM REQUIRING MEDICAL INTERVENTION WAS REPORTED. THE CUSTOMER ALSO REPORTED EXPERIENCING POWER ERROR FOR 3 DAYS. CUSTOMER WAS ABLE TO SUCCESSFULLY CLEAR POWER ERROR ALARM. CUSTOMER WAS ADVISED TO REVERT TO THE BACKUP PLAN. THE INSULIN PUMP WILL BE RETURNED FOR ANALYSIS. |  | (B)(4). CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. THE DEVICE WILL BE RETURNED FOR ANALYSIS AND FURTHER INFORMATION WILL FOLLOW ONCE THE ANALYSIS HAS BEEN COMPLETED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. DEVICE PASSED SLEEP CURRENT MEASUREMENT, ACTIVE CURRENT MEASUREMENT, SELF TEST AND DISPLACEMENT TEST. HOWEVER, DEVICE TRACE DOWNLOAD ANALYSIS CONFIRMED THE INSULIN PUMP ALARMED POWER ERROR 25, POWER LOSS ALARM, REPLACE BATTERY ALERT AND REPLACE BATTERY NOW ALARM. THE ADAPT TOOL CONFIRMED POWER ERROR 25, POWER LOSS ALARM, REPLACE BATTERY ALERT, REPLACE BATTERY NOW ALARM WAS TRIGGERED WHEN THE BACKUP BATTERY LOADED VOLTAGE (LOADED VLITH) WAS LESS THAN 3.5V FOR 4 CONSECUTIVE HOURS DUE TO CONNECTOR RESISTANCE ELECTRICAL BOARD. AFTER DISCONNECTING AND RECONNECTING THE INTERNAL BATTERY CONNECTOR ON ELECTRICAL BOARD, DEVICE WAS MONITORED AND FUNCTIONED PROPERLY. | ARTIFICIAL PANCREAS DEVICE SYSTEM, THRESHOLD SUSPEND | MEDTRONIC PUERTO RICO OPERATIONS CO. | MMT-1715K | HG1B44Y | OZO |  | 20200429 | Device was returned to manufacturer | LAY USER/PATIENT | Automated Insulin Dosing , Threshold Suspend
Malfunction | 20200421 | ['Excess Flow or Over-Infusion', 'Battery Problem', 'Adverse Event Without Identified Device or Use Problem', 'Appropriate Term/Code Not Available'] | INFORMATION RECEIVED BY MEDTRONIC INDICATED THAT THE INSULIN PUMP AND RESERVOIR HAD OVER DELIVERY. NO HARM REQUIRING MEDICAL INTERVENTION WAS REPORTED. TROUBLESHOOT WAS NOT PERFORMED. THE INSULIN PUMP AND RESERVOIR WILL NOT BE RETURNING FOR ANALYSIS. |  | CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. | PUMP, INFUSION, INSULIN, TO BE USED WITH INVASIVE GLUCOSE SENSOR | MEDTRONIC MINIMED | MMT-XXX |  | OYC |  |  | No | LAY USER/PATIENT | Pump, Infusion, Insulin, To Be Used With Invasive Glucose Sensor
Injury | 20200407 | ['Failure to Charge', 'Battery Problem'] | FOLLOWING THE BATTERY PERFORMANCE ALERT (BPA) ADVISORY, A BPA WAS RECEIVED BY THE CLINICIAN. THE DEVICE WAS EXPLANTED AND REPLACED. THE DEVICE WAS FOUND TO BE AT ELECTIVE REPLACEMENT INDICATOR (ERI) AND CAPACITOR EXCEEDED THE CHARGED TIME. THE PATIENT WAS RECOVERING. |  | THE RESULTS/METHOD AND CONCLUSION CODES ALONG WITH INVESTIGATION RESULTS WILL BE PROVIDED IN THE FINAL REPORT. THE DEVICE IS INCLUDED IN THE BATTERY PERFORMANCE ALERT ADVISORY ISSUED BY ABBOTT ON 28 AUGUST 2017. THE REPORTED EVENT OF CHARGE TIME OUT WAS CONFIRMED. THE CAUSE OF CHARGE TIME OUT WAS DUE TO PREMATURE BATTERY DEPLETION. PREMATURE BATTERY DEPLETION WAS CONFIRMED BY ANALYSIS. NO SOURCES OF HIGH CURRENT WERE NOTED. THE CAUSE OF THE PREMATURE BATTERY DEPLETION WAS CONSISTENT WITH LI CLUSTER FORMATION. FROM THESE ANALYSES, IN THE ABSENCE OF HIGH CURRENT DRAW, IT IS PROBABLE THAT THE PREMATURE BATTERY DEPLETION WAS CAUSED BY A LITHIUM CLUSTER INDUCED SHORT CIRCUIT. LI CLUSTERS ARE A KNOWN DEPLETION MECHANISM FOR THESE ADVISORY PRODUCTS THAT HAS BEEN INVESTIGATED AND ASSOCIATED WITH A FIELD ACTION IN OCTOBER 2016. ADDITIONAL INFORMATION: D10, H7, H9 | IMPLANTABLE CARDIOVERTER DEFIBRILLATOR | ST. JUDE MEDICAL, INC.(CRM-SUNNYVALE) | CD1357-40Q | 4443992 | LWS | 20160131 | 20200422 | Device was returned to manufacturer | HEALTH PROFESSIONAL | Implantable Cardioverter Defibrillator (Non-Crt)
Malfunction | 20200402 | ['Battery Problem'] | IT WAS REPORTED THAT THE PUMP BATTERY WAS DEPLETING QUICKLY. THERE WAS NO REPORTED IMPACT TO THE CUSTOMER¿S BLOOD GLUCOSE LEVEL. THE CUSTOMER DECLINED TO PERFORM TROUBLESHOOTING AND DECLINED FOLLOW UP FROM TANDEM TECHNICAL SUPPORT. |  | NO PRODUCT WAS RETURNED FOR EVALUATION. SHOULD NEW RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. | CONTINUOUS GLUCOSE MONITOR | TANDEM DIABETES CARE | 1000096 |  | OYC |  |  | No | LAY USER/PATIENT | Pump, Infusion, Insulin, To Be Used With Invasive Glucose Sensor
Malfunction | 20200414 | ['Battery Problem'] | IT WAS REPORTED THAT THIS DEVICE TRIPPED ERI ON (B)(6) 2020. THE DATA ANALYSIS PERFORMED BY THE MANUFACTURER INDICATED A BATTERY ISSUE. NO ADVERSE PATIENT EVENTS WERE REPORTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THIS FILE WILL BE UPDATED. |  | THE DEVICE WAS NOT RETURNED FOR ANALYSIS. THE ANALYSIS IS THEREFORE ONLY BASED ON THE RETURNED DEVICE DATA AS WELL AS THE INSPECTION OF THE QUALITY DOCUMENTS ASSOCIATED WITH THE MANUFACTURE OF THIS PARTICULAR DEVICE. THE MANUFACTURING PROCESS FOR THIS DEVICE WAS RE-INVESTIGATED. ALL PRODUCTION STEPS HAD BEEN PERFORMED ACCORDINGLY. THERE WAS NO SIGN OF ANY INCONSISTENCY DURING THE MANUFACTURING. IN A NEXT STEP, THE RETURNED DEVICE DATA WAS INSPECTED. DURING THE INSPECTION THE CLINICAL OBSERVATION COULD BE CONFIRMED. FURTHER ANALYSIS REVEALED AN UNEXPECTED BATTERY BEHAVIOR, WHICH LED TO THE CLINICAL OBSERVATION. BASED ON THE DATA AVAILABLE FOR ANALYSIS THE ROOT CAUSE FOR THAT BEHAVIOR WAS NOT DETERMINABLE AND CAN ONLY BE CLARIFIED BY AN ANALYSIS OF THE DEVICE ITSELF. HOWEVER, IT CANNOT BE EXCLUDED THAT THIS OCCURRENCE RESULTED FROM A DAMAGED COMPONENT, WHICH MAY COMPROMISE THE ABILITY OF THE DEVICE TO DELIVER THERAPY. THEREFORE AND IN ORDER TO EXCLUDE ANY POTENTIAL HARM FOR THE PATIENT WE WOULD LIKE TO RECOMMEND TO REPLACE THE DEVICE AND TO RETURN IT FOR ANALYSIS TO BIOTRONIK. OF COURSE, INDIVIDUAL CIRCUMSTANCES AND MEDICAL JUDGMENT DETERMINE DECISIONS ABOUT PATIENT CARE AND THE FREQUENCY OF FOLLOW-UP SESSIONS. THE ABOVE RECOMMENDATION IS BASED ON A PURELY TECHNICAL ASSESSMENT. SHOULD FURTHER RELEVANT INFORMATION OR THE DEVICE ITSELF BECOME AVAILABLE THIS INVESTIGATION WILL BE UPDATED. IT WAS REPORTED THAT THIS DEVICE WAS EXPLANTED ON (B)(6) 2020. PRIOR TO THE ANALYSIS OF THE DEVICE, THE QUALITY DOCUMENTS ACCOMPANYING THE MANUFACTURING PROCESS FOR THIS ICD WERE RE-INVESTIGATED. ALL PRODUCTION STEPS WERE PERFORMED ACCORDINGLY, AND IN PARTICULAR THE FINAL ACCEPTANCE TEST PROVED THE DEVICE FUNCTIONS TO BE AS SPECIFIED. THE DEVICE INTERROGATION REVEALED THE EOS BATTERY STATUS, DETECTED AFTER EXPLANTATION OF THE DEVICE ON (B)(6) 2020. THE DEVICE WAS IMPLANTED FOR APPROXIMATELY 73 MONTHS AND 31 CHARGING CYCLES WERE RECORDED IN THE DEVICES MEMORY. THE ICD WAS SUBJECTED TO AN ELECTRICAL ANALYSIS. THE CURRENT CONSUMPTION OF THE ICD WAS FOUND TO BE NORMAL AND AS EXPECTED. HOWEVER, AN INCONSISTENCY BETWEEN THE AMOUNT OF CHARGE TAKEN FROM THE BATTERY AND THE BATTERY VOLTAGE WAS NOTED. IN ORDER TO OBTAIN FURTHER INSIGHTS, THE ICD WAS OPENED AND THE INNER ASSEMBLY WAS INSPECTED. THE VISUAL INSPECTION SHOWED NO ANOMALIES. SUBSEQUENTLY THE CURRENT CONSUMPTION OF THE ELECTRONIC MODULE WAS VERIFIED BY DIRECT MEASUREMENT AND PROVED TO BE AS EXPECTED AND CONSISTENT WITH THE INTERROGATED ICD DATA. THERE WAS NO INDICATION OF A MALFUNCTION OF THE ELECTRONIC MODULE. THEREFORE, THE BATTERY WAS DISCONNECTED FROM THE ELECTRONIC MODULE AND SENT TO THE MANUFACTURER FOR FURTHER ANALYSIS. DURING ANALYSIS OF THE BATTERY LITHIUM PLATING WAS IDENTIFIED, WHICH LED TO AN ELEVATED INTERNAL SELF-DEPLETION AND, AS A RESULT, TO THE CLINICAL OBSERVATION. PLEASE NOTE, THIS ICD IS AFFECTED BY THE FIELD SAFETY CORRECTIVE ACTION, BIO-LQC, INITIATED IN MARCH 2021. IT WAS REPORTED THAT THIS DEVICE WAS EXPLANTED ON (B)(6) 2020. THE DEVICE WAS NOT RETURNED FOR ANALYSIS. THE ANALYSIS IS THEREFORE ONLY BASED ON THE RETURNED DEVICE DATA AS WELL AS THE INSPECTION OF THE QUALITY DOCUMENTS ASSOCIATED WITH THE MANUFACTURE OF THIS PARTICULAR DEVICE. THE MANUFACTURING PROCESS FOR THIS DEVICE WAS RE-INVESTIGATED. ALL PRODUCTION STEPS HAD BEEN PERFORMED ACCORDINGLY. THERE WAS NO SIGN OF ANY INCONSISTENCY DURING THE MANUFACTURING. IN A NEXT STEP, THE RETURNED DEVICE DATA WAS INSPECTED. DURING THE INSPECTION THE CLINICAL OBSERVATION COULD BE CONFIRMED. FURTHER ANALYSIS REVEALED AN UNEXPECTED BATTERY BEHAVIOR, WHICH LED TO THE CLINICAL OBSERVATION. BASED ON THE DATA AVAILABLE FOR ANALYSIS THE ROOT CAUSE FOR THAT BEHAVIOR WAS NOT DETERMINABLE AND CAN ONLY BE CLARIFIED BY AN ANALYSIS OF THE DEVICE ITSELF. HOWEVER, IT CANNOT BE EXCLUDED THAT THIS OCCURRENCE RESULTED FROM A DAMAGED COMPONENT, WHICH MAY COMPROMISE THE ABILITY OF THE DEVICE TO DELIVER THERAPY. THEREFORE AND IN ORDER TO EXCLUDE ANY POTENTIAL HARM FOR THE PATIENT WE WOULD LIKE TO RECOMMEND TO REPLACE THE DEVICE AND TO RETURN IT FOR ANALYSIS TO BIOTRONIK. OF COURSE, INDIVIDUAL CIRCUMSTANCES AND MEDICAL JUDGMENT DETERMINE DECISIONS ABOUT PATIENT CARE AND THE FREQUENCY OF FOLLOW-UP SESSIONS. THE ABOVE RECOMMENDATION IS BASED ON A PURELY TECHNICAL ASSESSMENT. SHOULD FURTHER RELEVANT INFORMATION OR THE DEVICE ITSELF BECOME AVAILABLE THIS INVESTIGATION WILL BE UPDATED. | ICD | BIOTRONIK SE & CO. KG | 383594 |  | LWS | 20150228 | 20210302 | Device was returned to manufacturer | HEALTH PROFESSIONAL | Implantable Cardioverter Defibrillator (Non-Crt)
Malfunction | 20200226 | ['No Display/Image', 'Battery Problem', 'Device Markings/Labelling Problem', 'Material Integrity Problem', 'Power Problem', 'Patient Device Interaction Problem'] | INFORMATION RECEIVED BY MEDTRONIC INDICATED THAT THE INSULIN PUMP HAD LOW BATTERY AND BLANK DISPLAY. THE CUSTOMER STATED WAS UNABLE TO GET THE INSULIN PUMP TURN ON ALSO CUSTOMER HAD THE BATTERY CAP CONTACT MISSING AND THE SN LABEL WORN OFF. NO HARM REQUIRING MEDICAL INTERVENTION WAS REPORTED. THE INSULIN PUMP WAS RETURNED FOR ANALYSIS. |  | UPDATED H9: 2032227-060322-002-C MEDTRONIC, INC. (MEDTRONIC) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY MEDTRONIC, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. MEDTRONIC HAS MADE REASONABLE EFFORTS TO OBTAIN MORE COMPLETE INFORMATION IN THE TIME ALLOTTED AND HAS PROVIDED AS MUCH INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE AN ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY "DEFECTS" OR HAS "MALFUNCTIONED". THESE WORDS ARE INCLUDED IN THE FDA 3500A FORM AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. MEDTRONIC OBJECTS TO THE USE OF THESE WORDS AND OTHERS LIKE IT BECAUSE OF THE LACK OF DEFINITION AND THE CONNOTATIONS IMPLIED BY THESE TERMS. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT. THIS REPORT IS PART OF A RETROSPECTIVE REVIEW AND REMEDIATION EFFORTS IN RESPONSE TO A WARNING LETTER. UPDATED H9: Z-0956-2020 MEDTRONIC, INC. (MEDTRONIC) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY MEDTRONIC, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. MEDTRONIC HAS MADE REASONABLE EFFORTS TO OBTAIN MORE COMPLETE INFORMATION IN THE TIME ALLOTTED AND HAS PROVIDED AS MUCH INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE AN ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY "DEFECTS" OR HAS "MALFUNCTIONED". THESE WORDS ARE INCLUDED IN THE FDA 3500A FORM AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. MEDTRONIC OBJECTS TO THE USE OF THESE WORDS AND OTHERS LIKE IT BECAUSE OF THE LACK OF DEFINITION AND THE CONNOTATIONS IMPLIED BY THESE TERMS. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT. UNIT PASSED SLEEP CURRENT MEASUREMENT, ACTIVE CURRENT MEASUREMENT, AND SELF TESTS. NO UNEXPECTED BLANK DISPLAYS OR UNEXPECTED ¿LOW BATTERY¿, ¿REPLACE BATTERY NOW¿, OR "POWER LOSS" ERRORS WERE NOTED DURING TESTING. UNIT RECEIVED WITH A PARTIALLY BROKEN OFF PIECE AT THE RESERVOIR TUBE LIP, A MISSING RETAINER RING, AND A MISSING RESERVOIR TUBE LIP O-RING. UNABLE TO PERFORM THE DISPLACEMENT TEST SINCE P-CAP / RESERVOIR WILL NOT LOCK PROPERLY DUE TO THE MISSING RETAINER. IN ADDITION TO THIS, THE S/N LABEL LOCATED BETWEEN THE BELT CLIP RAILS HAS RUBBED OFF AND BECOME ILLEGIBLE. (B)(4). | ARTIFICIAL PANCREAS DEVICE SYSTEM, SINGLE HORMONAL CONTROL | MEDTRONIC PUERTO RICO OPERATIONS CO. | MMT-1780KPK | HG29Y7N | OZP |  | 20200309 | Device was returned to manufacturer | LAY USER/PATIENT | Automated Insulin Dosing Device System, Single Hormonal Control
Injury | 20200422 | ['Battery Problem'] | FOLLOWING THE BATTERY PERFORMANCE ALERT (BPA) ADVISORY, A BPA WAS RECEIVED BY THE CLINICIAN AND THE DEVICE WAS EXPLANTED. THE PATIENT WAS IN STABLE CONDITION. |  | THE DEVICE IS INCLUDED IN THE BATTERY PERFORMANCE ALERT ADVISORY ISSUED BY ABBOTT ON (B)(6) 2017. THE RESULTS, METHOD, AND CONCLUSION CODES ALONG WITH INVESTIGATION RESULTS WILL BE PROVIDED IN THE FINAL REPORT. PREMATURE BATTERY DEPLETION WAS CONFIRMED BY ANALYSIS. NO SOURCES OF HIGH CURRENT WERE NOTED. THE CAUSE OF THE PREMATURE BATTERY DEPLETION WAS CONSISTENT WITH LITHIUM (LI) CLUSTER FORMATION. FROM THESE ANALYSES, IN THE ABSENCE OF HIGH CURRENT DRAW, IT IS PROBABLE THAT THE PREMATURE BATTERY DEPLETION WAS CAUSED BY A LITHIUM CLUSTER INDUCED SHORT CIRCUIT. LI CLUSTERS ARE A KNOWN DEPLETION MECHANISM FOR THESE ADVISORY PRODUCTS THAT HAS BEEN INVESTIGATED AND ASSOCIATED WITH A FIELD ACTION IN OCTOBER 2016. | IMPLANTABLE CARDIOVERTER DEFIBRILLATOR | ABBOTT | CD1233-40Q | 3122369 | LWS | 20111130 | 20200505 | Device was returned to manufacturer | HEALTH PROFESSIONAL | Implantable Cardioverter Defibrillator (Non-Crt)
Injury | 20200410 | ['Battery Problem', 'Insufficient Information'] | ADDITIONAL INFORMATION INDICATES THE IPG WAS EXPLANTED AND REPLACED TO ADDRESS THE ISSUE. | IT WAS REPORTED THE IPG HAS REACHED END OF SERVICE. IN TURN, SURGICAL INTERVENTION MAY BE PENDING. | THE RESULTS OF THE INVESTIGATION ARE INCONCLUSIVE SINCE THE DEVICE WAS NOT RETURNED FOR ANALYSIS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED. THE RESULTS/METHOD AND CONCLUSION CODES ALONG WITH INVESTIGATION RESULTS WILL BE PROVIDED IN THE FINAL REPORT. DATE OF EVENT IS ESTIMATED. DURING PROCESSING OF THIS COMPLAINT, ATTEMPTS WERE MADE TO OBTAIN COMPLETE PATIENT INFORMATION. FURTHER INFORMATION WAS REQUESTED BUT NOT RECEIVED. THE RESULTS OF THE INVESTIGATION ARE INCONCLUSIVE SINCE THE DEVICE WAS NOT RETURNED FOR ANALYSIS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED. | SCS IPG | ST. JUDE MEDICAL - NEUROMODULATION | 3662 | 5758833 | LGW | 20181213 |  | No | HEALTH PROFESSIONAL | Stimulator, Spinal-Cord, Totally Implanted For Pain Relief
Injury | 20200331 | ['Battery Problem'] | INFORMATION WAS RECEIVED FROM A CONSUMER (CON) REGARDING A PATIENT WHO WAS IMPLANTED WITH A NEUROSTIMULATOR (INS) FOR GASTRIC STIMULATION. IT WAS REPORTED THAT THE PATIENT THOUGHT THEIR INS WAS DEAD BECAUSE THEY HAD A LOT OF PAIN IN THEIR LOWER BACK. THEY STATED THEY WERE IN THE HOSPITAL WITH PAIN. NO FURTHER COMPLICATIONS WERE REPORTED OR ANTICIPATED. | ADDITIONAL INFORMATION WAS RECEIVED FROM THE PATIENT. THEY PROVIDED THEIR WEIGHT AT THE TIME OF THE EVENT. THEY STATED THERE WERE NO CIRCUMSTANCES LEADING TO THE DEAD INS AND PAIN. THEY WERE DOING NOTHING WHEN THE PAIN STARTED. IT WAS NOT RESOLVED AT THE TIME OF THE REPORT. NO PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. | IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. | INTESTINAL STIMULATOR | MEDTRONIC PUERTO RICO OPERATIONS CO. | 37800 |  | LNQ | 20181028 |  | No | HEALTH PROFESSIONAL | Intestinal Stimulator
Malfunction | 20200427 | ['Circuit Failure', 'Battery Problem', 'Material Integrity Problem'] | INFORMATION RECEIVED BY MEDTRONIC INDICATED THAT INSULIN PUMP HAD MULTIPLE PUMP ERROR ALARMS. CUSTOMER WAS ABLE TO SUCCESSFULLY CLEAR THE ALARM AND ABLE TO REWIND THE INSULIN PUMP. CUSTOMER REPORTED THAT BATTERY CAP METAL CONNECTOR WAS ALSO LOOSE. CUSTOMER ALSO REPORTED THAT TROUBLESHOOT WAS PERFORMED FOR PUMP ERROR ALARM. CUSTOMER WAS PERFORMING SELF TEST AND IT DID NOT PASS. NO HARM REQUIRING MEDICAL INTERVENTION WAS REPORTED. THE INSULIN PUMP WILL NOT BE RETURNED FOR ANALYSIS. |  | (B)(4). CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. THE INSULIN PUMP INVOLVED IN THIS EVENT IS THE 640G INSULIN INFUSION PUMP, WHICH IS NOT MARKETED IN THE UNITED STATES. HOWEVER, THE DEVICE IS SIMILAR TO THE PARADIGM REAL-TIME INSULIN INFUSION PUMP, WHICH IS MARKETED IN THE UNITED STATES. UPDATED H9: 2032227-060322-002-C MEDTRONIC, INC. (MEDTRONIC) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY MEDTRONIC, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. MEDTRONIC HAS MADE REASONABLE EFFORTS TO OBTAIN MORE COMPLETE INFORMATION IN THE TIME ALLOTTED AND HAS PROVIDED AS MUCH INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE AN ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY "DEFECTS" OR HAS "MALFUNCTIONED". THESE WORDS ARE INCLUDED IN THE FDA 3500A FORM AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. MEDTRONIC OBJECTS TO THE USE OF THESE WORDS AND OTHERS LIKE IT BECAUSE OF THE LACK OF DEFINITION AND THE CONNOTATIONS IMPLIED BY THESE TERMS. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT. DEVICE PASSED THE DISPLACEMENT TEST. HOWEVER, PUMP ERROR 63 ALARM (VARIABLE INFO=3) DURING SELF TEST AND CONFIRMED IN THE DEVICE HISTORY DUE COLD SOLDER JOINT ON U1 KEYPAD ASSEMBLY. NO ERROR 25 ALARM NOTED IN THE HISTORY FILE. UNABLE TO CONFIRM BATTERY CAP DAMAGE DUE TO INSULIN PUMP RECEIVED WITHOUT THE ORIGINAL BATTERY CAP. | PUMP, INFUSION, INSULIN, TO BE USED WITH INVASIVE GLUCOSE SENSOR | MEDTRONIC PUERTO RICO OPERATIONS CO. | MMT-1711K | HG1WT7E | OYC |  | 20200706 | Device was returned to manufacturer | LAY USER/PATIENT | Pump, Infusion, Insulin, To Be Used With Invasive Glucose Sensor