## Sheet1
Web Address | Report Number | Event Date | Event Type | Manufacturer | Date Received | Product Code | Brand Name | Device Problem | Patient Problem | PMA/PMN Number | Exemption Number | Number of Events | Event Text
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20808557&pc=LWP | 2124215-2024-75434 | 2024-11-05 05:00 | Injury | BOSTON SCIENTIFIC CORPORATION | 2024-11-29 05:00 | LWP | INCEPTA ICD | Premature Discharge of Battery; Battery Problem | No Clinical Signs, Symptoms or Conditions | P960040 |  | 1 | Event Description: IT WAS REPORTED THAT THIS IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) EXHIBITED PREMATURE BATTERY DEPLETION (PBD). THIS DEVICE WAS EXPLANTED AND REPLACED. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20810785&pc=LWP | 2124215-2024-75483 | 2023-09-08 04:00 | Injury | BOSTON SCIENTIFIC CORPORATION | 2024-11-29 05:00 | LWP | ACCOLADE MRI EL DR | Premature Discharge of Battery; Battery Problem | No Clinical Signs, Symptoms or Conditions | P150012/S000 |  | 1 | Event Description: IT WAS REPORTED THAT THE BATTERY OF THIS PACEMAKER WAS SUSPECTED TO BE DEPLETING PREMATURELY AS THE ESTIMATED LONGEVITY REMAINING DECREASED FROM 10.5 YEARS TO 8.5 YEARS DURING THE SAME MONTH. SLIGHT INCREASE IN ATRIAL PACING PERCENTAGE MAY ACCOUNT FOR SLIGHT INCREASE IN POWER CONSUMPTION. THE PATIENT WILL BE MONITORED. THE DEVICE WAS EXPLANTED AND SUCCESSFULLY REPLACED. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED.Event Description: IT WAS REPORTED THAT THE BATTERY OF THIS PACEMAKER WAS SUSPECTED TO BE DEPLETING PREMATURELY AS THE ESTIMATED LONGEVITY REMAINING DECREASED FROM 10.5 YEARS TO 8.5 YEARS DURING THE SAME MONTH. SLIGHT INCREASE IN ATRIAL PACING PERCENTAGE MAY ACCOUNT FOR SLIGHT INCREASE IN POWER CONSUMPTION BUT IT WAS NOT CONFIRMED. A SECOND ANALYSIS WAS PERFORMED APPROXIMATELY ONE YEAR LATER, WHICH CONFIRMED A GRADUAL INCREASE IN POWER CONSUMPTION. BASED ON THESE FINDINGS, THE EXPLANTATION OF THE DEVICE WAS RECOMMENDED. THE DEVICE WAS SUBSEQUENTLY EXPLANTED AND SUCCESSFULLY REPLACED. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. Manufacturer Narrative: CORRECTED FIELDS: B5 DESCRIBE EVENT OR PROBLEM. G3 BSC AWARE DATE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20810855&pc=LWP | 2124215-2024-75490 | 2024-10-25 04:00 | Injury | BOSTON SCIENTIFIC CORPORATION | 2024-11-29 05:00 | LWP | ENERGEN CRT-D | Premature Discharge of Battery; Delayed Charge Time; Battery Problem | Loss of consciousness; Syncope/Fainting | P010012 |  | 1 | Event Description: IT WAS REPORTED THAT THIS CARDIAC RESYNCHRONIZATION THERAPY DEFIBRILLATOR (CRT-D) WAS EXPLANTED DUE TO CONCERNS OF PREMATURE BATTERY DEPLETION (PBD) AS IT DECREASED FROM 1 YEAR REMAINING TO END OF LIFE (EOL) IN A 7 MONTHS TIME PERIOD. THE PATIENT EXPERIENCED TWO SYNCOPAL EPISODES DURING THIS TIME. A DEVICE CHANGEOUT PROCEDURE WAS PERFORMED AND IT WAS EXPLANTED AND REPLACED. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. IT IS EXPECTED TO RECEIVE THIS PRODUCT FOR ANALYSIS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20805256&pc=LWP | 2124215-2024-75234 | 2024-11-08 05:00 | Malfunction | BOSTON SCIENTIFIC CORPORATION | 2024-11-28 05:00 | LWP | PUNCTUA ICD | High impedance; Off-Label Use; Battery Problem | No Clinical Signs, Symptoms or Conditions | P960040 |  | 1 | Event Description: IT WAS REPORTED THAT ASSISTANCE WAS REQUESTED TO PROGRAM THIS IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD), WHICH WAS AT ELECTIVE REPLACEMENT INDICATOR (ERI), INTO MAGNETIC RESONANCE IMAGING (MRI) MODE. TECHNICAL SERVICES (TS) NOTED THAT THIS SYSTEM IS NON-MRI CONDITIONAL AND OBSERVED HIGH OUT-OF-RANGE SHOCK IMPEDANCES EXCEEDING 125 OHMS. DESPITE THIS, THE DEVICE WAS PROGRAMMED INTO MRI MODE AND ELECTROCAUTERY MODE AS DIRECTED BY CARDIOLOGY, AND THE MRI WAS SUCCESSFULLY PERFORMED. FOLLOWING THE MRI PROCEDURE, THE DEVICE WAS EXITED FROM MRI MODE, CHECKED, AND CONFIRMED TO BE FUNCTIONING AS PROGRAMMED. NO ADVERSE PATIENT EFFECTS WERE REPORTED. THIS ICD REMAINS IN SERVICE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20807184&pc=LWP | 2124215-2024-75371 | 2024-11-22 05:00 | Injury | BOSTON SCIENTIFIC CORPORATION | 2024-11-28 05:00 | LWP | ACCOLADE MRI DR | Signal Artifact/Noise; Low impedance; Battery Problem; Material Integrity Problem; High Capture Threshold | No Clinical Signs, Symptoms or Conditions | P150012 |  | 1 | Event Description: IT WAS REPORTED THAT THIS RECENTLY IMPLANTED PACEMAKER DROPPED ITS BATTERY LONGEVITY BY UP TO TWO YEARS IN A SPAN OF TWO MONTHS. AND THERE WERE CONCERNS OF NOISE ON THE RIGHT ATRIAL (RA) CHANNEL. TECHNICAL SERVICES (TS) WAS CONSULTED, AND IT WAS FOUND THAT THERE WERE LOW OUT OF RANGE PACING IMPEDANCE MEASUREMENTS ON THE RA CHANNEL AND HIGH PACING THRESHOLDS. IN ADDITION, THERE WAS A HIGH-POWER CONSUMPTION, THIS BEHAVIOR WAS CONSISTENT WITH A GATE OXIDE DEFECT OF INTERNAL CIRCUIT OF THE PACEMAKER WITH THE POTENTIAL FOR HEADER- LEAD CORROSION. DEVICE AND RA LEAD REPLACEMENT WAS RECOMMENDED. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. THIS PACEMAKER WAS EXPLANTED AND SUCCESSFULLY REPLACED. AT THIS MOMENT, THIS PACEMAKER HAS NOT RETURNED TO BOSTON SCIENTIFIC FOR FURTHER ANALYSIS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20793050&pc=LWP | 2124215-2024-74840 | 2024-11-04 05:00 | Malfunction | BOSTON SCIENTIFIC CORPORATION | 2024-11-27 05:00 | LWP | ACCOLADE MRI SR | Premature Discharge of Battery; Battery Problem | No Clinical Signs, Symptoms or Conditions | P150012/S000 |  | 1 | Event Description: IT WAS REPORTED THAT THIS IMPLANTABLE PACEMAKER WAS SUSPECTED TO HAVE PREMATURE BATTERY DEPLETION (PBD) DUE TO DIFFERING LONGEVITY ESTIMATES GIVEN SINCE AUGUST OF LAST YEAR. THE BATTERY LONGEVITY WAS 5.5 YEARS LAST AUGUST AND YESTERDAY WAS 3.5 YEARS. TECHNICAL SERVICES (TS) WAS CONSULTED AND ADVISED THIS COULD BE A TRANSITION PERIOD FROM BLENDED CAPACITY TO MONITORING VOLTAGE ALONE. IT WAS NOTED THE PATIENT WILL BE REEVALUATED AGAIN IN SIX MONTHS AND WILL VERIFY STABLE LONGEVITY DROPS. TS SUGGESTED TO SAVE DATA ANALYSIS, AND HAVE IT EVALUATED DUE TO CURRENT DATA WAS NOT SAVED AND THE PATIENT IS NOT ON EXTERNAL MONITORING. THE DEVICE REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20802806&pc=LWP | 2124215-2024-75174 | 2024-11-08 05:00 | Malfunction | BOSTON SCIENTIFIC CORPORATION | 2024-11-27 05:00 | LWP | ACCOLADE MRI EL DR | Premature Discharge of Battery; Battery Problem | No Clinical Signs, Symptoms or Conditions | P150012/S000 |  | 1 | Event Description: IT WAS REPORTED THAT THIS IMPLANTABLE PACEMAKER EXHIBITED PREMATURE BATTERY DEPLETION. THE POWER CONSUMPTION HAS INCREASED SHOWING A BATTERY LONGEVITY FROM FOUR YEARS TO TWO YEARS. A DEVICE REPLACEMENT PROCEDURE WAS SCHEDULED FOR THE FUTURE. AT THIS TIME, THIS DEVICE REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20784419&pc=LWP | 2124215-2024-74568 | 2024-10-31 04:00 | Malfunction | BOSTON SCIENTIFIC CORPORATION | 2024-11-26 05:00 | LWP | INGENIO | Battery Problem | No Clinical Signs, Symptoms or Conditions | N970003 |  | 1 | Event Description: IT WAS REPORTED THAT THIS IMPLANTABLE PACEMAKER EXHIBITED LONGEVITY CALCULATIONS FROM 2 YEARS REMAINING TO LESS THAN 3 MONTHS REMAINING IN A SPAN OF 1 YEAR. BOSTON SCIENTIFIC TECHNICAL SERVICES (TS) DISCUSSED THE CHANGE IN LONGEVITY IS DUE TO THE BATTERY VOLTAGE BEING HIGHER THAN THE ESTIMATED VOLTAGE AT THE TIME. DEVICE REPLACEMENT RECOMMENDED. THIS DEVICE REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20772431&pc=LWP | 2124215-2024-74160 | 2024-11-06 05:00 | Malfunction | BOSTON SCIENTIFIC CORPORATION | 2024-11-25 05:00 | LWP | ACCOLADE MRI DR | Premature Discharge of Battery; Battery Problem | No Clinical Signs, Symptoms or Conditions | P150012 |  | 1 | Event Description: IT WAS REPORTED THAT THE BATTERY OF THIS PACEMAKER WAS SUSPECTED TO BE DEPLETING PREMATURELY. A REQUEST WAS MADE TO HAVE DATA FROM THIS PACEMAKER ANALYZED. THE PACEMAKER IS STILL IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED. Manufacturer Narrative: IF PERTINENT INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20759451&pc=LWP | 2124215-2024-73876 | 2024-11-19 05:00 | Malfunction | BOSTON SCIENTIFIC CORPORATION | 2024-11-22 05:00 | LWP | INGENIO | Premature Discharge of Battery; Battery Problem | No Clinical Signs, Symptoms or Conditions |  |  | 1 | Event Description: IT WAS REPORTED THAT THE BATTERY OF THIS IMPLANTABLE DEVICE WAS SUSPECTED TO BE DEPLETING PREMATURELY DUE TO A DECREASE IN REMAINING LONGEVITY FROM 2.5 YEAR TO 1 YEAR OVER THE PAST 7 WEEKS. A REQUEST WAS MADE TO HAVE DATA FROM THIS DEVICE ANALYZED. DATA ANALYSIS CONFIRMED THE POWER CONSUMPTION IS NORMAL, BUT THE BATTERY PERFORMANCE IS LOWER THAN THE EXPECTED NOMINALS AND AT THE CURRENT TREND, ELECTIVE REPLACEMENT INDICATOR (ERI) WILL TRIGGER IN LESS THAN ONE YEAR. DEVICE REPLACEMENT OR RE-EVALUATION OF THE BATTERY STATUS IN 90 DAYS WAS RECOMMENDED. THIS DEVICE REMAINS IN SERVICE AT THIS TIME AND A REPLACEMENT INTERVAL WAS COMMUNICATED TO THE CALLER. NO ADVERSE PATIENT EFFECTS WERE REPORTED. Manufacturer Narrative: AS NO FURTHER INFORMATION CONCERNING THIS REPORT IS EXPECTED, OUR INVESTIGATION IS COMPLETE. THIS INVESTIGATION WILL BE UPDATED SHOULD FURTHER INFORMATION BE PROVIDED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20760821&pc=LWP | 2124215-2024-73914 | 2024-11-12 05:00 | Malfunction | BOSTON SCIENTIFIC CORPORATION | 2024-11-22 05:00 | LWP | INGENIO | Premature Discharge of Battery; Battery Problem | No Clinical Signs, Symptoms or Conditions | N970003 |  | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS FOUND TO BE DEPLETING FASTER THAN EXPECTED. THE PATIENT HAD BEEN SEEN ABOUT A YEAR AGO AND THE LONGEVITY REMAINING WAS ABOUT TWO YEARS. AT THE MOST RECENT FOLLOW-UP, HOWEVER, THE LONGEVITY REMAINING DECREASED TO ABOUT SIX MONTHS. A REQUEST WAS MADE TO HAVE DATA FROM THIS DEVICE ANALYZED. DATA ANALYSIS CONFIRMED THE BATTERY APPEARED TO BE DEPLETING MORE QUICKLY THAN EXPECTED. DEVICE REPLACEMENT WAS RECOMMENDED. NO ADVERSE PATIENT EFFECTS WERE REPORTED. CURRENTLY EVIDENCE SUGGESTS THAT THIS PRODUCT REMAINS IN-SERVICE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20748459&pc=LWP | 2124215-2024-73535 | 2023-07-16 04:00 | Malfunction | BOSTON SCIENTIFIC CORPORATION | 2024-11-21 05:00 | LWP | ACCOLADE MRI DR | Premature Discharge of Battery; Battery Problem | No Clinical Signs, Symptoms or Conditions | P150012 |  | 1 | Event Description: IT WAS REPORTED THAT THE BATTERY OF THIS PACEMAKER WAS SUSPECTED TO BE DEPLETING PREMATURELY. A REQUEST WAS MADE TO HAVE DATA FROM THIS DEVICE ANALYZED. DATA ANALYSIS CONFIRMED THE BATTERY APPEARED TO BE DEPLETING MORE QUICKLY THAN EXPECTED. DEVICE REPLACEMENT WAS RECOMMENDED. THIS DEVICE CURRENTLY REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20754304&pc=LWP | 2124215-2024-73697 | 2024-11-18 05:00 | Malfunction | BOSTON SCIENTIFIC CORPORATION | 2024-11-21 05:00 | LWP | PROPONENT MRI EL DR | Premature Discharge of Battery; Battery Problem | No Clinical Signs, Symptoms or Conditions | P150012/S000 |  | 1 | Event Description: IT WAS REPORTED THAT THIS IMPLANTABLE PACEMAKER'S BATTERY LONGEVITY WAS NOTED TO HAVE DECREASED FROM 6 YEARS REMAINING IN MAY 2024 TO 4.5 YEARS REMAINING AT A RECENT FOLLOW-UP APPOINTMENT. ADDITIONALLY, THE POWER CONSUMPTION WAS FOUND TO BE ELEVATED AT 164%. BOSTON SCIENTIFIC TECHNICAL SERVICES WAS CONTACTED AND CONFIRMED THAT THE BATTERY WAS DEPLETING PREMATURELY. IT WAS RECOMMENDED TO PERFORM DEVICE REPLACEMENT OR REPEAT DATA ANALYSIS WITHIN A PROVIDED TIMEFRAME. AT THIS TIME, THIS PACEMAKER REMAINS IMPLANTED AND IN-SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20737009&pc=LWP | 2124215-2024-73143 | 2024-10-21 04:00 | Injury | BOSTON SCIENTIFIC CORPORATION | 2024-11-20 05:00 | LWP | ACCOLADE MRI SR | Premature Discharge of Battery; High impedance; Battery Problem; High Capture Threshold | No Clinical Signs, Symptoms or Conditions | P150012/S000 |  | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS SUSPECTED OF PREMATURE BATTERY DEPLETION (PBD) AS THERE WAS A DECREASE IN ESTIMATED BATTERY OF 4 YEARS OVER A TIME PERIOD OF 6 MONTHS. ENGINEERING ANALYSIS OF DEVICE DATA FOUND THAT BATTERY DEPLETION APPEARED TO BE NORMAL AND THE INCREASED POWER CONSUMPTION WAS DUE TO HIGH LEVELS OF RIGHT VENTRICULAR (RV) LEAD IMPEDANCES RANGING FROM 260 TO 800 OHMS AND CAPTURE THRESHOLDS. THE PHYSICIAN ELECTED TO EXPLANT THE DEVICE AND THE RV LEAD AND REPLACE THEM WITH NEW ONES. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20738476&pc=LWP | 2124215-2024-73221 | 2024-11-03 04:00 | Malfunction | BOSTON SCIENTIFIC CORPORATION | 2024-11-20 05:00 | LWP | ACCOLADE MRI EL DR | Premature Discharge of Battery; Battery Problem | No Clinical Signs, Symptoms or Conditions | P150012/S000 |  | 1 | Event Description: IT WAS REPORTED THAT THE BATTERY OF THIS IMPLANTABLE DEVICE WAS SUSPECTED TO BE DEPLETING PREMATURELY AS REMAINING LONGEVITY DECREASED FROM 5.5 YEARS TO 3 YEARS OVER THE PAST YEAR. A REQUEST WAS MADE TO HAVE DATA FROM THIS DEVICE ANALYZED. DATA ANALYSIS CONFIRMED THE BATTERY APPEARED TO BE DEPLETING MORE QUICKLY THAN EXPECTED. DEVICE REPLACEMENT WAS RECOMMENDED. THIS DEVICE REMAINS IN SERVICE AT THIS TIME AND A REPLACEMENT INTERVAL WAS COMMUNICATED TO THE CALLER. NO ADVERSE PATIENT EFFECTS WERE REPORTED. Manufacturer Narrative: AS NO FURTHER INFORMATION CONCERNING THIS REPORT IS EXPECTED, OUR INVESTIGATION IS COMPLETE. THIS INVESTIGATION WILL BE UPDATED SHOULD FURTHER INFORMATION BE PROVIDED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20741621&pc=LWP | 2124215-2024-73338 | 2024-11-14 05:00 | Malfunction | BOSTON SCIENTIFIC CORPORATION | 2024-11-20 05:00 | LWP | ENERGEN ICD | Premature Discharge of Battery; Battery Problem | No Clinical Signs, Symptoms or Conditions |  |  | 1 | Event Description: IT WAS REPORTED THAT THIS IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) DEVICE ESTIMATED BATTERY LONGEVITY DROPPED FROM 5 YEARS REMAINING A YEAR AGO, TO 1 YEAR REMAINING TODAY. SUSPECTED PREMATURE BATTERY DEPLETION (PBD). A REQUEST WAS MADE TO HAVE DATA FROM THIS DEVICE ANALYZED. A TECHNICAL SERVICE (TS) CONSULTANT ADVISED TO ADDITIONAL INFORMATION IS NEEDED FROM THE DEVICE, FOR THEM TO REVIEW. AT THIS TIME THE DEVICE REMAINS IMPLANTED. NO ADVERSE PATIENT EFFECTS WERE REPORTED. Manufacturer Narrative: IF PERTINENT INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20722761&pc=LWP | 2124215-2024-72610 | 2024-10-21 04:00 | Malfunction | BOSTON SCIENTIFIC CORPORATION | 2024-11-19 05:00 | LWP | ESSENTIO MRI DR | Battery Problem | No Clinical Signs, Symptoms or Conditions | P150012 |  | 1 | Event Description: IT WAS REPORTED THAT THE IMPLANTABLE PACEMAKER WAS GETTING LOW ON BATTERY. BOSTON SCIENTIFIC TECHNICAL SERVICES (TS) REFERRED PATIENT TO GET DEVICE CHECKED BY THE PHYSICIAN. THIS DEVICE REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20723344&pc=LWP | 2124215-2024-72637 | 2024-10-24 04:00 | Injury | BOSTON SCIENTIFIC CORPORATION | 2024-11-19 05:00 | LWP | PROPONENT MRI DR | Premature Discharge of Battery; High impedance; Battery Problem | No Clinical Signs, Symptoms or Conditions | P150012/S000 |  | 1 | Event Description: IT WAS REPORTED THAT AFTER ONE MONTH OF IMPLANT TIME, THIS PACEMAKER EXHIBITED A REMAINING LONGEVITY OF 2.5 YEARS. THE POWER CONSUMPTION WAS HIGHER THAN EXPECTED AT 270% AND PREMATURE BATTERY DEPLETION WAS SUSPECTED. IT WAS ALSO NOTED THAT THE RIGHT ATRIAL (RA) PACING IMPEDANCE MEASUREMENTS WERE INTERMITTENTLY HIGH, OUT-OF-RANGE AT GREATER THAN 3,000 OHMS AND LOSS OF CAPTURE WAS ALSO OBSERVED INTERMITTENTLY. AT THIS TIME, THE DEVICE WAS REPROGRAMMED TO VVI MODE. A REQUEST WAS MADE FOR TECHNICAL SERVICES TO REVIEW THE DEVICE DATA AND DETERMINE IF THE DEVICE NEEDED TO BE REPLACED. NO ADVERSE PATIENT EFFECTS WERE REPORTED. THE DEVICE REMAINS IMPLANTED AND IN SERVICE. ADDITIONAL INFORMATION WAS RECEIVED. THE DEVICE WAS EXPLANTED AND REPLACED. THE NON-BOSTON SCIENTIFIC RA AND RIGHT VENTRICULAR (RV) LEADS WERE REMOVED AND NEW LEADS WERE IMPLANTED ON THE PATIENT'S RIGHT SIDE, DUE TO LEFT VENOUS STENOSIS. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. THE EXPLANTED PACEMAKER WAS NOT EXPECTED TO BE RETURNED. IT WAS LATER CLARIFIED THAT THE ORIGINAL DEVICE AND NON-BOSTON SCIENTIFIC LEADS WERE ABANDONED IN THE PATIENT AND WERE NOT REMOVED WHEN THE NEW SYSTEM WAS IMPLANTED ON THE PATIENT'S RIGHT SIDE. Manufacturer Narrative: IF PERTINENT INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20728380&pc=LWP | 2124215-2024-72887 | 2024-10-21 04:00 | Malfunction | BOSTON SCIENTIFIC CORPORATION | 2024-11-19 05:00 | LWP | ACCOLADE MRI DR | Premature Discharge of Battery; Battery Problem | No Clinical Signs, Symptoms or Conditions | P150012/S000 |  | 1 | Event Description: IT WAS REPORTED THAT THE BATTERY OF THIS PACEMAKER WAS SUSPECTED TO BE DEPLETING PREMATURELY FROM 8 YEARS TO 5 YEARS. THE RIGHT VENTRICULAR (RV) PACE PERCENTAGE WAS LOW. DUE TO ATRIAL TACHYCARDIA, THE ONLY ALTERNATIVE WAS ATRIAL ASYNCHRONOUS PACING (AAI), WHICH INCREASED POWER CONSUMPTION. THIS PACEMAKER REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20729696&pc=LWP | 2124215-2024-72939 | 2023-01-18 05:00 | Injury | BOSTON SCIENTIFIC CORPORATION | 2024-11-19 05:00 | LWP | VITALIO | Incorrect, Inadequate or Imprecise Result or Readings; Battery Problem; Data Problem | No Clinical Signs, Symptoms or Conditions | P150012/S000 |  | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER EXHIBITED FAULT CODE 1010. THERAPY HISTORY CORRUPTION WAS DETECTED AND PREVIOUSLY STORED THERAPY DATA WAS DELETED. THE DEVICE ALSO DISPLAYED BATTERY LONGEVITY AT ELEVEN YEARS DESPITE THE LAST REMAINING KNOWN PROPER LONGEVITY ESTIMATE BEING TWO YEARS. THE PACEMAKER SYSTEM WAS LATER EXPLANTED AND REPLACED DUE TO NORMAL BATTERY DEPLETION. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20708864&pc=LWP | 2124215-2024-72375 | 2024-11-02 04:00 | Malfunction | BOSTON SCIENTIFIC CORPORATION | 2024-11-18 05:00 | LWP | ACCOLADE MRI EL DR | Premature Discharge of Battery; Battery Problem | No Clinical Signs, Symptoms or Conditions | P150012 |  | 1 | Event Description: IT WAS REPORTED THE PACEMAKER RECORDED A CODE 1003 INDICATIVE OF BATTERY VOLTAGE TOO LOW FOR THE PROJECTED REMAINING CAPACITY. A REQUEST WAS MADE TO HAVE DATA FROM THIS DEVICE ANALYZED. DATA ANALYSIS SHOWED THE BATTERY APPEARED TO BE DEPLETING MORE QUICKLY THAN EXPECTED. DEVICE REPLACEMENT INTERVAL WAS PROVIDED. THIS DEVICE CURRENTLY REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED.Event Description: IT WAS REPORTED THE PACEMAKER RECORDED A CODE 1003 INDICATIVE OF BATTERY VOLTAGE TOO LOW FOR THE PROJECTED REMAINING CAPACITY. A REQUEST WAS MADE TO HAVE DATA FROM THIS DEVICE ANALYZED. DATA ANALYSIS SHOWED THE BATTERY APPEARED TO BE DEPLETING MORE QUICKLY THAN EXPECTED. DEVICE REPLACEMENT INTERVAL WAS PROVIDED. THIS DEVICE CURRENTLY REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED. ADDITIONAL INFORMATION WAS RECEIVED. THE DEVICE WAS EXPLANTED AND REPLACED. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20715895&pc=LWP | 2124215-2024-72519 | 2024-10-22 04:00 | Injury | BOSTON SCIENTIFIC CORPORATION | 2024-11-18 05:00 | LWP | ESSENTIO MRI EL DR | Incorrect, Inadequate or Imprecise Result or Readings; Battery Problem; Interrogation Problem | Arrhythmia; Dyspnea; Dizziness; Unspecified Heart Problem | P150012 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT IMPLANTED WITH THIS PACEMAKER SYSTEM PRESENTED TO THE EMERGENCY ROOM (ER) WITH SYMPTOMS OF DIZZINESS, SHORTNESS OF BREATH, AND A RAPID HEART RATE. THE PATIENT ALSO REPORTED A PULSING IN HER CHEST THAT WAS EXTERNALLY VISIBLE. PROGRAMMER INTERROGATION VIA CONSULT WAS UNSUCCESSFUL, EVEN AFTER UNPLUGGING/REPLUGGING THE WAND AND MOVING THE WAND IN CONCENTRIC CIRCLES. THE PATIENT NOTED THAT A DEVICE CHECK IN (B)(6) 2024 INDICATED THERE WAS APPROXIMATELY THREE YEARS REMAINING ON THE BATTERY. IT WAS DETERMINED THAT THE PACEMAKER HAD ENTERED SAFETY MODE, AND DEVICE REPLACEMENT WAS SCHEDULED. DURING THE DEVICE REPLACEMENT PROCEDURE, THE CHRONIC RIGHT VENTRICULAR (RV) LEAD WAS DISCONNECTED, AND PLUGGED INTO THE NEW DEVICE. NO PACING WAS NOTED FOR THIS PACER DEPENDENT PATIENT. THE RV LEAD WAS THEN UNPLUGGED FROM THE NEW DEVICE, AND CONNECTED TO THE PACING SYSTEM ANALYZER (PSA) CABLES. ONCE CONNECTED TO THE PSA, THE RV LEAD EXHIBITED NON-CAPTURE AT MAXIMUM OUTPUT. THE RV LEAD WAS THEN RECONNECTED TO THE OLD DEVICE, PLACED BACK INTO THE POCKET, AND PACING RESUMED. IT WAS NOTED THAT DURING THE PROCEDURE, THE PATIENT WAS RECEIVING TEMPORARY PACING VIA EXTERNAL PADS DUE TO ASYSTOLE. IT WAS NOTED THAT THE OLD PACEMAKER SYSTEM HAD NOT CONNECTED TO THE REMOTE MONITORING SYSTEM SINCE (B)(6) 2024 , THUS DEVICE DATA WAS NOT AVAILABLE FOR REVIEW TO EVALUATE PAST RV LEAD BEHAVIOR. SUBSEQUENTLY, THE RV LEAD WAS SURGICALLY ABANDONED AND REPLACED DURING THE DEVICE CHANGEOUT PROCEDURE AS WELL. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. THE EXPLANTED PACEMAKER WAS RETURNED FOR ANALYSIS. Manufacturer Narrative: THE PRODUCT HAS BEEN RECEIVED FOR ANALYSIS. THIS REPORT WILL BE UPDATED UPON COMPLETION OF ANALYSIS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20696210&pc=LWP | 2124215-2024-72005 | 2024-11-04 05:00 | Malfunction | BOSTON SCIENTIFIC CORPORATION | 2024-11-15 05:00 | LWP | PROPONENT MRI EL DR | Premature Discharge of Battery; Battery Problem | No Clinical Signs, Symptoms or Conditions | P150012/S000 |  | 1 | Event Description: IT WAS REPORTED THAT THERE WAS A CONCERN OF THIS PACEMAKER'S LONGEVITY AND EXHIBITED PREMATURE BATTERY DEPLETION (PBD). THE HEALTH CARE PROFESSIONAL (HCP) REQUESTED ANALYSIS OF THE DEVICE; HOWEVER, THE HCP DID NOT SEND THE FILES FOR ANALYSIS. TS PROVIDED INSTRUCTIONS ON HOW TO OBTAIN THE FILES. THE DEVICE REMAINS IN USE. NO ADVERSE PATIENT EFFECTS WERE REPORTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20684460&pc=LWP | 2124215-2024-71630 | 2024-11-04 05:00 | Injury | BOSTON SCIENTIFIC CORPORATION | 2024-11-14 05:00 | LWP | ACCOLADE MRI EL DR | Premature Discharge of Battery; Battery Problem | No Clinical Signs, Symptoms or Conditions | P150012/S000 |  | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS SUSPECTED OF EXHIBITING PREMATURE BATTERY DEPLETION, AS A DECREASE IN ESTIMATED BATTERY LONGEVITY WAS NOTED BETWEEN FOLLOW-UP VISITS. TECHNICAL SERVICES REVIEWED THE DEVICE DATA AND CONFIRMED THESE OBSERVATIONS. IT WAS RECOMMENDED THAT THE DEVICE BE REPLACED WITHIN 90 DAYS. CURRENTLY, THERAPY DELIVERY REMAINS UNAFFECTED, AND THE DEVICE CONTINUES TO BE IMPLANTED AND IN SERVICE WITH NO REPORTED ADVERSE EFFECTS ON THE PATIENT. ADDITIONAL INFORMATION: THIS DEVICE WAS EXPLANTED AND REPLACED. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. THIS DEVICE IS ANTICIPATED TO BE RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20686037&pc=LWP | 2124215-2024-71711 | 2024-09-17 04:00 | Malfunction | BOSTON SCIENTIFIC CORPORATION | 2024-11-14 05:00 | LWP | ESSENTIO MRI DR | Battery Problem | No Clinical Signs, Symptoms or Conditions | P150012 |  | 1 | Event Description: IT WAS REPORTED THAT THIS IMPLANTABLE PACEMAKER DEVICE EXHIBITED A BATTERY FLUCTUATION. IN ADDITION, THIS DEVICE WAS AT LESS THAN 3 MONTHS REMAINING THE FOLLOWING DAY AFTER THIS DEVICE HAS REACHED TO ELECTIVE REPLACEMENT INDICATOR (ERI). DATA ANALYSIS WAS REQUESTED AND SHOWED THAT THE POWER CONSUMPTION HAS DECREASED FROM 54 MICRO WATTS TO 43 MICRO WATTS AND THE BATTERY RECOVERED. TECHNICAL SERVICES CONSULTANT RECOMMENDED REPLACEMENT. AS OF THIS TIME, THIS DEVICE REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20691047&pc=LWP | 2124215-2024-71815 | 2024-05-31 04:00 | Injury | BOSTON SCIENTIFIC CORPORATION | 2024-11-14 05:00 | LWP | ACCOLADE MRI EL DR | Premature Discharge of Battery; Battery Problem | No Clinical Signs, Symptoms or Conditions | P150012 |  | 1 | Event Description: IT WAS REPORTED THAT THE BATTERY OF THIS PACEMAKER WAS SUSPECTED TO HAVE DEPLETED PREMATURELY. A REQUEST WAS MADE TO HAVE DATA FROM THIS DEVICE ANALYZED AND DATA ANALYSIS CONFIRMED THE BATTERY APPEARED TO HAVE DEPLETED MORE QUICKLY THAN EXPECTED; THE DEVICE WAS AT END-OF-LIFE. THERE WAS ALSO A RECENT MESSAGE TO CHECK THE ATRIAL LEAD, AND IT WAS NOTED THAT THE AUTOMATIC THRESHOLD TEST HAD BEEN PREVIOUSLY SUSPENDED IN APRIL 2023. TECHNICAL SERVICES DISCUSSED THAT GIVEN THE SHORT LONGEVITY OF THE DEVICE DUE TO AN UNKNOWN FACTOR IT WOULD BE APPROPRIATE TO ASSUME THAT THE ATRIAL LEAD COULD HAVE BEEN DAMAGED. THE PACEMAKER WAS REPLACED; HOWEVER, IT WAS DECIDED TO LEAVE THE ATRIAL LEAD IMPLANTED AND IN SERVICE. THE PACEMAKER WAS RECEIVED FOR ANALYSIS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20673529&pc=LWP | 2124215-2024-71295 | 2024-09-13 04:00 | Malfunction | BOSTON SCIENTIFIC CORPORATION | 2024-11-13 05:00 | LWP | PROPONENT MRI DR | Premature Discharge of Battery; Battery Problem; High Capture Threshold | No Clinical Signs, Symptoms or Conditions | P150012/S000 |  | 1 | Event Description: IT WAS REPORTED THAT THERE WERE CONCERNS FOR PREMATURE BATTERY DEPLETION (PBD) FOR THIS PACEMAKER DEVICE. THE HEALTH CARE PROFESSIONAL (HCP) CALLED IN FOR ASSISTANCE IF THERE WOULD BE ANY CHANCE ON EXTENDING THE BATTERY LONGEVITY TO PREVENT THE DEVICE HITTING ELECTIVE REPLACEMENT INDICATOR (ERI). THE HCP REQUESTED DATA ANALYSIS ON THIS DEVICE. UPON REVIEW, THE RIGHT VENTRICULAR (RV) LEAD THRESHOLD MEASUREMENT WAS 2.5V @ 1.0MS AND AUTO THRESHOLD WAS APPARENTLY STRUGGLING TO OPTIMIZE AND THEREFORE PROGRAMMED OUTPUT TO 5.5V @ 1.0MS. TROUBLESHOOTING AND DEVICE OPTIMIZATION OPTIONS WERE PROVIDED, WHICH WOULD ADD AT LEAST A YEAR TO THE BATTERY LONGEVITY AND POSSIBLY ELIMINATE THE NEED TO REPLACE THE DEVICE. THIS DEVICE CURRENTLY REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20675215&pc=LWP | 2124215-2024-71364 | 2024-10-17 04:00 | Malfunction | BOSTON SCIENTIFIC CORPORATION | 2024-11-13 05:00 | LWP | ESSENTIO SR | Premature Discharge of Battery; Battery Problem | No Clinical Signs, Symptoms or Conditions | N970003 |  | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER RECORDED A CODE 1003 INDICATIVE OF BATTERY VOLTAGE TOO LOW FOR THE PROJECTED REMAINING CAPACITY. IT WAS REPORTED THAT A REQUEST WAS MADE TO HAVE DATA FROM THIS DEVICE ANALYZED. DATA ANALYSIS SHOWED THE BATTERY APPEARED TO BE DEPLETING MORE QUICKLY THAN EXPECTED. TECHNICAL SERVICES (TS) RECOMMENDED THAT THE DEVICE EITHER GET REPLACED OR HAVE BATTERY STATUS RE-REVALUATED IN NINETY DAYS. THIS DEVICE REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20667615&pc=LWP | 2124215-2024-71018 | 2024-10-16 04:00 | Malfunction | BOSTON SCIENTIFIC CORPORATION | 2024-11-12 05:00 | LWP | ACCOLADE MRI EL DR | Battery Problem | No Clinical Signs, Symptoms or Conditions | P150012 |  | 1 | Event Description: IT WAS REPORTED THAT THIS IMPLANTABLE PACEMAKER EXHIBITED A LOW BATTERY. AS OF THIS TIME, THIS DEVICE REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20670471&pc=LWP | 2124215-2024-71200 | 2024-10-14 04:00 | Malfunction | BOSTON SCIENTIFIC CORPORATION | 2024-11-12 05:00 | LWP | ACCOLADE MRI EL DR | Premature Discharge of Battery; Battery Problem | No Clinical Signs, Symptoms or Conditions | P150012/S000 |  | 1 | Event Description: IT WAS REPORTED THAT THIS IMPLANTABLE PACEMAKER WAS SUSPECTED OF PREMATURE BATTERY DEPLETION (PBD). THE DEVICE BATTERY LONGEVITY DROPPED FROM 6 YEARS TO 3 YEARS IN A SPAN OF 1.5 YEARS. A REQUEST WAS MADE TO HAVE DATA FROM THE DEVICE ANALYZED. DATA ANALYSIS CONFIRMED PREMATURE BATTERY DEPLETION (PBD) AND THAT THIS DEVICE POWER CONSUMPTION WAS INCREASING DURING THE PAST YEAR AND IS EXPECTED TO CONTINUE TO INCREASE. DEVICE REPLACEMENT WAS RECOMMENDED. TO DATE, THIS DEVICE REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20659573&pc=LWP | 2124215-2024-70819 | 2024-10-21 04:00 | Malfunction | BOSTON SCIENTIFIC CORPORATION | 2024-11-11 05:00 | LWP | PROPONENT MRI DR | Premature Discharge of Battery; Battery Problem | No Clinical Signs, Symptoms or Conditions | P150012/S000 |  | 1 | Event Description: IT WAS REPORTED THAT THE BATTERY OF THIS PACEMAKER WAS SUSPECTED TO BE DEPLETING PREMATURELY AS THE ESTIMATED REMAINING LONGEVITY DECREASED FROM TWO AND A HALF YEARS TO LESS THAN FOUR MONTHS IN APPROXIMATELY 14 MONTHS. A REQUEST WAS MADE TO HAVE DATA FROM THIS DEVICE ANALYZED, BUT TECHNICAL SERVICES (TS) INDICATED THAT AT THIS TIME, DATA IS NOT AVAILABLE. TS REQUESTED THE REPRESENTATIVE INVOLVED TO INTERROGATE THE DEVICE TO OBTAIN UPDATED DATA. AT THIS TIME, THERE IS NO EVIDENCE TO SUGGEST THAT THIS REQUEST HAS BEEN COMPLETED. NO ADVERSE PATIENT EFFECTS WERE REPORTED. THE DEVICE REMAINS IN SERVICE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20660328&pc=LWP | 2124215-2024-70859 | 2024-09-24 04:00 | Malfunction | BOSTON SCIENTIFIC CORPORATION | 2024-11-11 05:00 | LWP | ACCOLADE MRI EL DR | Premature Discharge of Battery; Battery Problem | No Clinical Signs, Symptoms or Conditions | P150012/S000 |  | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER DEVICE EXHIBITED A DROP IN ESTIMATED BATTERY LONGEVITY. IN 2021, THERE WAS 7 YEARS REMAINING, IN 2022, THERE WAS 3.5 YEARS REMAINING IN 2023, AND 1 YEAR REMAINING IN 2024. THE PHYSICIAN SCHEDULED THE PATIENT FOR A FOLLOW UP APPOINTMENT IN THE NEAR FUTURE TO RECHECK THE DEVICE. AT THIS TIME THE DEVICE REMAINS IMPLANTED. NO ADVERSE PATIENT EFFECTS WERE REPORTED. ATTEMPTS TO OBTAIN THE SERIAL NUMBER OF THE DEVICE, HAVE BEEN UNSUCCESSFUL. Manufacturer Narrative: IF PERTINENT INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20638786&pc=LWP | 2124215-2024-70354 | 2024-10-09 04:00 | Injury | BOSTON SCIENTIFIC CORPORATION | 2024-11-08 05:00 | LWP | ACCOLADE MRI DR | Battery Problem | No Clinical Signs, Symptoms or Conditions | P150012 |  | 1 | Event Description: IT WAS REPORTED THAT THE IMPLANTABLE PACEMAKER REACHED END OF LIFE (EOL) IN A SPAN OF 6 YEARS. THIS PACEMAKER WAS EXPLANTED AND NEW PACEMAKER WAS IMPLANTED. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20640962&pc=LWP | 2124215-2024-70423 | 2024-11-04 05:00 | Malfunction | BOSTON SCIENTIFIC CORPORATION | 2024-11-08 05:00 | LWP | ADVANTIO | Battery Problem | No Clinical Signs, Symptoms or Conditions |  |  | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER DEVICE EXHIBITED LONGEVITY DROP TO CURRENTLY REMAINING AT ONE YEAR. THE PATIENT IS NOT DEPENDENT ON THIS PACEMAKER SYSTEM, SO THE PLAN IS TO CONTINUE MONITORING AND CHECK THE PATIENT AGAIN IN THREE MONTHS. THE PHYSICIAN STATED THAT THE BATTERY CONSUMPTION WAS INCREASED IN LAST YEAR AND REQUESTED TO INVESTIGATE MORE. THIS DEVICE REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20626323&pc=LWP | 2124215-2024-70044 | 2024-10-16 04:00 | Injury | BOSTON SCIENTIFIC CORPORATION | 2024-11-07 05:00 | LWP | INGENIO | Premature Discharge of Battery; Battery Problem | No Clinical Signs, Symptoms or Conditions |  |  | 1 | Event Description: IT WAS REPORTED THAT THIS IMPLANTABLE PACEMAKER WAS SUSPECTED TO HAVE PREMATURE BATTERY DEPLETION (PBD). FURTHERMORE, THERE WERE DIFFERING LONGEVITY ESTIMATES GIVEN OVER THE FIVE MONTHS. A REQUEST WAS MADE TO HAVE DATA FROM THIS DEVICE ANALYZED. NO ADDITIONAL INFORMATION AT THIS TIME. THE DEVICE REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED. RECEIVED ADDITIONAL INFORMATION THE DEVICE WAS EXPLANTED AND REPLACED DUE TO THE SUSPECTED PBD BEFORE DATA ANALYSIS WAS COMPLETED. DATA ANALYSIS WAS PERFORMED, AND THE PACEMAKER HAS HAD NO UNUSUAL FAULTS OR RESETS, THE POWER CONSUMPTION IS STABLE AND WITHIN EXPECTATIONS BASED ON PROGRAMMING AND THERAPY DELIVERY WITH NO EVIDENCE OF PREMATURE BATTERY DEPLETION (PBD). TECHNICAL SERVICES (TS) PROVIDED INFORMATION ON THE DIFFERING LONGEVITY ESTIMATES AND ADVISED THIS IS NORMAL DEVICE EXPECTATION. ADDITIONAL INFORMATION HAS BEEN REQUESTED ON THE DEVICE EXPLANT. OUTSIDE OF THE REVISION, NO ADVERSE PATIENT EFFECTS WERE REPORTED.Event Description: IT WAS REPORTED THAT THIS IMPLANTABLE PACEMAKER WAS SUSPECTED TO HAVE PREMATURE BATTERY DEPLETION (PBD). FURTHERMORE, THERE WERE DIFFERING LONGEVITY ESTIMATES GIVEN OVER THE FIVE MONTHS. A REQUEST WAS MADE TO HAVE DATA FROM THIS DEVICE ANALYZED. NO ADDITIONAL INFORMATION AT THIS TIME. THE DEVICE REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED. RECEIVED ADDITIONAL INFORMATION THE DEVICE WAS EXPLANTED AND REPLACED DUE TO THE SUSPECTED PBD BEFORE DATA ANALYSIS WAS COMPLETED. DATA ANALYSIS WAS PERFORMED, AND THE PACEMAKER HAS HAD NO UNUSUAL FAULTS OR RESETS, THE POWER CONSUMPTION IS STABLE AND WITHIN EXPECTATIONS BASED ON PROGRAMMING AND THERAPY DELIVERY WITH NO EVIDENCE OF PREMATURE BATTERY DEPLETION (PBD). TECHNICAL SERVICES (TS) PROVIDED INFORMATION ON THE DIFFERING LONGEVITY ESTIMATES AND ADVISED THIS IS NORMAL DEVICE EXPECTATION. ADDITIONAL INFORMATION HAS BEEN REQUESTED ON THE DEVICE EXPLANT. OUTSIDE OF THE REVISION, NO ADVERSE PATIENT EFFECTS WERE REPORTED. ADDITIONAL INFORMATION RECEIVED ADVISED THE DEVICE WAS DISPOSED OF AND WILL NOT RETURN FOR ANALYSIS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20627311&pc=LWP | 2124215-2024-70065 | 2024-07-01 04:00 | Injury | BOSTON SCIENTIFIC CORPORATION | 2024-11-07 05:00 | LWP | ENERGEN CRT-D | Premature Discharge of Battery; Battery Problem | No Clinical Signs, Symptoms or Conditions |  |  | 1 | Event Description: IT WAS REPORTED THAT THE BATTERY OF THIS CARDIAC RESYNCHRONIZATION THERAPY DEFIBRILLATOR (CRT-D) WAS SUSPECTED TO BE DEPLETING PREMATURELY. WITHIN THE LAST YEAR, THE ESTIMATED LONGEVITY DECREASED FROM 1.5 YEARS, TO EOL. A REQUEST WAS MADE TO HAVE DATA FROM THIS DEVICE ANALYZED, HOWEVER DATA ANALYSIS IS NOT YET AVAILABLE. THIS CRT-D WAS SUBSEQUENTLY REPLACED, BUT UNAVAILABLE FOR RETURN. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. ADDITIONAL INFORMATION WAS RECEIVED. A REQUEST WAS MADE TO HAVE DATA FROM THIS DEVICE ANALYZED. DATA ANALYSIS FOUND THIS PACEMAKER TO HAVE NORMAL BATTERY DEPLETION. NO ADVERSE PATIENT EFFECTS WERE REPORTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20629723&pc=LWP | 2124215-2024-70093 | 2024-09-20 04:00 | Injury | BOSTON SCIENTIFIC CORPORATION | 2024-11-07 05:00 | LWP | INCEPTA ICD | Premature Discharge of Battery; Battery Problem | No Clinical Signs, Symptoms or Conditions | P960040 |  | 1 | Event Description: IT WAS REPORTED THAT THIS IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) DROPPED ITS BATTERY LONGEVITY FROM 2 YEARS TO ELECTIVE REPLACEMENT INDICATOR (ERI) IN A SPAN OF 7 MONTHS. A REQUEST WAS MADE TO HAVE DATA FROM THIS DEVICE ANALYZED. DATA ANALYSIS CONFIRMED LOW VOLTAGE FAULT WAS TRIGGERED AND THAT THE DEVICE POWER CONSUMPTION WAS INCREASING IN A GRADUAL FASHION. FURTHERMORE, DEVICE REPLACEMENT WAS RECOMMENDED. THIS DEVICE REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20631050&pc=LWP | 2124215-2024-70131 | 2024-10-08 04:00 | Malfunction | BOSTON SCIENTIFIC CORPORATION | 2024-11-07 05:00 | LWP | INGENIO | Battery Problem | No Clinical Signs, Symptoms or Conditions |  |  | 1 | Event Description: IT WAS REPORTED THAT THIS IMPLANTABLE PACEMAKER DEVICE DROPPED IT BATTERY LONGEVITY FOR 1 YEAR IN JUST 5 MONTHS. BOSTON SCIENTIFIC TECHNICAL SERVICE (TS) CONFIRMED FROM THE HEALTHCARE PROFESSIONAL (HCP) THE DEVICE SETTING. PATIENT WAS PACING AT 50 PERCENT IN THE ATRIAL AND 100 PERCENT IN VENTRICULAR. ALSO, HIGH THRESHOLD IN VENTRICULAR WAS NOTED. TS THEN DISCUSSED TO HAVE ANALYSIS FOR ASSESS DEVICE DEPLETION OR THIS COULD BE DEVICE SWITCHED TO USING VOLTAGE-BASED ASSESSMENT IN TIME REMAINING WHICH CAN CAUSE THIS CHANGE IN TIME REMAINING. THIS DEVICE REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20635401&pc=LWP | 2124215-2024-70238 | 2024-10-24 04:00 | Injury | BOSTON SCIENTIFIC CORPORATION | 2024-11-07 05:00 | LWP | ACCOLADE MRI EL DR | Premature Discharge of Battery; Battery Problem | No Clinical Signs, Symptoms or Conditions | P150012 |  | 1 | Event Description: IT WAS REPORTED THAT THE BATTERY OF THIS PACEMAKER WAS SUSPECTED TO BE DEPLETING PREMATURELY. THE PACEMAKER WAS SHOWING A RECOMMENDED SAFE REPLACEMENT INTERVAL EQUAL TO 90 DAYS. THERAPY DELIVERY IS CURRENTLY UNAFFECTED. ULTIMATELY, THIS PACEMAKER WAS EXPLANTED AND A NEW DEVICE WAS IMPLANTED WHICH REMAINS IN SERVICE. THIS PACEMAKER IS IN THE HOSPITAL'S POSSESSION. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20619987&pc=LWP | 2124215-2024-69758 | 2024-10-07 04:00 | Injury | BOSTON SCIENTIFIC CORPORATION | 2024-11-06 05:00 | LWP | ACCOLADE EL DR | Premature Discharge of Battery; Battery Problem | No Clinical Signs, Symptoms or Conditions | N970003 |  | 1 | Event Description: IT WAS REPORTED THAT THIS IMPLANTABLE PACEMAKER LONGEVITY DROPPED FROM 3 YEARS TO 8 MONTHS IN A SPAN OF 2 MONTHS. A REQUEST WAS MADE TO HAVE DATA FROM THIS DEVICE ANALYZED. DATA ANALYSIS CONFIRMED THAT THE POWER CONSUMPTION IS INCREASING IN A GRADUAL FASHION. THIS DEVICE WAS EXPLANTED AND REPLACED. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20620207&pc=LWP | 2124215-2024-69770 | 2024-10-30 04:00 | Malfunction | BOSTON SCIENTIFIC CORPORATION | 2024-11-06 05:00 | LWP | PROPONENT MRI EL DR | Premature Discharge of Battery; Battery Problem | No Clinical Signs, Symptoms or Conditions | P150012/S000 |  | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER DEVICE WAS DEPLETING PREMATURELY. DURING THE FOLLOW-UP VISIT IT WAS NOTED THAT THERE WAS AN INCREASE IN ENERGY CONSUMPTION FROM LAST FOLLOW-UP AND THE PROJECTED LONGEVITY HAD DECREASED FROM 7 YR TO 3 YR REMAINING. THIS DEVICE REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20624373&pc=LWP | 2124215-2024-69903 | 2024-11-04 05:00 | Malfunction | BOSTON SCIENTIFIC CORPORATION | 2024-11-06 05:00 | LWP | ACCOLADE MRI DR | Premature Discharge of Battery; Battery Problem | No Clinical Signs, Symptoms or Conditions | P150012/S000 |  | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS SUSPECTED TO BE DEPLETING PREMATURELY. IT WAS NOTED THAT THE PATIENT HAD A RECENT ONSET OF ATRIAL FIBRILLATION (AF). THE HEALTH CARE PROFESSIONAL (HCP) OPTED NOT TO HAVE DATA ANALYSIS PERFORMED AT THIS TIME AND PLANS TO CONTINUE MONITORING. NO ADVERSE PATIENT EFFECTS WERE REPORTED. THIS DEVICE REMAINS IN SERVICE. IF PERTINENT INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20597253&pc=LWP | 2124215-2024-68922 | 2024-01-08 05:00 | Injury | BOSTON SCIENTIFIC CORPORATION | 2024-11-04 05:00 | LWP | ACCOLADE SR | Battery Problem | No Clinical Signs, Symptoms or Conditions | N970003/S167 |  | 1 | Event Description: IT WAS REPORTED THAT THIS IMPLANTABLE PACEMAKER BATTERY LONGEVITY WENT FROM THREE TO TWO YEARS IN JUST THREE MONTHS. DATA ANALYSIS CONFIRMED THAT THIS DEVICE WAS DEPLETING NORMALLY. THE POWER CONSUMPTION IS STABLE, AND VOLTAGE IS TRACKING NORMALLY, AND THE TRANSITION IN THE BLENDING PERIOD RESULTED IN THE DROP IN ESTIMATED LONGEVITY REMAINING. THIS DEVICE WAS EXPLANTED AND REPLACED. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20604047&pc=LWP | 2124215-2024-69191 | 2024-10-12 04:00 | Malfunction | BOSTON SCIENTIFIC CORPORATION | 2024-11-04 05:00 | LWP | ACCOLADE MRI EL DR | Premature Discharge of Battery; Battery Problem | No Clinical Signs, Symptoms or Conditions | P150012 |  | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER RECORDED A CODE 1003 INDICATIVE OF BATTERY VOLTAGE TOO LOW FOR THE PROJECTED REMAINING CAPACITY. IT WAS REPORTED THAT A REQUEST WAS MADE TO HAVE DATA FROM THIS DEVICE ANALYZED. DATA ANALYSIS SHOWED THE BATTERY APPEARED TO BE DEPLETING MORE QUICKLY THAN EXPECTED. DEVICE REPLACEMENT WAS RECOMMENDED. AT THIS TIME, THIS PACEMAKER REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED.Event Description: IT WAS REPORTED THAT THIS PACEMAKER RECORDED A CODE 1003 INDICATIVE OF BATTERY VOLTAGE TOO LOW FOR THE PROJECTED REMAINING CAPACITY. IT WAS REPORTED THAT A REQUEST WAS MADE TO HAVE DATA FROM THIS DEVICE ANALYZED. DATA ANALYSIS SHOWED THE BATTERY APPEARED TO BE DEPLETING MORE QUICKLY THAN EXPECTED. DEVICE REPLACEMENT WAS RECOMMENDED. AT THIS TIME, THIS PACEMAKER REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED. ADDITIONAL INFORMATION WAS RECEIVED. THE DEVICE WAS EXPLANTED AND REPLACED. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20591885&pc=LWP | 2124215-2024-68800 | 2024-10-03 04:00 | Malfunction | BOSTON SCIENTIFIC CORPORATION | 2024-11-01 04:00 | LWP | ACCOLADE MRI EL DR | Battery Problem | No Clinical Signs, Symptoms or Conditions | P150012 |  | 1 | Event Description: IT WAS REPORTED THAT THIS IMPLANTABLE PACEMAKER DEVICE EXHIBITED HIGH POWER CONSUMPTION AT 99 MICROWATTS, 232 PERCENT. BOSTON SCIENTIFIC TECHNICAL SERVICE (TS) REVIEWED AND NOTED RIGHT VENTRICULAR (RV) OUTPUT WAS SET TO 5 VOLTS AND PATIENT IS 100 PERCENT PACED. TO DATE, THIS DEVICE REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED.