## FDA Data
event_type | date_of_event | product_problems | initial_text | follow_up | manufacturer_narrative | device_generic_name | device_manufacturer_name | device_model_number | device_lot_number | device_report_product_code | expiration_date_of_device | date_returned_to_manufacturer | device_availability | device_operator | device_name
Malfunction | 20200418 | ['Battery Problem', 'Power Problem'] | INFORMATION RECEIVED BY MEDTRONIC INDICATED THAT THE INSULIN PUMP ALARMED WITH ALARM GENERATED WHEN A POWER SYSTEM ERROR WAS DETECTED AND THIS ERROR DOES NOT PREVENT THE PUMP FROM RUNNING. NO HARM REQUIRING MEDICAL INTERVENTION WAS REPORTED. THE CUSTOMER ALSO REPORTED EXPERIENCING POWER ERROR FOR 3 DAYS. CUSTOMER WAS ABLE TO SUCCESSFULLY CLEAR POWER ERROR ALARM. CUSTOMER WAS ADVISED TO REVERT TO THE BACKUP PLAN. THE INSULIN PUMP WILL BE RETURNED FOR ANALYSIS. |  | (B)(4). CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. THE DEVICE WILL BE RETURNED FOR ANALYSIS AND FURTHER INFORMATION WILL FOLLOW ONCE THE ANALYSIS HAS BEEN COMPLETED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. DEVICE PASSED SLEEP CURRENT MEASUREMENT, ACTIVE CURRENT MEASUREMENT, SELF TEST AND DISPLACEMENT TEST. HOWEVER, DEVICE TRACE DOWNLOAD ANALYSIS CONFIRMED THE INSULIN PUMP ALARMED POWER ERROR 25, POWER LOSS ALARM, REPLACE BATTERY ALERT AND REPLACE BATTERY NOW ALARM. THE ADAPT TOOL CONFIRMED POWER ERROR 25, POWER LOSS ALARM, REPLACE BATTERY ALERT, REPLACE BATTERY NOW ALARM WAS TRIGGERED WHEN THE BACKUP BATTERY LOADED VOLTAGE (LOADED VLITH) WAS LESS THAN 3.5V FOR 4 CONSECUTIVE HOURS DUE TO CONNECTOR RESISTANCE ELECTRICAL BOARD. AFTER DISCONNECTING AND RECONNECTING THE INTERNAL BATTERY CONNECTOR ON ELECTRICAL BOARD, DEVICE WAS MONITORED AND FUNCTIONED PROPERLY. | ARTIFICIAL PANCREAS DEVICE SYSTEM, THRESHOLD SUSPEND | MEDTRONIC PUERTO RICO OPERATIONS CO. | MMT-1715K | HG1B44Y | OZO |  | 20200429 | Device was returned to manufacturer | LAY USER/PATIENT | Automated Insulin Dosing , Threshold Suspend
Malfunction | 20200421 | ['Excess Flow or Over-Infusion', 'Battery Problem', 'Adverse Event Without Identified Device or Use Problem', 'Appropriate Term/Code Not Available'] | INFORMATION RECEIVED BY MEDTRONIC INDICATED THAT THE INSULIN PUMP AND RESERVOIR HAD OVER DELIVERY. NO HARM REQUIRING MEDICAL INTERVENTION WAS REPORTED. TROUBLESHOOT WAS NOT PERFORMED. THE INSULIN PUMP AND RESERVOIR WILL NOT BE RETURNING FOR ANALYSIS. |  | CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. | PUMP, INFUSION, INSULIN, TO BE USED WITH INVASIVE GLUCOSE SENSOR | MEDTRONIC MINIMED | MMT-XXX |  | OYC |  |  | No | LAY USER/PATIENT | Pump, Infusion, Insulin, To Be Used With Invasive Glucose Sensor
Injury | 20200407 | ['Failure to Charge', 'Battery Problem'] | FOLLOWING THE BATTERY PERFORMANCE ALERT (BPA) ADVISORY, A BPA WAS RECEIVED BY THE CLINICIAN. THE DEVICE WAS EXPLANTED AND REPLACED. THE DEVICE WAS FOUND TO BE AT ELECTIVE REPLACEMENT INDICATOR (ERI) AND CAPACITOR EXCEEDED THE CHARGED TIME. THE PATIENT WAS RECOVERING. |  | THE RESULTS/METHOD AND CONCLUSION CODES ALONG WITH INVESTIGATION RESULTS WILL BE PROVIDED IN THE FINAL REPORT. THE DEVICE IS INCLUDED IN THE BATTERY PERFORMANCE ALERT ADVISORY ISSUED BY ABBOTT ON 28 AUGUST 2017. THE REPORTED EVENT OF CHARGE TIME OUT WAS CONFIRMED. THE CAUSE OF CHARGE TIME OUT WAS DUE TO PREMATURE BATTERY DEPLETION. PREMATURE BATTERY DEPLETION WAS CONFIRMED BY ANALYSIS. NO SOURCES OF HIGH CURRENT WERE NOTED. THE CAUSE OF THE PREMATURE BATTERY DEPLETION WAS CONSISTENT WITH LI CLUSTER FORMATION. FROM THESE ANALYSES, IN THE ABSENCE OF HIGH CURRENT DRAW, IT IS PROBABLE THAT THE PREMATURE BATTERY DEPLETION WAS CAUSED BY A LITHIUM CLUSTER INDUCED SHORT CIRCUIT. LI CLUSTERS ARE A KNOWN DEPLETION MECHANISM FOR THESE ADVISORY PRODUCTS THAT HAS BEEN INVESTIGATED AND ASSOCIATED WITH A FIELD ACTION IN OCTOBER 2016. ADDITIONAL INFORMATION: D10, H7, H9 | IMPLANTABLE CARDIOVERTER DEFIBRILLATOR | ST. JUDE MEDICAL, INC.(CRM-SUNNYVALE) | CD1357-40Q | 4443992 | LWS | 20160131 | 20200422 | Device was returned to manufacturer | HEALTH PROFESSIONAL | Implantable Cardioverter Defibrillator (Non-Crt)
Malfunction | 20200402 | ['Battery Problem'] | IT WAS REPORTED THAT THE PUMP BATTERY WAS DEPLETING QUICKLY. THERE WAS NO REPORTED IMPACT TO THE CUSTOMER¿S BLOOD GLUCOSE LEVEL. THE CUSTOMER DECLINED TO PERFORM TROUBLESHOOTING AND DECLINED FOLLOW UP FROM TANDEM TECHNICAL SUPPORT. |  | NO PRODUCT WAS RETURNED FOR EVALUATION. SHOULD NEW RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. | CONTINUOUS GLUCOSE MONITOR | TANDEM DIABETES CARE | 1000096 |  | OYC |  |  | No | LAY USER/PATIENT | Pump, Infusion, Insulin, To Be Used With Invasive Glucose Sensor
Malfunction | 20200414 | ['Battery Problem'] | IT WAS REPORTED THAT THIS DEVICE TRIPPED ERI ON (B)(6) 2020. THE DATA ANALYSIS PERFORMED BY THE MANUFACTURER INDICATED A BATTERY ISSUE. NO ADVERSE PATIENT EVENTS WERE REPORTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THIS FILE WILL BE UPDATED. |  | THE DEVICE WAS NOT RETURNED FOR ANALYSIS. THE ANALYSIS IS THEREFORE ONLY BASED ON THE RETURNED DEVICE DATA AS WELL AS THE INSPECTION OF THE QUALITY DOCUMENTS ASSOCIATED WITH THE MANUFACTURE OF THIS PARTICULAR DEVICE. THE MANUFACTURING PROCESS FOR THIS DEVICE WAS RE-INVESTIGATED. ALL PRODUCTION STEPS HAD BEEN PERFORMED ACCORDINGLY. THERE WAS NO SIGN OF ANY INCONSISTENCY DURING THE MANUFACTURING. IN A NEXT STEP, THE RETURNED DEVICE DATA WAS INSPECTED. DURING THE INSPECTION THE CLINICAL OBSERVATION COULD BE CONFIRMED. FURTHER ANALYSIS REVEALED AN UNEXPECTED BATTERY BEHAVIOR, WHICH LED TO THE CLINICAL OBSERVATION. BASED ON THE DATA AVAILABLE FOR ANALYSIS THE ROOT CAUSE FOR THAT BEHAVIOR WAS NOT DETERMINABLE AND CAN ONLY BE CLARIFIED BY AN ANALYSIS OF THE DEVICE ITSELF. HOWEVER, IT CANNOT BE EXCLUDED THAT THIS OCCURRENCE RESULTED FROM A DAMAGED COMPONENT, WHICH MAY COMPROMISE THE ABILITY OF THE DEVICE TO DELIVER THERAPY. THEREFORE AND IN ORDER TO EXCLUDE ANY POTENTIAL HARM FOR THE PATIENT WE WOULD LIKE TO RECOMMEND TO REPLACE THE DEVICE AND TO RETURN IT FOR ANALYSIS TO BIOTRONIK. OF COURSE, INDIVIDUAL CIRCUMSTANCES AND MEDICAL JUDGMENT DETERMINE DECISIONS ABOUT PATIENT CARE AND THE FREQUENCY OF FOLLOW-UP SESSIONS. THE ABOVE RECOMMENDATION IS BASED ON A PURELY TECHNICAL ASSESSMENT. SHOULD FURTHER RELEVANT INFORMATION OR THE DEVICE ITSELF BECOME AVAILABLE THIS INVESTIGATION WILL BE UPDATED. IT WAS REPORTED THAT THIS DEVICE WAS EXPLANTED ON (B)(6) 2020. PRIOR TO THE ANALYSIS OF THE DEVICE, THE QUALITY DOCUMENTS ACCOMPANYING THE MANUFACTURING PROCESS FOR THIS ICD WERE RE-INVESTIGATED. ALL PRODUCTION STEPS WERE PERFORMED ACCORDINGLY, AND IN PARTICULAR THE FINAL ACCEPTANCE TEST PROVED THE DEVICE FUNCTIONS TO BE AS SPECIFIED. THE DEVICE INTERROGATION REVEALED THE EOS BATTERY STATUS, DETECTED AFTER EXPLANTATION OF THE DEVICE ON (B)(6) 2020. THE DEVICE WAS IMPLANTED FOR APPROXIMATELY 73 MONTHS AND 31 CHARGING CYCLES WERE RECORDED IN THE DEVICES MEMORY. THE ICD WAS SUBJECTED TO AN ELECTRICAL ANALYSIS. THE CURRENT CONSUMPTION OF THE ICD WAS FOUND TO BE NORMAL AND AS EXPECTED. HOWEVER, AN INCONSISTENCY BETWEEN THE AMOUNT OF CHARGE TAKEN FROM THE BATTERY AND THE BATTERY VOLTAGE WAS NOTED. IN ORDER TO OBTAIN FURTHER INSIGHTS, THE ICD WAS OPENED AND THE INNER ASSEMBLY WAS INSPECTED. THE VISUAL INSPECTION SHOWED NO ANOMALIES. SUBSEQUENTLY THE CURRENT CONSUMPTION OF THE ELECTRONIC MODULE WAS VERIFIED BY DIRECT MEASUREMENT AND PROVED TO BE AS EXPECTED AND CONSISTENT WITH THE INTERROGATED ICD DATA. THERE WAS NO INDICATION OF A MALFUNCTION OF THE ELECTRONIC MODULE. THEREFORE, THE BATTERY WAS DISCONNECTED FROM THE ELECTRONIC MODULE AND SENT TO THE MANUFACTURER FOR FURTHER ANALYSIS. DURING ANALYSIS OF THE BATTERY LITHIUM PLATING WAS IDENTIFIED, WHICH LED TO AN ELEVATED INTERNAL SELF-DEPLETION AND, AS A RESULT, TO THE CLINICAL OBSERVATION. PLEASE NOTE, THIS ICD IS AFFECTED BY THE FIELD SAFETY CORRECTIVE ACTION, BIO-LQC, INITIATED IN MARCH 2021. IT WAS REPORTED THAT THIS DEVICE WAS EXPLANTED ON (B)(6) 2020. THE DEVICE WAS NOT RETURNED FOR ANALYSIS. THE ANALYSIS IS THEREFORE ONLY BASED ON THE RETURNED DEVICE DATA AS WELL AS THE INSPECTION OF THE QUALITY DOCUMENTS ASSOCIATED WITH THE MANUFACTURE OF THIS PARTICULAR DEVICE. THE MANUFACTURING PROCESS FOR THIS DEVICE WAS RE-INVESTIGATED. ALL PRODUCTION STEPS HAD BEEN PERFORMED ACCORDINGLY. THERE WAS NO SIGN OF ANY INCONSISTENCY DURING THE MANUFACTURING. IN A NEXT STEP, THE RETURNED DEVICE DATA WAS INSPECTED. DURING THE INSPECTION THE CLINICAL OBSERVATION COULD BE CONFIRMED. FURTHER ANALYSIS REVEALED AN UNEXPECTED BATTERY BEHAVIOR, WHICH LED TO THE CLINICAL OBSERVATION. BASED ON THE DATA AVAILABLE FOR ANALYSIS THE ROOT CAUSE FOR THAT BEHAVIOR WAS NOT DETERMINABLE AND CAN ONLY BE CLARIFIED BY AN ANALYSIS OF THE DEVICE ITSELF. HOWEVER, IT CANNOT BE EXCLUDED THAT THIS OCCURRENCE RESULTED FROM A DAMAGED COMPONENT, WHICH MAY COMPROMISE THE ABILITY OF THE DEVICE TO DELIVER THERAPY. THEREFORE AND IN ORDER TO EXCLUDE ANY POTENTIAL HARM FOR THE PATIENT WE WOULD LIKE TO RECOMMEND TO REPLACE THE DEVICE AND TO RETURN IT FOR ANALYSIS TO BIOTRONIK. OF COURSE, INDIVIDUAL CIRCUMSTANCES AND MEDICAL JUDGMENT DETERMINE DECISIONS ABOUT PATIENT CARE AND THE FREQUENCY OF FOLLOW-UP SESSIONS. THE ABOVE RECOMMENDATION IS BASED ON A PURELY TECHNICAL ASSESSMENT. SHOULD FURTHER RELEVANT INFORMATION OR THE DEVICE ITSELF BECOME AVAILABLE THIS INVESTIGATION WILL BE UPDATED. | ICD | BIOTRONIK SE & CO. KG | 383594 |  | LWS | 20150228 | 20210302 | Device was returned to manufacturer | HEALTH PROFESSIONAL | Implantable Cardioverter Defibrillator (Non-Crt)
Malfunction | 20200226 | ['No Display/Image', 'Battery Problem', 'Device Markings/Labelling Problem', 'Material Integrity Problem', 'Power Problem', 'Patient Device Interaction Problem'] | INFORMATION RECEIVED BY MEDTRONIC INDICATED THAT THE INSULIN PUMP HAD LOW BATTERY AND BLANK DISPLAY. THE CUSTOMER STATED WAS UNABLE TO GET THE INSULIN PUMP TURN ON ALSO CUSTOMER HAD THE BATTERY CAP CONTACT MISSING AND THE SN LABEL WORN OFF. NO HARM REQUIRING MEDICAL INTERVENTION WAS REPORTED. THE INSULIN PUMP WAS RETURNED FOR ANALYSIS. |  | UPDATED H9: 2032227-060322-002-C MEDTRONIC, INC. (MEDTRONIC) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY MEDTRONIC, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. MEDTRONIC HAS MADE REASONABLE EFFORTS TO OBTAIN MORE COMPLETE INFORMATION IN THE TIME ALLOTTED AND HAS PROVIDED AS MUCH INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE AN ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY "DEFECTS" OR HAS "MALFUNCTIONED". THESE WORDS ARE INCLUDED IN THE FDA 3500A FORM AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. MEDTRONIC OBJECTS TO THE USE OF THESE WORDS AND OTHERS LIKE IT BECAUSE OF THE LACK OF DEFINITION AND THE CONNOTATIONS IMPLIED BY THESE TERMS. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT. THIS REPORT IS PART OF A RETROSPECTIVE REVIEW AND REMEDIATION EFFORTS IN RESPONSE TO A WARNING LETTER. UPDATED H9: Z-0956-2020 MEDTRONIC, INC. (MEDTRONIC) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY MEDTRONIC, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. MEDTRONIC HAS MADE REASONABLE EFFORTS TO OBTAIN MORE COMPLETE INFORMATION IN THE TIME ALLOTTED AND HAS PROVIDED AS MUCH INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE AN ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY "DEFECTS" OR HAS "MALFUNCTIONED". THESE WORDS ARE INCLUDED IN THE FDA 3500A FORM AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. MEDTRONIC OBJECTS TO THE USE OF THESE WORDS AND OTHERS LIKE IT BECAUSE OF THE LACK OF DEFINITION AND THE CONNOTATIONS IMPLIED BY THESE TERMS. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT. UNIT PASSED SLEEP CURRENT MEASUREMENT, ACTIVE CURRENT MEASUREMENT, AND SELF TESTS. NO UNEXPECTED BLANK DISPLAYS OR UNEXPECTED ¿LOW BATTERY¿, ¿REPLACE BATTERY NOW¿, OR "POWER LOSS" ERRORS WERE NOTED DURING TESTING. UNIT RECEIVED WITH A PARTIALLY BROKEN OFF PIECE AT THE RESERVOIR TUBE LIP, A MISSING RETAINER RING, AND A MISSING RESERVOIR TUBE LIP O-RING. UNABLE TO PERFORM THE DISPLACEMENT TEST SINCE P-CAP / RESERVOIR WILL NOT LOCK PROPERLY DUE TO THE MISSING RETAINER. IN ADDITION TO THIS, THE S/N LABEL LOCATED BETWEEN THE BELT CLIP RAILS HAS RUBBED OFF AND BECOME ILLEGIBLE. (B)(4). | ARTIFICIAL PANCREAS DEVICE SYSTEM, SINGLE HORMONAL CONTROL | MEDTRONIC PUERTO RICO OPERATIONS CO. | MMT-1780KPK | HG29Y7N | OZP |  | 20200309 | Device was returned to manufacturer | LAY USER/PATIENT | Automated Insulin Dosing Device System, Single Hormonal Control
Injury | 20200422 | ['Battery Problem'] | FOLLOWING THE BATTERY PERFORMANCE ALERT (BPA) ADVISORY, A BPA WAS RECEIVED BY THE CLINICIAN AND THE DEVICE WAS EXPLANTED. THE PATIENT WAS IN STABLE CONDITION. |  | THE DEVICE IS INCLUDED IN THE BATTERY PERFORMANCE ALERT ADVISORY ISSUED BY ABBOTT ON (B)(6) 2017. THE RESULTS, METHOD, AND CONCLUSION CODES ALONG WITH INVESTIGATION RESULTS WILL BE PROVIDED IN THE FINAL REPORT. PREMATURE BATTERY DEPLETION WAS CONFIRMED BY ANALYSIS. NO SOURCES OF HIGH CURRENT WERE NOTED. THE CAUSE OF THE PREMATURE BATTERY DEPLETION WAS CONSISTENT WITH LITHIUM (LI) CLUSTER FORMATION. FROM THESE ANALYSES, IN THE ABSENCE OF HIGH CURRENT DRAW, IT IS PROBABLE THAT THE PREMATURE BATTERY DEPLETION WAS CAUSED BY A LITHIUM CLUSTER INDUCED SHORT CIRCUIT. LI CLUSTERS ARE A KNOWN DEPLETION MECHANISM FOR THESE ADVISORY PRODUCTS THAT HAS BEEN INVESTIGATED AND ASSOCIATED WITH A FIELD ACTION IN OCTOBER 2016. | IMPLANTABLE CARDIOVERTER DEFIBRILLATOR | ABBOTT | CD1233-40Q | 3122369 | LWS | 20111130 | 20200505 | Device was returned to manufacturer | HEALTH PROFESSIONAL | Implantable Cardioverter Defibrillator (Non-Crt)
Injury | 20200410 | ['Battery Problem', 'Insufficient Information'] | ADDITIONAL INFORMATION INDICATES THE IPG WAS EXPLANTED AND REPLACED TO ADDRESS THE ISSUE. | IT WAS REPORTED THE IPG HAS REACHED END OF SERVICE. IN TURN, SURGICAL INTERVENTION MAY BE PENDING. | THE RESULTS OF THE INVESTIGATION ARE INCONCLUSIVE SINCE THE DEVICE WAS NOT RETURNED FOR ANALYSIS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED. THE RESULTS/METHOD AND CONCLUSION CODES ALONG WITH INVESTIGATION RESULTS WILL BE PROVIDED IN THE FINAL REPORT. DATE OF EVENT IS ESTIMATED. DURING PROCESSING OF THIS COMPLAINT, ATTEMPTS WERE MADE TO OBTAIN COMPLETE PATIENT INFORMATION. FURTHER INFORMATION WAS REQUESTED BUT NOT RECEIVED. THE RESULTS OF THE INVESTIGATION ARE INCONCLUSIVE SINCE THE DEVICE WAS NOT RETURNED FOR ANALYSIS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED. | SCS IPG | ST. JUDE MEDICAL - NEUROMODULATION | 3662 | 5758833 | LGW | 20181213 |  | No | HEALTH PROFESSIONAL | Stimulator, Spinal-Cord, Totally Implanted For Pain Relief
Injury | 20200331 | ['Battery Problem'] | INFORMATION WAS RECEIVED FROM A CONSUMER (CON) REGARDING A PATIENT WHO WAS IMPLANTED WITH A NEUROSTIMULATOR (INS) FOR GASTRIC STIMULATION. IT WAS REPORTED THAT THE PATIENT THOUGHT THEIR INS WAS DEAD BECAUSE THEY HAD A LOT OF PAIN IN THEIR LOWER BACK. THEY STATED THEY WERE IN THE HOSPITAL WITH PAIN. NO FURTHER COMPLICATIONS WERE REPORTED OR ANTICIPATED. | ADDITIONAL INFORMATION WAS RECEIVED FROM THE PATIENT. THEY PROVIDED THEIR WEIGHT AT THE TIME OF THE EVENT. THEY STATED THERE WERE NO CIRCUMSTANCES LEADING TO THE DEAD INS AND PAIN. THEY WERE DOING NOTHING WHEN THE PAIN STARTED. IT WAS NOT RESOLVED AT THE TIME OF THE REPORT. NO PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. | IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. | INTESTINAL STIMULATOR | MEDTRONIC PUERTO RICO OPERATIONS CO. | 37800 |  | LNQ | 20181028 |  | No | HEALTH PROFESSIONAL | Intestinal Stimulator
Malfunction | 20200422 | ['Device Alarm System', 'Device Difficult to Program or Calibrate', 'Battery Problem'] | INFORMATION RECEIVED BY MEDTRONIC INDICATED THAT THE INSULIN PUMP HAD KEYPAD ANOMALY. CUSTOMER STATED THAT THE BUTTONS WERE NOT AS EASY TO PUSH AS THEY USE TO BE AND SOMETIMES I DIDN'T GOT A LOW RESERVOIR WARNING. IT WAS REPORTED THAT THE INSULIN PUMP WAS REJECTED SEVERAL BATTERY. NO HARM REQUIRING MEDICAL INTERVENTION WAS REPORTED. THE INSULIN PUMP WILL NOT BE RETURNED FOR ANALYSIS. |  | (B)(4). CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. | ARTIFICIAL PANCREAS DEVICE SYSTEM, THRESHOLD SUSPEND | MEDTRONIC PUERTO RICO OPERATIONS CO. | MMT-751NAS | A4751NASJ | OZO |  |  | No | LAY USER/PATIENT | Automated Insulin Dosing , Threshold Suspend
Malfunction | 20200422 | ['Device Difficult to Program or Calibrate', 'Battery Problem'] | INFORMATION RECEIVED BY MEDTRONIC INDICATED THAT THE INSULIN PUMP HAD BUTTON ERROR ALARM. NO HARM REQUIRING MEDICAL INTERVENTION WAS REPORTED. THE INSULIN PUMP WILL NOT BE RETURNED FOR ANALYSIS. |  | (B)(4). CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. | ARTIFICIAL PANCREAS DEVICE SYSTEM, THRESHOLD SUSPEND | MEDTRONIC PUERTO RICO OPERATIONS CO. | MMT-751LNAS | A4751LNASJ | OZO |  |  | No | LAY USER/PATIENT | Automated Insulin Dosing , Threshold Suspend
Malfunction | 20200404 | ['Battery Problem'] | IT WAS REPORTED THAT THE PUMP BATTERY WAS DEPLETING QUICKLY. THE CUSTOMER'S BLOOD GLUCOSE WAS 101 MG/DL. MULTIPLE ATTEMPTS WERE MADE TO FOLLOW UP WITH THE CUSTOMER REGARDING THE REPORTED ISSUE; HOWEVER, NO RESPONSE FROM THE CUSTOMER WAS RECEIVED. |  | NO PRODUCT WAS RETURNED FOR EVALUATION. SHOULD NEW RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. | CONTINUOUS GLUCOSE MONITOR | TANDEM DIABETES CARE | 1000096 |  | OYC |  |  | No | LAY USER/PATIENT | Pump, Infusion, Insulin, To Be Used With Invasive Glucose Sensor
Malfunction | 20200422 | ['Battery Problem', 'Patient Device Interaction Problem'] | INFORMATION RECEIVED BY MEDTRONIC INDICATED THAT INSULIN PUMP RETAINER RING WAS DAMAGED. CUSTOMER STATED THAT THE INSULIN PUMP HAD NOT HELD A CHARGE, NOW THEY HAD TO CHANGE THE BATTERY AFTER ONLY FOUR DAYS. NO HARM REQUIRING MEDICAL INTERVENTION WAS REPORTED. THE DEVICE WILL BE RETURNED FOR ANALYSIS. |  | (B)(4). CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. THE DEVICE WILL BE RETURNED FOR ANALYSIS AND FURTHER INFORMATION WILL FOLLOW ONCE THE ANALYSIS HAS BEEN COMPLETED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. DEVICE WAS RECEIVED WITH A SCRATCHED CASE. THE TEST RESERVOIR DOES LOCK PROPERLY INSIDE THE RESERVOIR TUBE. DEVICE PASSED THE DISPLACEMENT TEST, ACTIVE CURRENT MEASUREMENT AND SELF TEST. HOWEVER, THE DEVICE FAILED THE SLEEP CURRENT MEASUREMENT DUE TO A PROBLEM ISOLATED TO THE ELECTRONIC ASSEMBLY. NO DAMAGED RESERVOIR RETAINER RING WAS FOUND DURING VISUAL INSPECTION. | ARTIFICIAL PANCREAS DEVICE SYSTEM, THRESHOLD SUSPEND | MEDTRONIC PUERTO RICO OPERATIONS CO. | MMT-1715K | HG3ZRHT | OZO |  | 20200709 | Device was returned to manufacturer | LAY USER/PATIENT | Automated Insulin Dosing , Threshold Suspend
Malfunction | 20200201 | ['Battery Problem'] | IT WAS REPORTED THAT THE PUMP BATTERY INTERMITTENTLY DEPLETED QUICKLY. THERE WAS NO REPORTED IMPACT TO THE CUSTOMER¿S BLOOD GLUCOSE LEVEL. THE CUSTOMER CONTINUED TO USE THE PUMP FOR INSULIN THERAPY. |  | NO PRODUCT WAS RETURNED FOR EVALUATION. SHOULD NEW RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. | AUTOMATED INSULIN DOSING, THRESHOLD SUSPEND | TANDEM DIABETES CARE | 1000096 |  | OZO |  |  | No | LAY USER/PATIENT | Automated Insulin Dosing , Threshold Suspend
Malfunction | 20200101 | ['Battery Problem'] | IT WAS REPORTED THAT THE PUMP BATTERY INTERMITTENTLY DEPLETED QUICKLY. THE CUSTOMER¿S BLOOD GLUCOSE RANGED FROM AROUND 100-200 MG/DL. THE CUSTOMER REVERTED TO AN ALTERNATE METHOD OF INSULIN THERAPY. |  | THE DEVICE IS EXPECTED TO BE RETURNED; HOWEVER, THE DEVICE HAS NOT YET BEEN RECEIVED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED IF THE DEVICE IS RECEIVED. | INSULIN PUMP | TANDEM DIABETES CARE | 4628-003 |  | OYC |  | 20200428 | Device was returned to manufacturer | LAY USER/PATIENT | Pump, Infusion, Insulin, To Be Used With Invasive Glucose Sensor
Injury |  | ['Battery Problem'] | INFORMATION WAS RECEIVED FROM A CONSUMER CONCERNING PATIENT WITH AN IMPLANTABLE NEUROSTIMULATOR (INS). IT WAS REPORTED THAT THE PATIENT¿S PAIN STIMULATION DEVICE WAS DEAD. THE PATIENT STATED THAT HER DOCTOR WAS PLANNING TO IMPLANT A NEW PAIN STIMULATION DEVICE SOON. NO PATIENT SYMPTOMS WERE REPORTED. NO FURTHER COMPLICATIONS WERE REPORTED/ANTICIPATED. |  | IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. | STIMULATOR, SPINAL-CORD, TOTALLY IMPLANTED FOR PAIN RELIEF | MEDTRONIC MED REL MEDTRONIC PUERTO RICO | 97712 |  | LGW | 20141028 |  | No | HEALTH PROFESSIONAL | Stimulator, Spinal-Cord, Totally Implanted For Pain Relief
Injury | 20200201 | ['Battery Problem', 'Electromagnetic Compatibility Problem', 'Insufficient Information'] | INFORMATION WAS RECEIVED FROM HEALTHCARE PROVIDER (HCP) AND A MANUFACTURER REPRESENTATIVE (REP) REGARDING A PATIENT WITH AN IMPLANTABLE NEUROSTIMULATOR (INS) FOR URINARY DYSFUNCTION/SACRAL NERVE STIMULATION AND GASTROINTESTINAL/PELVIC FLOOR. A POWER ON RESET (POR) WAS REPORTED. THE REP WAS WITH THE PATIENT THAT DAY TO ADDRESS IT. THE CALLER NOTED GETTING ¿INVALID IMPLANTABLE NEUROSTIMULATOR¿ ON THE CLINICIAN PROGRAMMER (8840) AFTER INTERROGATING THE INS POST POR AND THE 8840 ASKED THE CALLER TO ENTER THE SERIAL NUMBER. DURING THE CALL, THE CALLER ENTERED THE SERIAL NUMBER MULTIPLE TIMES (WITHOUT THE PREFIX NJY, AS INTENDED,) AND THE POR CONTINUED TO OCCUR. TECHNICAL SERVICES INQUIRED WITH THE CALLER AND THE CALLER STATED THAT THE PATIENT HAD BEEN RUNNING STIMULATION AT 7-8V IN THE PAST AND THE PATIENT HAD CONFIRMED SEEING THE ¿LOW INS¿ SYMBOL ON THE ICON PATIENT PROGRAMMER. TECHNICAL SERVICES ADVISED THAT THE PERSISTENT POR OCCURRING THAT DAY OF THE CALL WAS LIKELY RELATED TO A LOW INS BATTERY. THE CALLER DID NOTE SEEING ¿UNCONFIRMED PARAMETER VALUES,¿ AND ¿INVALID PARAMETERS,¿ ON THE 8840 AS WELL DURING THIS PROCESS HOWEVER TECHNICAL SERVICES WAS UNABLE TO FIND ANY DOCUMENTATION OF THAT MESSAGE IN THE 8840 OR TECHNICAL SERVICES RESOURCE GUIDE. THE CALLER ALSO SAW ¿INTERFERENCE¿ ON THE 8840 DURING THIS PROCESS AND THE MESSAGE CAME UP MULTIPLE TIMES WHEN THE CLINIC LIGHTS WERE ON/OFF AND WHEN THE PATIENT M OVED AWAY FROM A COMPUTER. THE POR WAS NOTED TO HAVE OCCURRED IN (B)(6) 2020 BUT ALL 8840 MESSAGES HAPPENED DURING THE CALL ON (B)(6) 2020. IT WAS NOTED THAT BOTH THE CALLER (WHO HAD THE HEALTH CARE PHYSICIAN (HCP) ON THE PHONE WITH THEM) AND THE HCP WERE UNCLEAR IF A FALL ON (B)(6) 2019, WHERE THE PATIENT FELL ON THEIR SIDE AND THEN ON THEIR BACK, HAD ANY EFFECT ON THE PATIENT¿S SYSTEM. THERE WERE NO SYMPTOMS AND THERE WERE NO FURTHER COMPLICATIONS REPORTED. ADDITIONAL INFORMATION WAS RECEIVED FROM A HEALTH CARE PHYSICIAN (HCP) VIA A MANUFACTURER REPRESENTATIVE (REP). IN RESPONSE TO THE INQUIRY FOR CLARIFICATION ON IF THE REPORT OF THE POWER ON RESET (POR) MESSAGE AND THE LOW BATTERY SYMBOL ON THE ICON WAS THE RESULT OF NORMAL BATTERY DEPLETION DUE TO THE SETTINGS THE PATIENT HAD BEEN USING THE REP REPLIED THAT ¿YES, IT WAS DETERMINED THAT IT WAS DUE TO HIGHER AMPLITUDES DEPLETING THE INS CAUSING THE POR. THERE WERE NO FURTHER COMPLICATIONS REPORTED. ADDITIONAL INFORMATION WAS RECEIVED FROM THE MANUFACTURER REPRESENTATIVE. THE REP STATED THAT THE PATIENT NEEDS REPLACEMENT HOWEVER DUE TO THE PANDEMIC OF COVID-19, THE REP IS UNABLE TO PROVIDE ADDITIONAL INFORMATION. NO FURTHER COMPLICATIONS WERE REPORTED. |  | DATE IS APPROXIMATE. IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. | STIMULATOR, ELECTRICAL, IMPLANTABLE, FOR INCONTINENCE | MEDTRONIC PUERTO RICO OPERATIONS CO. | 3058 |  | EZW | 20190728 |  | No | HEALTH PROFESSIONAL | Stimulator, Electrical, Implantable, For Incontinence
Malfunction | 20200122 | ['Battery Problem'] | THE BATTERY WAS RETURNED TO THE MANUFACTURER BECAUSE IT WAS NOT CHARGING AND WAS FLASHING RED. THE DEVICE SUBSEQUENTLY TESTED OUT OF SPECIFICATION DUE TO MANUFACTURER¿S ANALYSIS. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. |  | THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANSPLANTATION AND DESTINATION THERAPY IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY END-STAGE LEFT VENTRICULAR HEART FAILURE. THE SYSTEM IS DESIGNED FOR IN-HOSPITAL AND OUT-OF-HOSPITAL SETTINGS. THE BATTERY (BAT217222) WAS RETURNED FOR EVALUATION. VARIOUS ANALYSES WERE CONDUCTED AND REVIEWED IN ORDER TO EVALUATE THE PERFORMANCE OF THE DEVICE IN RELATION TO THE REPORTED EVENT. FAILURE ANALYSIS OF THE RETURNED BATTERY REVEALED THAT THE UNIT PASSED VISUAL EXAMINATION. FUNCTIONAL TESTING REVEALED THAT THE BATTERY FLASHED RED ON THE BATTERY CHARGER DUE TO A HIGH MAX ERROR. THE MAX ERROR IS AN INDICATOR OF HOW WELL THE BATTERY¿S RELATIVE STATE OF CHARGE (RSOC) IS CALIBRATED. THE BATTERY IS STILL ABLE TO CHARGE AND PROVIDE POWER TO A CONTROLLER. A TENDENCY TO EXCHANGE BATTERIES BEFORE REACHING 25% RELATIVE STATE OF CHARGE (RSOC) MAY CONTRIBUTE TO AN OUT-OF-CALIBRATION CONDITION. A HIGH MAX ERROR WILL CAUSE THE BATTERY TO FLASH RED WHEN CONNECTED TO A BATTERY CHARGER. ADDITIONALLY, THE HIGH MAX ERROR WAS UNABLE TO BE RESET DURING FUNCTIONAL TESTING, LIKELY DUE TO A PROBLEM WITH THE MAIN INTEGRATED CIRCUIT (IC). AS A RESULT, THE REPORTED FLASHING RED EVENT WAS CONFIRMED. THE REPORTED BATTERY NOT CHARGING EVENT COULD NOT BE CONFIRMED; HOWEVER, THE HIGH MAX ERROR MIGHT BE WHAT THE PATIENT EXPERIENCED DURING THE REPORTED EVENT. THE MOST LIKELY ROOT CAUSE OF THE REPORTED EVENT CAN BE ATTRIBUTED TO A BATTERY THAT IS OUT OF CALIBRATION. THE MOST LIKELY ROOT CAUSE OF THE MAX ERROR UNABLE TO BE RESET DURING TESTING CAN BE ATTRIBUTED TO A FAULTY INTEGRATED CIRCUIT. AN INTERNAL INVESTIGATION EVALUATED BATTERY MAIN INTEGRATED CIRCUIT MALFUNCTIONS. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. | VENTRICULAR (ASSIST) BYPASS | HEARTWARE, INC. | 1650DE |  | DSQ | 20170430 | 20200403 | Device was returned to manufacturer | LAY USER/PATIENT | Ventricular (Assist) Bypass
Injury | 20200424 | ['Migration or Expulsion of Device', 'Energy Output Problem', 'Battery Problem', 'Charging Problem', 'Communication or Transmission Problem'] | ADDITIONAL INFORMATION WAS RECEIVED FROM A MANUFACTURER REPRESENTATIVE (REP) REPORTING THEY ARE WITH THE PATIENT NOW AND THE REPLACEMENT DESKTOP CHARGER (DTC) WILL NOT FIT INTO THE RECHARGER. THE PORT IS LOOSE AND THE PATIENT'S IMPLANTABLE NEUROSTIMULATOR RECHARGER (INSR) WILL NOT CONNECT. THEY USED THEIR SPARE RECHARGE AND CONNECTED TO THE INS RIGHT AWAY WITHOUT ISSUE. A REPLACEMENT RECHARGER AND DTC WAS SENT. | THE HEALTH CARE PROVIDER (HCP) WAS CALLING FROM A NURSING HOME, PATIENT HAD A SUCCESSFUL CHARGE TWO DAYS AGO (B)(6) 2020. IT WAS REPORTED THAT THE PATIENT STARTED HAVING BILATERAL ARM TREMOR WHICH STARTED SUDDENLY LAST NIGHT. THE CALLER WAS ATTEMPTING TO CHARGE PATIENT'S DEVICE AND SHE IS SEEING REPOSITION ANTENNA-PRESS THERAPY OFF ICON. IT WAS REPORTED THAT THE PATIENT PROGRAMMER SHOWS RECHARGE STIMULATOR. ANTENNA LOCATOR WAS PERFORMED, RANGING FROM 52-54. CALLER REPORTS THAT THEY CONTINUE TO SEE REPOSITION ANTENNA ICON. THEY ATTEMPTED MULTIPLE TIMES WITH THE RECHARGE ANTENNA AT DIFFERENT LOCATIONS, BUT CONTINUES TO SEE REPOSITION ANTENNA ICON. IT WAS REPORTED THAT THEY WOULD ADMIT THE PATIENT TO HOSPITAL AS THE TREMORS WERE GETTING WORSE ADDITIONAL INFORMATION WAS RECEIVED FROM A MANUFACTURER REPRESENTATIVE (REP) ON (B)(6) 2020. IT WAS REPORTED THE CAUSE OF THE RECHARGING ISSUE AND SUDDEN TREMORS WAS, BASED OFF THE RECORD PULLED FROM THE DEVICE, THERE WAS NO CHARGING SESSION FROM (B)(6) TO (B)(6), SO THE BATTERY LEVEL DROPPED TO 25%. THERE WERE SHORT SESSIONS IN THE RECORD ON (B)(6)AND THEN NONE AFTER THAT. THEIR DEVICE DIED AND THEIR THERAPY TURNED OFF SOMETIME AROUND (B)(6). STAFF REPORTED THE PATIENT WAS CHARGING THEMSELVES AND NO ONE WAS HELPING THEM AT THEIR NURSING HOME. THE PATIENT HAS VERY ADVANCED PARKINSON'S AND DEMENTIA. THE REP SPOKE WITH THE MANAGING PHYSICIAN ABOUT HAVING THE STAFF CHARGE THE PATIENT DAILY SINCE THEY ARE NOT ABLE TO KEEP UP WITH IT THEMSELVES. THEY HAVE UPDATED THEIR ORDERS FOR THE STAFF TO CHARGE THE PATIENT DAILY, BUT THE STAFF IS STILL HAVING PROBLEMS GETTING THE RECHARGER TO CONNECT. TROUBLESHOOTING WAS PERFORMED OVER PHONE AND FACETIME AND HAVE SPOKEN TO TECHNICAL SUPPORT MULTIPLE TIMES; HOWEVER, NO ONE CAN FIGURE OUT WHY THE STAFF IS UNABLE TO GET THE RECHARGER TO COUPLE. THE REP HAD NO ISSUE RECHARGING THE PATIENT ON (B)(6). DUE TO COVID-19, THE REP IS NOT ALLOWED INTO THE FACILITY. THE ONLY SUDDEN TREMORS WERE FROM THE ORIGINAL ISSUE WHEN THEIR DEVICE DEPLETED ON (B)(6)TO (B)(6). THE THERAPY TURNED OFF BECAUSE THEIR DEVICE WAS DEPLETED. THERE HAS BEEN NO TREMOR ISSUES SINCE THEY CHARGED THE PATIENT BACK TO 100% ON THE 28TH DUE TO THEIR THERAPY BEING ON AND HAVE BEEN ABLE TO CHARGE THEIR DEVICE AT LEAST 2 TIMES SINCE THE (B)(6). THE RECHARGING ISSUES HAVE CONTINUED BECAUSE THE STAFF CANNOT GET THE RECHARGER TO CONNECT TO THE BATTERY CONSISTENTLY. ADDITIONAL INFORMATION WAS RECEIVED BY THE MANUFACTURING REPRESENTATIVE (REP). IT WAS REPORTED THAT THE REP MET THE PATIENT AT THE HOSPITAL FOR A POSSIBLE REVISION DUE TO THE INS NOT TAKING A CHARGE AND IT WAS REPORTED THE INS WAS DEAD. THE REP WAS ABLE TO CHARGE THE DEVICE WITH NO ISSUE WITH ALL COUPLING BARS. THE DEVICE WAS CHARGED TO 75-100% AND THERAPY WAS ON. EARLIER IN THE MONTH WHEN THE REP INTERROGATED THE DEVICE, THERE WAS NO CHARGE ON THE INS AND THE BATTERY DROPPED TO 25%. OVER THE WEEKEND, TWO NURSES WERE ABLE TO CHARGE AND TWO NURSES WERE NOT ABLE TO. THE PATIENT HAD THERAPY, THEY JUST WEREN'T ABLE TO CHARGE. THE CALLER SAID SHE CONTINUED TO SEE REPOSITION ANTENNA ICON ON THE RECHARGER. SOMEONE FROM THE NURSING HOME CALLED AND STATED THE RECHARGER HAD S SW 4.1.00. THE CALLER REPORTED THE INS CAN MOVE AROUND IN THE POCKET, SIDE TO SIDE. THEY INDICATED THE INS COULD NOT BE FLIPPED. ANTENNA LOCATION (AL) WAS PERFORMED AND SHOWED 64/66 AND WHEN ITS AWAY FROM THE INS, AL WAS 48. WHEN THE CALLER PUTS THE RECHARGER DIRECTLY ON THE INS, SHE CONTINUES TO SEE REPOSITION ANTENNA. THE AGENT CALLED THE REP BACK AND SUGGESTED POSSIBLE X-RAY. NO SYMPTOMS WERE REPORTED. | IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. MEDTRONIC IS SUBMITTING THIS REPORT TO COMPLY WITH FDA REPORTING REGULATIONS UNDER 21 CFR PARTS 4 AND 803. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY MEDTRONIC, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. MEDTRONIC HAS MADE REASONABLE EFFORTS TO OBTAIN MORE COMPLETE INFORMATION AND HAS PROVIDED AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE AN ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY ¿DEFECTS¿ OR HAS ¿MALFUNCTIONED¿. THESE WORDS ARE INCLUDED IN THE FDA 3500A FORM AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REGULATORY REPORTING. MEDTRONIC OBJECTS TO THE USE OF THESE WORDS AND OTHERS LIKE THEM BECAUSE OF THE LACK OF DEFINITION AND THE CONNOTATIONS IMPLIED BY THESE TERMS. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. A GOOD FAITH EFFORT WILL BE MADE TO OBTAIN THE APPLICABLE INFORMATION RELEVANT TO THE REPORT. IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. PRODUCT ID 37754, SERIAL# (B)(4). PRODUCT TYPE RECHARGER. IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. | STIMULATOR, ELECTRICAL, IMPLANTED, FOR PARKINSONIAN TREMOR | MEDTRONIC MED REL MEDTRONIC PUERTO RICO | 37612 |  | MHY | 20140728 |  | No | HEALTH PROFESSIONAL | Stimulator, Electrical, Implanted, For Parkinsonian Tremor
Injury | 20200407 | ['Battery Problem'] | FOLLOWING THE BATTERY PERFORMANCE ALERT (BPA) ADVISORY, A BPA WAS RECEIVED BY THE CLINICIAN AND THE DEVICE WAS EXPLANTED. PATIENT WAS STABLE POST-PROCEDURE. |  | THE DEVICE IS INCLUDED IN THE BATTERY PERFORMANCE ALERT ADVISORY ISSUED BY ABBOTT ON 28 AUGUST 2017. THE RESULTS OF THE INVESTIGATION ARE INCONCLUSIVE SINCE THE DEVICE WAS NOT RETURNED FOR ANALYSIS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED. PREMATURE BATTERY DEPLETION WAS CONFIRMED BY ANALYSIS. NO SOURCES OF HIGH CURRENT WERE NOTED. THE CAUSE OF THE PREMATURE BATTERY DEPLETION WAS CONSISTENT WITH LI CLUSTER FORMATION. FROM THESE ANALYSES, IN THE ABSENCE OF HIGH CURRENT DRAW, IT IS PROBABLE THAT THE PREMATURE BATTERY DEPLETION WAS CAUSED BY A LITHIUM CLUSTER INDUCED SHORT CIRCUIT. LI CLUSTERS ARE A KNOWN DEPLETION MECHANISM FOR THESE ADVISORY PRODUCTS THAT HAS BEEN INVESTIGATED AND ASSOCIATED WITH A FIELD ACTION IN (B)(6) 2016. | IMPLANTABLE CARDIOVERTER DEFIBRILLATOR | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | CD3365-40C | A000002613 | NIK | 20170228 | 20200514 | Device was returned to manufacturer | HEALTH PROFESSIONAL | Defibrillator, Automatic Implantable Cardioverter, With Cardiac Resynchronization (Crt-D)
Malfunction | 20200422 | ['Obstruction of Flow', 'Battery Problem', 'Patient Device Interaction Problem', 'Priming Problem'] | INFORMATION RECEIVED BY MEDTRONIC INDICATED THAT THE INSULIN PUMP SHOOTED OUT INSULIN WHILE PRIMING. CUSTOMER ALSO REPORTED CRACK ON RESERVOIR WINDOW, NO DELIVERY ALARMS AND BATTERY LIFE HAD DIMINISHED. NO HARM REQUIRING MEDICAL INTERVENTION WAS REPORTED. THE DEVICE WILL NOT BE RETURNED FOR ANALYSIS. |  | (B)(4). CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. | ARTIFICIAL PANCREAS DEVICE SYSTEM, THRESHOLD SUSPEND | MEDTRONIC PUERTO RICO OPERATIONS CO. | MMT-551NAS | A4551NASJ | OZO |  |  | No | LAY USER/PATIENT | Automated Insulin Dosing , Threshold Suspend
Malfunction | 20200422 | ['Battery Problem', 'Power Problem'] | INFORMATION RECEIVED BY MEDTRONIC INDICATED THAT THE THEY DID NOT RECEIVED A LOW BATTERY ALERT PRIOR TO THE REPLACE BATTERY ALERT. NO HARM REQUIRING MEDICAL INTERVENTION WAS REPORTED. THE INSULIN PUMP WILL NOT BE RETURNED FOR ANALYSIS. |  | (B)(4). CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. THE INSULIN PUMP INVOLVED IN THIS EVENT IS THE 640G INSULIN INFUSION PUMP, WHICH IS NOT MARKETED IN THE UNITED STATES. HOWEVER, THE DEVICE IS SIMILAR TO THE PARADIGM REAL-TIME INSULIN INFUSION PUMP, WHICH IS MARKETED IN THE UNITED STATES. | PUMP, INFUSION, INSULIN, TO BE USED WITH INVASIVE GLUCOSE SENSOR | MEDTRONIC PUERTO RICO OPERATIONS CO. | MMT-1712K | HG1QQQA | OYC |  |  | No | LAY USER/PATIENT | Pump, Infusion, Insulin, To Be Used With Invasive Glucose Sensor
Malfunction | 20200404 | ['Loss of Power', 'Battery Problem', 'Electrical Power Problem'] | RELATED MANUFACTURER REFERENCE NUMBER: (B)(4). THE BACKUP BATTERY WAS EXCHANGED. IF THE MESSAGE TO CHANGE THE BACKUP BATTERY REAPPEARED THEN THE SITE WOULD CHANGE THE CONTROLLER NEXT. | IT WAS REPORTED THAT THE PATIENT'S PUMP STOPPED FOR A SHORT PERIOD. THERE WAS ALSO A MESSAGE FOR THE BACKUP BATTERY NEEDING TO BE REPLACED BUT THE BACKUP BATTERY WAS REPLACED A FEW MONTHS AGO FOR THE SAME PROBLEM. LOG FILE ANALYSIS SHOWED NUMEROUS LOW VOLTAGE ADVISORY EVENTS FROM NORMAL BATTERY DEPLETION AND THAT THE PATIENT APPEARS TO RUN THE BATTERIES PRETTY LOW BEFORE CHANGING TO FULLY CHARGED BATTERIES. INTERMITTENT BACKUP BATTERY FAULTS WERE NOTED STARTING ON (B)(6) 2020 AND MORE BACKUP BATTERY FAULTS WERE NOTED ON (B)(6) 2020. THERE WERE ALSO SOME BACKUP BATTERY FAULTS NOTED AND IT APPEARED THAT THE PUMP STOP BUTTON WAS PRESSED ON THE SYSTEM MONITOR, BRIEFLY STOPPING THE PUMP FOR APPROXIMATELY 3 SECONDS AND THEN THE PUMP RESUMED OPERATION. | MANUFACTURER'S INVESTIGATION CONCLUSION: THE REPORTED EVENT OF BACKUP BATTERY FAULT ALARMS WAS CONFIRMED VIA THE LOG FILE. THE SYSTEM CONTROLLER WAS NOT RETURNED FOR ANALYSIS; HOWEVER, THE 11V BACKUP BATTERY (SERIAL #: (B)(6)) WAS RETURNED FOR ANALYSIS AND A LOG FILE WAS SUBMITTED FOR REVIEW WITH EVENTS SPANNING APPROXIMATELY 11 DAYS ((B)(6) 2020 ¿ (B)(6) 2020 PER TIME STAMP). BACKUP BATTERY FAULT ALARMS WERE ACTIVE DUE TO A LOAD TEST FAILURE. THE BACKUP BATTERY FAULT ALARMS WERE INTERMITTENTLY ACTIVE ON (B)(6) 2020 BETWEEN 06:08:48 ¿ 10:11:44, (B)(6) 2020 BETWEEN 06:14:40 ¿ 21:24:02, AND ON (B)(6) 2020 BETWEEN 01:21:14 ¿ 07:54:17. THE BACKUP BATTERY FAILED THE LOAD TEST DUE TO THE LOADED VOLTAGE BEING UNDER THE ACCEPTABLE THRESHOLD AS COMPARE TO THE UNLOADED VOLTAGE. THE BACKUP BATTERY FAULT ALARMS DID NOT AFFECT PUMP OPERATION. THE 11V BACKUP BATTERY WAS TESTED AND PERFORMED AS INTENDED. THE 11V BACKUP BATTERY RETURNED WITH A MEASURED VOLTAGE OF 11.86V. THE REPORTED BACKUP BATTERY FAULT ALARMS WERE NOT ABLE TO BE REPRODUCED DURING THIS INVESTIGATION. THE ROOT CAUSE FOR THE REPORTED EVENT WAS UNABLE TO BE CONCLUSIVELY DETERMINED THROUGH THIS ANALYSIS. HEARTMATE III INSTRUCTIONS FOR USE SECTION 7 ENTITLED ¿ALARMS AND TROUBLESHOOTING¿ AND HEARTMATE III PATIENT HANDBOOK SECTION 5 ENTITLED ¿ALARMS AND TROUBLESHOOTING¿ ADDRESSES HOW TO PROPERLY INTERPRET AND TROUBLESHOOT ALL SYSTEM ALARMS. HEARTMATE III INSTRUCTIONS FOR USE SECTION 8 ENTITLED ¿EQUIPMENT STORAGE AND CARE¿ AND HEARTMATE III PATIENT HANDBOOK SECTION 6 ENTITLED ¿CARING FOR THE EQUIPMENT¿ ADDRESSES HOW TO PROPERLY CARE FOR, MAINTAIN, AND STORE THE EQUIPMENT FOR PROPER USE. THE PATIENT HANDBOOK CAUTIONS THE USERS TO CALL THEIR HOSPITAL CONTACTS IF THEY THINK, FOR ANY REASON, ANY PORTION OF THEIR EQUIPMENT IS NOT FUNCTIONING AS USUAL, IS BROKEN, OR THEY ARE UNCOMFORTABLE WITH THE OPERATION OF THE EQUIPMENT. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER¿S INVESTIGATION IS COMPLETED. | VENTRICULAR (ASSISST) BYPASS | THORATEC CORPORATION | 106531US | 6566340 | DSQ | 20210710 | 20200416 | Device was returned to manufacturer | LAY USER/PATIENT | Ventricular (Assist) Bypass
Injury | 20200326 | ['Improper or Incorrect Procedure or Method', 'Battery Problem', 'Charging Problem', 'Insufficient Information'] | INFORMATION WAS RECEIVED FROM A CONSUMER REGARDING A PATIENT WITH AN IMPLANTABLE NEUROSTIMULATOR (INS). IT WAS REPORTED THAT THE PATIENT¿S INS WAS NOT CHARGING AND THEY WERE GETTING ¿A SCREEN¿ ON THEIR RECHARGER WHEN THEY TRIED TO CHARGE. THE PATIENT DESCRIBED SEEING THE REPOSITION ANTENNA SCREEN, WHICH THEY FIRST SAW ABOUT THREE WEEKS AGO. PATIENT SERVICES HAD THE PATIENT TRY TO CONNECT USING THEIR PATIENT PROGRAMMER DURING THE CALL AND THE PATIENT DESCRIBED SEEING THE POOR COMMUNICATION SCREEN. PATIENT SERVICES ASKED THE PATIENT WHEN THEY LAST RECHARGED THEIR INS AND THE PATIENT STATED BACK IN FEBRUARY. PATIENT SERVICES REVIEWED THAT THE PATIENT¿S INS WAS POSSIBLY OVERDISCHARGED AND THEY WERE REDIRECTED TO FOLLOW-UP WITH THEIR HEALTHCARE PROVIDER (HCP) TO CHECK THEIR INS. AN EMAIL WAS ALSO SENT OUT TO A LOCAL MANUFACTURER REPRESENTATIVE (REP) WHO RESPONDED INDICATING THEY¿D REACH OUT TO THE PATIENT. THE EVENT OCCURRED IN (B)(6) 2020. NO SYMPTOMS WERE REPORTED. NO FURTHER COMPLICATIONS WERE REPORTED/ANTICIPATED. ADDITIONAL INFORMATION WAS RECEIVED FROM THE REP. IT WAS REPORTED THAT THE DEVICE WAS OVERDISCHARGED AND THE REP ATTEMPTED A JUMPSTART. REP REPORTED THAT (B)(6), THE REP SENT PATIENT TO HCP TO GET REPLACED. THE CAUSE WAS DUE TO A DEAD BATTERY. |  | IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. | STIMULATOR, SPINAL-CORD, TOTALLY IMPLANTED FOR PAIN RELIEF | MEDTRONIC MED REL MEDTRONIC PUERTO RICO | 37714 |  | LGW | 20140128 |  | No | HEALTH PROFESSIONAL | Stimulator, Spinal-Cord, Totally Implanted For Pain Relief
Malfunction | 20200422 | ['Mechanical Problem', 'Device Difficult to Program or Calibrate', 'Battery Problem', 'Patient Device Interaction Problem'] | INFORMATION RECEIVED BY MEDTRONIC INDICATED THAT THE INSULIN PUMP STOPPED IN MIDDLE DURING THE REWIND. CUSTOMER ALSO STATED THAT RESERVOIR HAD CRACKED. INSULIN PUMP'S SCREEN HAD SCRATCHES. NO HARM REQUIRING MEDICAL INTERVENTION WAS REPORTED. THE INSULIN PUMP WILL NOT BE RETURNED FOR ANALYSIS. |  | CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. | PUMP, INFUSION, INSULIN, TO BE USED WITH INVASIVE GLUCOSE SENSOR | MEDTRONIC MINIMED | MMT-XXX |  | OYC |  |  | No | LAY USER/PATIENT | Pump, Infusion, Insulin, To Be Used With Invasive Glucose Sensor
Malfunction | 20200422 | ['Mechanical Problem', 'Device Difficult to Program or Calibrate', 'Battery Problem', 'Power Problem', 'Patient Device Interaction Problem'] | INFORMATION RECEIVED BY MEDTRONIC INDICATED THAT THE INSULIN PUMP'S BUTTONS WERE HARDER TO PRESS, UNEXPLAINED NO DELIVERY ALARM, REWOUND PROCESS WERE SLOWER AND FAILED BATTERY TEST. NO HARM REQUIRING MEDICAL INTERVENTION WAS REPORTED. THE INSULIN PUMP WILL NOT BE RETURNED FOR ANALYSIS. |  | DEVICE MONITORED FOR SEVERAL DAYS. DEVICE RECEIVED WITH INTERMITTENT BUTTON RESPONSE DUE TO FLATTENED ESC, ACT BUTTON DOME SWITCH (CREASED). NO BUTTON ERROR ALARM DURING TESTING. NO UNLOCKED J2/LCD KEYPAD CONNECTOR. DEVICE RECEIVED WITH ALL OPERATING CURRENTS WITHIN SPEC AND PASSED A21 ERROR TEST, REWIND AND DISPLACEMENT TEST. NO UNEXPECTED REWIND ANOMALIES NOTED. NO UNEXPECTED LOW BATTERY ALARM ANOMALIES NOTED. NO UNEXPECTED FAILED BATTERY ALARM ANOMALIES NOTED. (B)(4). (B)(4). CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. DEVICE MONITORED FOR SEVERAL DAYS. DEVICE RECEIVED WITH INTERMITTENT BUTTON RESPONSE DUE TO FLATTENED ESC, ACT BUTTON DOME SWITCH (CREASED). NO BUTTON ERROR ALARM DURING TESTING. NO UNLOCKED LCD KEYPAD CONNECTOR. DEVICE RECEIVED WITH ALL OPERATING CURRENTS WITHIN SPEC AND PASSED A21 ERROR TEST, REWIND AND DISPLACEMENT TEST. NO UNEXPECTED REWIND ANOMALIES NOTED. NO UNEXPECTED LOW BATTERY ALARM ANOMALIES NOTED. NO UNEXPECTED FAILED BATTERY ALARM ANOMALIES NOTED. (B)(4). | ARTIFICIAL PANCREAS DEVICE SYSTEM, THRESHOLD SUSPEND | MEDTRONIC PUERTO RICO OPERATIONS CO. | MMT-751NAS | A4751NASJ | OZO |  | 20200521 | Device was returned to manufacturer | LAY USER/PATIENT | Automated Insulin Dosing , Threshold Suspend
Injury | 20200404 | ['Battery Problem'] | IT WAS REPORTED THAT THE PUMP BATTERY WAS DEPLETING QUICKLY. ADDITIONALLY, THE CUSTOMER EXPERIENCED AN ELEVATED BLOOD GLUCOSE (BG) LEVEL (VALUE NOT PROVIDED). CUSTOMER DECLINED TROUBLESHOOTING WITH TANDEM TECHNICAL SUPPORT AND DECLINED TO PROVIDE FURTHER INFORMATION. |  | NO PRODUCT WAS RETURNED FOR EVALUATION. SHOULD NEW RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. | AUTOMATED INSULIN DOSING, THRESHOLD SUSPEND | TANDEM DIABETES CARE | 1002717 |  | OZO |  |  | No | LAY USER/PATIENT | Automated Insulin Dosing , Threshold Suspend
Injury | 20190725 | ['Delayed Charge Time', 'Battery Problem', 'Charging Problem', 'Communication or Transmission Problem', 'Electromagnetic Compatibility Problem', 'Insufficient Information'] | ADDITIONAL INFORMATION RECEIVED FROM THE MANUFACTURER REPRESENTATIVE REPORTED THAT THE REASON THE DEVICES WERE SO CLOSE TOGETHER WAS THAT THE PATIENT ORIGINALLY HAD TWO IMPLANTS THAT WERE REPLACED WITH THE CURRENT DEVICES. WHEN THAT REPLACEMENT OCCURRED THE NEW DEVICES WENT INTO THE SAME POCKET AS THE ONES THAT WERE REMOVED. | INFORMATION WAS RECEIVED FROM A MANUFACTURER REPRESENTATIVE REGARDING A PATIENT WHO IS IMPLANTED WITH A NEUROSTIMULATOR FOR NON-MALIGNANT PAIN. IT WAS REPORTED THAT THE PATIENT IS HAVING DIFFICULTY COMMUNICATING AND RECHARGING THEIR TWO IMPLANTABLE NEUROSTIMULATORS (INSS). THERE WAS DISCREPANCY WITH THE EVENT DATE WHICH WAS ORIGINALLY NOTED AS THIS ISSUE OCCURRING SINCE THE DAY AFTER IMPLANT; HOWEVER, IT WAS ALSO NOTED AS STARTING ABOUT THREE WEEKS PRIOR TO THE REPORT. THE PATIENT HAS TWO PATIENT CONTROLLERS AND TWO RECHARGERS; THEY ARE NOT LABELED, BUT THEY HAVE MARKINGS TO WHERE THEY CAN TELL WHICH ONE IS FOR EACH INS. THE REP HAD USED THE TABLET AND CONFIRMED THAT THE RIGHT INS IS (B)(4), AND THE LEFT INS IS (B)(4). THESE TWO INSS WERE IMPLANTED ABOUT 2-3 INCHES APART. THE REP WENT THROUGH THE PAIRING SET-UP PROCESS USING THE RIGHT PATIENT CONTROLLER (PC) AND THE RIGHT INS. THE RIGHT INS DID NOT HAVE ANY CHARGE INITIALLY, BUT THEY WERE ABLE TO CHARGE TO 40%. HOWEVER, WHEN THE RECHARGER THERAPY MODULE (RTM) WAS DISCONNECTED, SOMETIMES THE PC WOULD CONNECT WITH THE INS, AND SOMETIMES IT WOULD SHOW ¿NO DEVICE FOUND¿ OR GO BACK TO THE ¿SET-UP¿ SCREEN. THE REP HAD ALREADY RESEATED THE BATTERY PACK IN THE PC MULTIPLE TIMES AND IS STILL HAVING ISSUES. THEY THEN USED THE OTHER PC TO ATTEMPT TO COMMUNICATE WITH THE LEFT INS, BUT IT WAS SHOWING THAT THE LEFT INS WAS DEPLETED. THE REP HAD DONE 20 MINUTES OF PASSIVE RECHARGE CYCLE ON THE LEFT INS AND HAS NOT GOTTEN ANYWHERE. THE REP TRIED TO RESEAT THE BATTERY PACK IN THE RIGHT PC AGAIN, AND THEN IT WAS DISPLAYING THE ¿NO DEVICE FOUND¿ SCREEN EVEN WHEN IT WAS PLACED RIGHT OVER THE INS, BUT THEN THEY WERE ABLE TO PROCEED WITH THE NORMAL CHARGING WHEN THE RTM WAS CONNECTED. IT WAS CONFIRMED THAT THE LEFT PC WAS LOCKED WHEN ATTEMPTING TO COMMUNICATE WITH THE RIGHT INS. THE REP STOPPED CHARGING THE RIGHT INS AND LOCKED THE RIGHT PC. THE REP WALKED THROUGH THE ¿SET-UP¿ PROCESS FOR THE LEFT INS AND LEFT PC. THE PC FOUND BOTH INS SERIAL NUMBERS, AND THE REP SELECTED (B)(4) TO COMPLETE SET-UP ON THE LEFT SIDE. AT THAT POINT, IT SHOWED THAT THE INS WAS EMPTY, BUT WHEN THE REP ATTEMPTED TO CHARGE IT, THE PC SHOWED ¿UNFAMILIAR DEVICE¿ AND CONTINUED TO SPIN, BUT THEN IT SHOWED POOR RECHARGE QUALITY. WHEN THE REP MOVED THE SAME RTM OVER THE RIGHT INS, IT LOCATED THE RIGHT INS IMMEDIATELY. THE LEFT PC HAD PAIRED TO THE RIGHT INS. THE REP COMPLETED ANOTHER ¿SET-UP¿ PROCESS WITH THE LEFT PC AGAIN TO RE-PAIR THE LEFT PC TO THE LEFT INS, AND AGAIN, THE LEFT PC COMMUNICATED WITH THE RIGHT INS AND WAS ABLE TO CHARGE NORMALLY WITH EXCELLENT QUALITY AND THE INS WAS AT 40% CHARGE. THE REP WAS THEN ABLE TO SUCCESSFULLY PAIR THE LEFT PC TO THE LEFT INS (B)(4). AFTER THE PAIRING WAS COMPLETED, THE REP WAS THEN ABLE TO USE THE LEFT PC TO COMMUNICATE WITH THE LEFT INS AND WAS ABLE TO CHARGE NORMALLY WHICH SHOWED THAT THE INS WAS AT 50% CHARGE. IT WAS REVIEWED THAT THE CHARGE LEVEL LIKELY CAME FROM THE PASSIVE RECHARGE SESSIONS ALREADY PERFORMED. THERE WERE NO PATIENT SYMPTOMS OR COMPLICATIONS REPORTED. ADDITIONAL INFORMATION WAS RECEIVED FROM A REP ON NOVEMBER 25, 2020. IT WAS REPORTED THE RIGHT PC WAS NOT COMMUNICATING WITH THE RIGHT INS ON THE DAY OF THE REPORT. IT WAS ONLY COMMUNICATING IF THE RTM WAS CONNECTED. AFTER GOING INTO TECH MODE AND RESELECTING THE IMPLANT DEVICE, THE PC STARTED WORKING FINE. IT WAS REVIEWED THAT IF THE WRONG RTM IS USED TO CHARGE THE INS, IT WILL DEPLETE THE TEL-M KEY. THE CAUSE OF THE COMMUNICATION ISSUES BETWEEN THE PC AND INS WAS DETERMINED TO BE THAT THE DEVICES ARE IMPLANTED TOO CLOSE TO EACH OTHER. THE PATIENT USES A HOMEMADE SHIELD MADE OF ALUMINUM FOIL IN A CLOTH NAPKIN WHICH SEEMS TO WORK, BUT IT IS VERY CUMBERSOME. THE ISSUE HAS NOT YET BEEN RESOLVED. RESETTING UP THE PC WAS PERFORMED FOR BOTH INSS, WHICH WOULD ONLY RESOLVE THE COMMUNICATION ISSUE TEMPORARILY. IF THE RTM EVER GOT NEAR THE INS, THE ISSUE WOULD REPEAT ITSELF. AFTER SPEAKING WITH THE PATIENT ON (B)(6) 2019, THE REP INDICATED THAT THE HOMEMADE SHIELD SEEMED TO BE WORKING PROPERLY FOR NOW. THERE WERE NO PATIENT SYMPTOMS OR COMPLICATIONS REPORTED AT THAT TIME. ADDITIONAL INFORMATION WAS OBTAINED FROM THE PATIENT ON APRIL 27, 2020. THEY PROVIDED THE SERIAL NUMBERS OF THEIR TWO RTMS. THEY ALSO MENTIONED THAT DUE TO THE TWO INS DEVICES BEING TOO CLOSE TOGETHER, THEY WILL BE HAVING SURGICAL INTERVENTION TO REVISE THE INS ON THEIR RIGHT SIDE. NO FURTHER INFORMATION COULD BE OBTAINED FROM THE PATIENT. NO FURTHER COMPLICATIONS WERE REPORTED OR ANTICIPATED. | IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. CONCOMITANT MEDICAL PRODUCTS: PRODUCT ID: 97755, SERIAL#: (B)(4), PRODUCT TYPE: RECHARGER; PRODUCT ID: 97755, SERIAL#: (B)(4), PRODUCT TYPE: RECHARGER; PRODUCT ID: 97715, SERIAL#: (B)(4), IMPLANTED: (B)(6) 2019, PRODUCT TYPE: IMPLANTABLE NEUROSTIMULATOR. IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. | STIMULATOR, SPINAL-CORD, TOTALLY IMPLANTED FOR PAIN RELIEF | MEDTRONIC PUERTO RICO OPERATIONS CO. | 97715 |  | LGW | 20200514 |  | No | HEALTH PROFESSIONAL | Stimulator, Spinal-Cord, Totally Implanted For Pain Relief
Malfunction | 20200422 | ['Battery Problem'] | INFORMATION RECEIVED BY MEDTRONIC INDICATED THAT THE INSULIN PUMP HAD POWER ERROR DETECTED ALARM. THE CUSTOMER WAS ABLE TO CLEAR THE ALARM AND REWIND THE INSULIN PUMP. THE CUSTOMER HAD CALLED PREVIOUSLY TO REPORT THE ERROR THAT RECURRED WITHIN ONE WEEK OF TROUBLESHOOTING. NO HARM REQUIRING MEDICAL INTERVENTION WAS REPORTED. THE INSULIN PUMP WILL BE RETURNED FOR ANALYSIS. |  | THIS REPORT IS PART OF A RETROSPECTIVE REVIEW AND REMEDIATION EFFORTS IN RESPONSE TO A WARNING LETTER. UPDATED H9: Z-0955-2020. MEDTRONIC, INC. (MEDTRONIC) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY MEDTRONIC, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. MEDTRONIC HAS MADE REASONABLE EFFORTS TO OBTAIN MORE COMPLETE INFORMATION IN THE TIME ALLOTTED AND HAS PROVIDED AS MUCH INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE AN ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY "DEFECTS" OR HAS "MALFUNCTIONED". THESE WORDS ARE INCLUDED IN THE FDA 3500A FORM AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. MEDTRONIC OBJECTS TO THE USE OF THESE WORDS AND OTHERS LIKE IT BECAUSE OF THE LACK OF DEFINITION AND THE CONNOTATIONS IMPLIED BY THESE TERMS. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT. SLEEP CURRENT MEASUREMENT AND ACTIVE CURRENT MEASUREMENT WITHIN SPECIFICATIONS. THE POWER MANAGEMENT TOOL CONFIRMED POWER ERROR 25, LOW BATTERY ALERT, POWER LOSS ALARM, WAS TRIGGERED WHEN THE BACKUP BATTERY LOADED VOLTAGE (LOADED VLITH) WAS LESS THAN 3.5V FOR 4 CONSECUTIVE HOURS DUE TO CONNECTOR RESISTANCE ELECTRICAL BOARD. AFTER DISCONNECTING AND RECONNECTING THE INTERNAL BATTERY CONNECTOR ON ELECTRICAL BOARD, DEVICE WAS MONITORED AND FUNCTIONED PROPERLY. UNABLE TO PERFORM THE DISPLACEMENT TEST AND THE P-CAP/RESERVOIR WILL NOT LOCK PROPERLY DUE TO MISSING RETAINER. DEVICE RECEIVED WITH MISSING RESERVOIR TUBE O-RING, CRACKED BATTERY TUBE THREADS, BROKEN RESERVOIR TUBE LIP, FADING SERIAL NUMBER LABEL AND CRACKED CASE CORNER OF BELT CLIP RAILS. (B)(4). CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. THE DEVICE WILL BE RETURNED FOR ANALYSIS AND FURTHER INFORMATION WILL FOLLOW ONCE THE ANALYSIS HAS BEEN COMPLETED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. | ARTIFICIAL PANCREAS DEVICE SYSTEM, THRESHOLD SUSPEND | MEDTRONIC PUERTO RICO OPERATIONS CO. | MMT-1715K | HG1KZM2 | OZO |  | 20200508 | Device was returned to manufacturer | LAY USER/PATIENT | Automated Insulin Dosing , Threshold Suspend
Malfunction | 20200422 | ['Device Alarm System', 'Computer Software Problem', 'Device Difficult to Program or Calibrate', 'Battery Problem', 'Power Problem'] | "INFORMATION RECEIVED BY MEDTRONIC INDICATED THAT THE INSULIN PUMP HAD DOWN ARROW BUTTON STICKY AND NON RESPONSIVE AND RECEIVED ERROR CODE." CUSTOMER STATED THE INSULIN PUMP HAD DIMINISHED BATTERY LIFE, REJECTED NEW BATTERIES AND NO LOW RESERVOIR WARNING. NO HARM REQUIRING MEDICAL INTERVENTION WAS REPORTED. THE INSULIN PUMP WILL NOT BE RETURNED FOR ANALYSIS. |  | (B)(4). CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. | PUMP, INFUSION, INSULIN, TO BE USED WITH INVASIVE GLUCOSE SENSOR | MEDTRONIC PUERTO RICO OPERATIONS CO. | MMT-723NAS | B5723NASJ | OYC |  |  | No | LAY USER/PATIENT | Pump, Infusion, Insulin, To Be Used With Invasive Glucose Sensor
Malfunction | 20200422 | ['Battery Problem', 'Power Problem', 'Patient Device Interaction Problem'] | "INFORMATION RECEIVED BY MEDTRONIC INDICATED THAT THE SCREEN OF THE INSULIN PUMP WAS SCRATCHED, THEY COULD SEE THROUGH UNDER THE RESERVOIR AND A PIECE WAS MISSING, HAD BEEN CHANGING THE BATTERY EVERY 2 WEEKS AND HAD REJECTED THE BATTERIES." NO HARM REQUIRING MEDICAL INTERVENTION WAS REPORTED. TROUBLESHOOTING WAS DECLINED. THE INSULIN PUMP WILL BE RETURNED FOR ANALYSIS. |  | DEVICE MONITORED FOR SEVERAL DAYS. DEVICE RECEIVED WITH ALL OPERATING CURRENTS WITHIN SPECIFICATION AND PASSED A21 ERROR TEST AND DISPLACEMENT TEST. NO UNEXPECTED FAILED BATTERY ALARM ANOMALIES NOTED. DEVICE RECEIVED WITH MINOR SCRATCHED LCD WINDOW, CRACKED RESERVOIR TUBE LIP AND CRACKED BATTERY TUBE THREADS. THE TEST INFUSION SET AND RESERVOIR DOES LOCK INTO PLACE. (B)(4). (B)(4). CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. | ARTIFICIAL PANCREAS DEVICE SYSTEM, THRESHOLD SUSPEND | MEDTRONIC PUERTO RICO OPERATIONS CO. | MMT-751LNAS | A6751LNASJ | OZO |  | 20200616 | Device was returned to manufacturer | LAY USER/PATIENT | Automated Insulin Dosing , Threshold Suspend
Malfunction | 20200422 | ['Battery Problem', 'Power Problem'] | INFORMATION RECEIVED BY MEDTRONIC INDICATED THAT THE INSULIN PUMP HAD POWER ERROR DETECTED ALARM. NO HARM REQUIRING MEDICAL INTERVENTION WAS REPORTED. THE INSULIN PUMP WILL BE RETURNED FOR ANALYSIS. |  | DEVICE TRACE DOWNLOAD ANALYSIS CONFIRMED THE PUMP ALARMED POWER ERROR 25 ALARM AND BATTERY POWER LOSS ALARM. THE ADAPT TOOL CONFIRMED POWER ERROR 25 ALARM AND BATTERY POWER LOSS ALARM WAS TRIGGERED WHEN THE BACKUP BATTERY LOADED VOLTAGE (LOADED VLITH) WAS LESS THAN 3.5V FOR 4 CONSECUTIVE HOURS DUE TO CONNECTOR RESISTANCE ELECTRICAL BOARD. AFTER DISCONNECTING AND RECONNECTING THE INTERNAL BATTERY CONNECTOR ON ELECTRICAL BOARD, DEVICE WAS MONITORED AND FUNCTIONED PROPERLY. (B)(4). CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. THE DEVICE WILL BE RETURNED FOR ANALYSIS AND FURTHER INFORMATION WILL FOLLOW ONCE THE ANALYSIS HAS BEEN COMPLETED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. THE INSULIN PUMP INVOLVED IN THIS EVENT IS THE 640G INSULIN INFUSION PUMP, WHICH IS NOT MARKETED IN THE UNITED STATES. HOWEVER, THE DEVICE IS SIMILAR TO THE PARADIGM REAL-TIME INSULIN INFUSION PUMP, WHICH IS MARKETED IN THE UNITED STATES. | PUMP, INFUSION, INSULIN, TO BE USED WITH INVASIVE GLUCOSE SENSOR | MEDTRONIC PUERTO RICO OPERATIONS CO. | MMT-1711K | HG1226C | OYC |  | 20200519 | Device was returned to manufacturer | LAY USER/PATIENT | Pump, Infusion, Insulin, To Be Used With Invasive Glucose Sensor
Malfunction | 20200406 | ['Battery Problem'] | IT WAS REPORTED THAT THE PUMP BATTERY WAS INTERMITTENTLY OBSERVED TO BE DEPLETING RAPIDLY. CUSTOMER'S BLOOD GLUCOSE LEVEL WAS 156 MG/DL. REPORTEDLY, CUSTOMER CONTINUED TO USE THE PUMP FOR INSULIN THERAPY. |  | THE DEVICE IS EXPECTED TO BE RETURNED; HOWEVER, THE DEVICE HAS NOT YET BEEN RECEIVED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED IF THE DEVICE IS RECEIVED. | AUTOMATED INSULIN DOSING, THRESHOLD SUSPEND | TANDEM DIABETES CARE | 1002717 |  | OZO |  | 20200520 | Device was returned to manufacturer | LAY USER/PATIENT | Automated Insulin Dosing , Threshold Suspend
Malfunction | 20200406 | ['Battery Problem'] | IT WAS REPORTED THAT THE PUMP BATTERY WAS DEPLETING QUICKLY. THE CUSTOMER DECLINED FOLLOW-UP FROM TANDEM TECHNICAL SUPPORT REGARDING THE ISSUE. THE CUSTOMER CONTINUED TO USE THE PUMP FOR INSULIN THERAPY. THERE WAS NO REPORTED ADVERSE IMPACT TO THE CUSTOMER¿S BLOOD GLUCOSE LEVEL. |  | NO PRODUCT WAS RETURNED FOR EVALUATION. SHOULD NEW RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. | INSULIN PUMP | TANDEM DIABETES CARE | 4628-003 |  | OYC |  |  | No | LAY USER/PATIENT | Pump, Infusion, Insulin, To Be Used With Invasive Glucose Sensor
Malfunction | 20200406 | ['Battery Problem'] | IT WAS REPORTED THAT THE PUMP BATTERY WAS DEPLETING QUICKLY. THE CUSTOMER DECLINED FOLLOW-UP FROM TANDEM TECHNICAL SUPPORT REGARDING THE ISSUE. THE CUSTOMER CONTINUED TO USE THE PUMP FOR INSULIN THERAPY. THERE WAS NO REPORTED ADVERSE IMPACT TO THE CUSTOMER¿S BLOOD GLUCOSE LEVEL. |  | NO PRODUCT WAS RETURNED FOR EVALUATION. SHOULD NEW RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. | INSULIN PUMP | TANDEM DIABETES CARE | 4628-003 |  | OYC |  |  | No | LAY USER/PATIENT | Pump, Infusion, Insulin, To Be Used With Invasive Glucose Sensor
Malfunction | 20200422 | ['Battery Problem', 'Power Problem'] | INFORMATION RECEIVED BY MEDTRONIC INDICATED THAT THE INSULIN PUMP HAD A REPLACE BATTERY ALARM. THE CUSTOMER REPORTED THAT THEY DID NOT RECEIVED THE LOW BATTERY ALERT PRIOR TO THE REPLACE BATTERY ALARM. THE CUSTOMER REPORTED THAT THUS WAS THE SECOND OCCURRENCE OF REPLACE BATTERY ALERT OR REPLACE BATTERY NOW WITHOUT A LOW BATTERY WARNING. NO HARM REQUIRING MEDICAL INTERVENTION WAS REPORTED. THE INSULIN PUMP WILL BE RETURNED FOR ANALYSIS. |  | DEVICE PASSED THE SELF TEST, SLEEP CURRENT MEASUREMENT, ACTIVE CURRENT MEASUREMENT, AND THE DISPLACEMENT TEST. THE ADAPT TOOL INDICATED ABNORMAL BEHAVIOR OF THE LITHIUM BACKUP BATTERY LOADED VLITH VOLTAGE AS SHOWN ON THE POWER MANAGEMENT GRAPH AND CONFIRMED LOW BATTERY AND POWER LOSS ALARMS DUE TO CONNECTOR RESISTANCE ELECTRICAL BOARD. AFTER DISCONNECTING AND RECONNECTING THE INTERNAL BATTERY CONNECTOR ON ELECTRICAL BOARD, DEVICE WAS MONITORED AND FUNCTIONED PROPERLY. NO UNEXPECTED REPLACE BATTERY NOW OR BATTERY FAILED ALARMS NOTED. (B)(4). CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. THE DEVICE WILL BE RETURNED FOR ANALYSIS AND FURTHER INFORMATION WILL FOLLOW ONCE THE ANALYSIS HAS BEEN COMPLETED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. | ARTIFICIAL PANCREAS DEVICE SYSTEM, SINGLE HORMONAL CONTROL | MEDTRONIC PUERTO RICO OPERATIONS CO. | MMT-1780KPK | HG2SY7M | OZP | 20210905 | 20200504 | Device was returned to manufacturer | LAY USER/PATIENT | Automated Insulin Dosing Device System, Single Hormonal Control
Malfunction | 20200321 | ['Battery Problem'] | IT WAS REPORTED THAT THE PUMP BATTERY DEPLETES FASTER THAN EXPECTED. THERE WAS NO REPORTED IMPACT TO THE CUSTOMER'S BLOOD GLUCOSE LEVEL. CUSTOMER REVERTED TO MANUAL INJECTIONS FOR INSULIN THERAPY. |  | THE DEVICE IS EXPECTED TO BE RETURNED; HOWEVER, THE DEVICE HAS NOT YET BEEN RECEIVED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED IF THE DEVICE IS RECEIVED. | AUTOMATED INSULIN DOSING, THRESHOLD SUSPEND | TANDEM DIABETES CARE | 1000096 |  | OZO |  | 20200511 | Device was returned to manufacturer | LAY USER/PATIENT | Automated Insulin Dosing , Threshold Suspend
Malfunction | 20200422 | ['Battery Problem'] | INFORMATION RECEIVED BY MEDTRONIC INDICATED THAT THE INSULIN PUMP HAD A POWER ERROR DETECTED ALARM. CUSTOMER STATED THE BATTERY WAS CHANGED HOWEVER THE ISSUE WAS NOT RESOLVED. NO HARM REQUIRING MEDICAL INTERVENTION WAS REPORTED. THE INSULIN PUMP WILL BE RETURNED FOR ANALYSIS. |  | (B)(4). CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. THE DEVICE WILL BE RETURNED FOR ANALYSIS AND FURTHER INFORMATION WILL FOLLOW ONCE THE ANALYSIS HAS BEEN COMPLETED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. THE INSULIN PUMP INVOLVED IN THIS EVENT IS THE 640G INSULIN INFUSION PUMP, WHICH IS NOT MARKETED IN THE UNITED STATES. HOWEVER, THE DEVICE IS SIMILAR TO THE PARADIGM REAL-TIME INSULIN INFUSION PUMP, WHICH IS MARKETED IN THE UNITED STATES. THIS REPORT IS PART OF A RETROSPECTIVE REVIEW AND REMEDIATION EFFORTS IN RESPONSE TO A WARNING LETTER. UPDATED H9: Z-0958-2020. MEDTRONIC, INC. (MEDTRONIC) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY MEDTRONIC, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. MEDTRONIC HAS MADE REASONABLE EFFORTS TO OBTAIN MORE COMPLETE INFORMATION IN THE TIME ALLOTTED AND HAS PROVIDED AS MUCH INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE AN ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY "DEFECTS" OR HAS "MALFUNCTIONED". THESE WORDS ARE INCLUDED IN THE FDA 3500A FORM AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. MEDTRONIC OBJECTS TO THE USE OF THESE WORDS AND OTHERS LIKE IT BECAUSE OF THE LACK OF DEFINITION AND THE CONNOTATIONS IMPLIED BY THESE TERMS. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT. DEVICE PASSED DISPLACEMENT TEST, SLEEP CURRENT MEASUREMENT, ACTIVE CURRENT MEASUREMENT AND SELFTEST. DEVICE TRACE DOWNLOAD ANALYSIS CONFIRMED THE DEVICE ALARMED POWER ERROR 25. THE POWER MANAGEMENT TOOL CONFIRMED POWER ERROR 25 DUE TO NON-SLEEP ANOMALY SOFTWARE ERROR. | PUMP, INFUSION, INSULIN, TO BE USED WITH INVASIVE GLUCOSE SENSOR | MEDTRONIC PUERTO RICO OPERATIONS CO. | MMT-1712B | HG0Q5YM | OYC |  | 20200610 | Device was returned to manufacturer | LAY USER/PATIENT | Pump, Infusion, Insulin, To Be Used With Invasive Glucose Sensor
Malfunction | 20200422 | ['Obstruction of Flow', 'Battery Problem'] | CUSTOMER REPORTED VIA PHONE CALL LOW BATTERY AND NO DELIVERY ALARMS ON THE INSULIN PUMP. TROUBLESHOOTING WAS NOT PERFORMED. NO HARM REQUIRING MEDICAL INTERVENTION WAS REPORTED. THE INSULIN PUMP WILL NOT BE RETURNED FOR ANALYSIS. |  | (B)(4). CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. | ARTIFICIAL PANCREAS DEVICE SYSTEM, THRESHOLD SUSPEND | MEDTRONIC PUERTO RICO OPERATIONS CO. | MMT-751NAH | A2751NAHJ | OZO |  |  | No | LAY USER/PATIENT | Automated Insulin Dosing , Threshold Suspend
Malfunction | 20200423 | ['Battery Problem'] | INFORMATION RECEIVED BY MEDTRONIC INDICATED THAT THE INSULIN PUMP RECEIVED POWER ERROR DETECTED ALARM. CUSTOMER WAS ABLE TO SUCCESSFULLY CLEAR THE ALARM, BUT UNABLE TO COMPLETE INSULIN PUMP REWIND. NO HARM REQUIRING MEDICAL INTERVENTION WAS REPORTED. THE INSULIN PUMP WILL BE RETURNED FOR ANALYSIS. |  | (B)(4). CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. THE DEVICE WILL BE RETURNED FOR ANALYSIS AND FURTHER INFORMATION WILL FOLLOW ONCE THE ANALYSIS HAS BEEN COMPLETED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. DEVICE PASSED SLEEP CURRENT MEASUREMENT, ACTIVE CURRENT MEASUREMENT AND DISPLACEMENT TEST. DEVICE RECEIVED ERROR 25 DUE TO CONNECTOR RESISTANCE ISSUE. | ARTIFICIAL PANCREAS DEVICE SYSTEM, SINGLE HORMONAL CONTROL | MEDTRONIC PUERTO RICO OPERATIONS CO. | MMT-1780KPK | HG1V3E9 | OZP |  | 20200507 | Device was returned to manufacturer | LAY USER/PATIENT | Automated Insulin Dosing Device System, Single Hormonal Control
Malfunction | 20191105 | ['Battery Problem', 'Material Integrity Problem'] | INFORMATION RECEIVED BY MEDTRONIC INDICATED THAT THE INSULIN PUMP HAD BATTERY CAP WORN AND FREQUENTLY BATTERY CHANGED. NO HARM REQUIRING MEDICAL INTERVENTION WAS REPORTED. THE INSULIN PUMP WAS NOT RETURNED FOR ANALYSIS. |  | UNIT RECEIVED WITH BLANK DISPLAY, PROBLEM ISOLATED TO LCD BOARD. UNABLE TO PERFORM OPERATING CURRENTS, SELF-TEST, UNEXPECTED RESTART ERROR TEST AND DISPLACEMENT TEST OR VERIFY LOW BATTERY ALARM DUE TO BLANK DISPLAY. UNIT HAD STRIPPED BATTERY CAP COIN SLOT (P). THE RESERVOIR TUBE LIP WAS TESTED WITH THE TEST RESERVOIR LOCKED IN PLACE PROPERLY AND NO ANOMALY NOTED. (B)(4). CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. | ARTIFICIAL PANCREAS DEVICE SYSTEM, THRESHOLD SUSPEND | MEDTRONIC PUERTO RICO OPERATIONS CO. | MMT-751LNAH | A4751LNAHJ | OZO |  | 20200326 | Device was returned to manufacturer | LAY USER/PATIENT | Automated Insulin Dosing , Threshold Suspend
Malfunction | 20200423 | ['Device Difficult to Program or Calibrate', 'Obstruction of Flow', 'Battery Problem'] | INFORMATION RECEIVED BY MEDTRONIC INDICATED THAT THE BUTTONS WERE HARD TO PRESS AND UNRESPONSIVE AT TIMES. THEY ALSO RECEIVED DIMINISHED BATTERY LIFE, REJECTED NEW BATTERIES AND UNEXPLAINED NO DELIVERY ALERTS. NO HARM REQUIRING MEDICAL INTERVENTION WAS REPORTED. THE INSULIN PUMP WILL NOT BE RETURNED FOR ANALYSIS. |  | (B)(4). CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. | PUMP, INFUSION, INSULIN, TO BE USED WITH INVASIVE GLUCOSE SENSOR | MEDTRONIC PUERTO RICO OPERATIONS CO. | MMT-723LNAS | A7723LNASJ | OYC |  |  | No | LAY USER/PATIENT | Pump, Infusion, Insulin, To Be Used With Invasive Glucose Sensor
Malfunction | 20200423 | ['Failure to Cycle', 'Device Difficult to Program or Calibrate', 'Battery Problem'] | INFORMATION RECEIVED BY MEDTRONIC INDICATED THAT INSULIN PUMP HAD COMPROMISED FORCE SENSOR SYSTEM ALARM AND BUTTON ERROR ALARM. NO HARM REQUIRING MEDICAL INTERVENTION WAS REPORTED. THE CUSTOMER WAS ASSISTED WITH TROUBLESHOOTING. CUSTOMER STATED THAT DRIVE SUPPORT CAP WAS NORMAL. THE DEVICE WILL BE RETURNED FOR ANALYSIS. |  | DEVICE RECEIVED WITH INTERMITTENT BUTTON RESPONSE DUE TO FLATTENED BUTTON DOME SWITCH AND PARTIAL UNLOCKED J2/LCD KEYPAD CONNECTOR NOTED DURING VISUAL INSPECT. UNABLE TO PERFORM ALL FUNCTIONAL TESTING INCLUDING THE DISPLACEMENT, OPERATING CURRENTS, A21 ERROR TEST, REWIND, BASIC OCCLUSION, OCCLUSION, PRIME OR A33 AND EXCESSIVE NO DELIVERY TEST OR VERIFY A33 ALARM AND CHANGE OR BATTERY LASTS LESS THAN EXPECTED DUE TO BUTTONS NOT RESPONDING. (B)(4). CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. THE DEVICE WILL BE RETURNED FOR ANALYSIS AND FURTHER INFORMATION WILL FOLLOW ONCE THE ANALYSIS HAS BEEN COMPLETED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. | ARTIFICIAL PANCREAS DEVICE SYSTEM, THRESHOLD SUSPEND | MEDTRONIC PUERTO RICO OPERATIONS CO. | MMT-551NAS | A3551NASJ | OZO |  | 20200505 | Device was returned to manufacturer | LAY USER/PATIENT | Automated Insulin Dosing , Threshold Suspend
Malfunction | 20200422 | ['Battery Problem'] | INFORMATION RECEIVED BY MEDTRONIC INDICATED THAT THE INSULIN PUMP HAD A POWER ERROR DETECTED ALARM. CUSTOMER STATED THE BATTERY WAS CHANGED HOWEVER THE ALARM RECURRED WITHIN A WEEK AND THE ISSUE WAS NOT RESOLVED. NO HARM REQUIRING MEDICAL INTERVENTION WAS REPORTED. THE INSULIN PUMP WILL BE RETURNED FOR ANALYSIS. |  | THIS REPORT IS PART OF A RETROSPECTIVE REVIEW AND REMEDIATION EFFORTS IN RESPONSE TO A WARNING LETTER. UPDATED H9: Z-0955-2020. MEDTRONIC, INC. (MEDTRONIC) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY MEDTRONIC, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. MEDTRONIC HAS MADE REASONABLE EFFORTS TO OBTAIN MORE COMPLETE INFORMATION IN THE TIME ALLOTTED AND HAS PROVIDED AS MUCH INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE AN ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY "DEFECTS" OR HAS "MALFUNCTIONED". THESE WORDS ARE INCLUDED IN THE FDA 3500A FORM AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. MEDTRONIC OBJECTS TO THE USE OF THESE WORDS AND OTHERS LIKE IT BECAUSE OF THE LACK OF DEFINITION AND THE CONNOTATIONS IMPLIED BY THESE TERMS. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT. (B)(4). CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. THE DEVICE WILL BE RETURNED FOR ANALYSIS AND FURTHER INFORMATION WILL FOLLOW ONCE THE ANALYSIS HAS BEEN COMPLETED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. DEVICE PASSED SLEEP CURRENT MEASUREMENT, ACTIVE CURRENT MEASUREMENT, SELF TEST AND DISPLACEMENT TEST. HOWEVER, DEVICE TRACE DOWNLOAD ANALYSIS CONFIRMED THE PUMP ALARMED POWER ERROR 25 AND REPLACE BATTERY ALERT/REPLACE BATTERY NOW ALARM. THE POWER MANAGEMENT TOOL CONFIRMED POWER ERROR 25 AND REPLACE BATTERY ALERT/REPLACE BATTERY NOW ALARM WAS TRIGGERED WHEN THE BACKUP BATTERY LOADED VOLTAGE (LOADED VLITH) WAS LESS THAN 3.5V FOR 4 CONSECUTIVE HOURS DUE TO CONNECTOR RESISTANCE ELECTRICAL BOARD. AFTER DISCONNECTING AND RECONNECTING THE INTERNAL BATTERY CONNECTOR ON ELECTRICAL BOARD, DEVICE WAS MONITORED AND FUNCTIONED PROPERLY. | ARTIFICIAL PANCREAS DEVICE SYSTEM, THRESHOLD SUSPEND | MEDTRONIC PUERTO RICO OPERATIONS CO. | MMT-1715K | HG206HN | OZO |  | 20200501 | Device was returned to manufacturer | LAY USER/PATIENT | Automated Insulin Dosing , Threshold Suspend
Malfunction | 20200423 | ['Mechanical Problem', 'Device Difficult to Program or Calibrate', 'Battery Problem'] | INFORMATION RECEIVED BY MEDTRONIC INDICATED THAT THE ONCE OR TWICE BUTTONS WERE HARDER TO PRESS. THE CUSTOMER ALSO STATED THAT THE INSULIN PUMP HAD REJECTED NEW BATTERIES AND MOTOR SOMETIMES MAKES A GRINDING NOISE. NO HARM REQUIRING MEDICAL INTERVENTION WAS REPORTED. THE INSULIN PUMP WILL NOT BE RETURNED FOR ANALYSIS. |  | (B)(4). CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. | PUMP, INFUSION, INSULIN, TO BE USED WITH INVASIVE GLUCOSE SENSOR | MEDTRONIC PUERTO RICO OPERATIONS CO. | MMT-723RNAB | A1723RNABJ | OYC |  |  | No | LAY USER/PATIENT | Pump, Infusion, Insulin, To Be Used With Invasive Glucose Sensor
Malfunction | 20200422 | ['No Display/Image', 'Battery Problem'] | INFORMATION RECEIVED BY MEDTRONIC INDICATED THAT THE INSULIN PUMP HAD A BLANK DISPLAY AND BATTERY ALERT. CUSTOMER STATED THE BATTERY WAS CHANGED BUT THE ISSUE WAS NOT RESOLVED. NO HARM REQUIRING MEDICAL INTERVENTION WAS REPORTED. THE INSULIN PUMP WILL BE RETURNED FOR ANALYSIS. |  | THE PUMP WAS RECEIVED WITH A BLANK DISPLAY DUE TO CRACKED CASE AT BATTERY TUBE SIDE AND CRACKED BATTERY TUBE THREADS. THE CRACKED CASE AT BATTERY TUBE SIDE SHIFTED THE BATTERY TUBE OUT OF POSITION NOT ALLOWING PROPER CONTACT BETWEEN THE BATTERY CAP CONTACT AND BATTERY TUBE. UNABLE TO PERFORM THE SELF TEST, SLEEP CURRENT MEASUREMENT, ACTIVE CURRENT MEASUREMENT, AND THE DISPLACEMENT TEST OR VERIFY UNEXPECTED BATTERY POWER LOSS DUE TO BLANK DISPLAY. (B)(4). CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. THE DEVICE WILL BE RETURNED FOR ANALYSIS AND FURTHER INFORMATION WILL FOLLOW ONCE THE ANALYSIS HAS BEEN COMPLETED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. | ARTIFICIAL PANCREAS DEVICE SYSTEM, SINGLE HORMONAL CONTROL | MEDTRONIC PUERTO RICO OPERATIONS CO. | MMT-1780KPK | HG239B7 | OZP |  | 20200430 | Device was returned to manufacturer | LAY USER/PATIENT | Automated Insulin Dosing Device System, Single Hormonal Control
Malfunction | 20200422 | ['Battery Problem', 'Power Problem'] | INFORMATION RECEIVED BY MEDTRONIC INDICATED THAT THE INSULIN PUMP WAS OUT OF INSULIN AND IT HAD NOT ALARMED THAT THE RESERVOIR WAS LOW. THE CUSTOMER EXPERIENCED HIGH BLOOD GLUCOSE. CUSTOMER ALSO STATED THEY HAD TO CHANGE BATTERY FREQUENTLY. THE INSULIN PUMP WILL NOT BE RETURNED FOR ANALYSIS. |  | DEVICE MONITORED FOR SEVERAL DAYS. DEVICE RECEIVED WITH ALL OPERATING CURRENTS WITHIN SPEC AND PASSED A21 ERROR TEST AND DISPLACEMENT TEST. NO UNEXPECTED BATTERY OUT LIMIT ALARM ANOMALIES NOTED. NO UNEXPECTED LOW BATTERY ALARM ANOMALIES NOTED. NO UNEXPECTED AUDIO/VIBRATE ALARM ANOMALIES NOTED. (B)(4). (B)(4). CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. | ARTIFICIAL PANCREAS DEVICE SYSTEM, THRESHOLD SUSPEND | MEDTRONIC PUERTO RICO OPERATIONS CO. | MMT-751LNAS | A6751LNASJ | OZO |  | 20200630 | Device was returned to manufacturer | LAY USER/PATIENT | Automated Insulin Dosing , Threshold Suspend
Malfunction | 20200423 | ['Mechanical Problem', 'Obstruction of Flow', 'Battery Problem'] | INFORMATION RECEIVED BY MEDTRONIC INDICATED THAT THE INSULIN PUMP HAD MOTOR ERROR. THE CUSTOMER ALSO STATED THAT THE INSULIN PUMP HAD REJECTED NEW BATTERIES AND RECEIVED RANDOM NO DELIVERY ALARM. NO HARM REQUIRING MEDICAL INTERVENTION WAS REPORTED. THE INSULIN PUMP WILL NOT BE RETURNED FOR ANALYSIS. |  | (B)(4). CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. | PUMP, INFUSION, INSULIN, TO BE USED WITH INVASIVE GLUCOSE SENSOR | MEDTRONIC PUERTO RICO OPERATIONS CO. | MMT-723NAS | B5723NASJ | OYC |  |  | No | LAY USER/PATIENT | Pump, Infusion, Insulin, To Be Used With Invasive Glucose Sensor
Malfunction | 20200424 | ['Battery Problem'] | INFORMATION RECEIVED BY MEDTRONIC INDICATED THAT THE INSULIN PUMP HAD POWER LOSS ALARM. NO HARM REQUIRING MEDICAL INTERVENTION WAS REPORTED. THE INSULIN PUMP WILL NOT BE RETURNED FOR ANALYSIS. |  | THE INSULIN PUMP INVOLVED IN THIS EVENT IS THE 640G INSULIN INFUSION PUMP, WHICH IS NOT MARKETED IN THE UNITED STATES. HOWEVER, THE DEVICE IS SIMILAR TO THE PARADIGM REAL-TIME INSULIN INFUSION PUMP, WHICH IS MARKETED IN THE UNITED STATES. (B)(4). CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. | PUMP, INFUSION, INSULIN, TO BE USED WITH INVASIVE GLUCOSE SENSOR | MEDTRONIC PUERTO RICO OPERATIONS CO. | MMT-1711K | HG0HRDZ | OYC |  |  | No | LAY USER/PATIENT | Pump, Infusion, Insulin, To Be Used With Invasive Glucose Sensor
Injury | 20190601 | ['Battery Problem', 'Adverse Event Without Identified Device or Use Problem'] | ADDITIONAL INFORMATION WAS RECEIVED FROM A CONSUMER REGARDING THE PATIENT. IT WAS REPORTED THAT THE STIMULATOR BATTERY WORKED GREAT AND WAS REMOVED ON (B)(6) 2019 DUE TO NORMAL BATTERY DEPLETION. THE LEADS WERE LEFT IMPLANTED. THE PATIENT NOTED THAT AFTER REMOVAL, HE CRASHED ON THE TABLE, SO HE WAS SENT TO THE TRAUMA UNIT FOR A HEART RATE OF 25. THE PATIENT WAS SHOCKED TWO TIMES AND SPENT 4 DAYS IN THE INTENSIVE CARE UNIT. HE NOTED THAT HE WALKED OUT AND IS OKAY NOW. THE PATIENT IS AWAITING REPLACEMENT OF THE IMPLANTABLE NEUROSTIMULATOR. THE PATIENT NOTED THAT THE REASON FOR THE INITIAL CALL TO PATIENT SERVICES ON (B)(6) 2020 WAS BECAUSE SINCE THE LEADS WERE STILL IN PLACE, HE COULD ONLY HAVE A HEAD-ONLY MRI DONE. THE PATIENT WAS IN THE HOSPITAL FOR UNRELATED LOW BLOOD SUGAR AND IS OKAY NOW. THERE WERE NO FURTHER COMPLICATIONS REPORTED OR ANTICIPATED. | INFORMATION WAS RECEIVED FROM A CONSUMER AND HEALTHCARE PROFESSIONAL (HCP) REGARDING A PATIENT WITH AN IMPLANTABLE NEUROSTIMULATOR (INS). IT WAS REPORTED THAT IN (B)(6) 2020, THE PATIENT WAS GETTING A REPLACEMENT INS DUE TO SEEING THE END OF SERVICE (EOS) MESSAGE. THE PROCEDURE WAS ABORTED. A LEAD REMAINED IMPLANTED. NO FURTHER COMPLICATIONS WERE REPORTED. | H6. DEVICE CODES HAVE BEEN UPDATED TO REFLECT THE ADDITIONAL INFORMATION RECEIVED. MEDTRONIC IS SUBMITTING THIS REPORT TO COMPLY WITH FDA REPORTING REGULATIONS UNDER 21 CFR PARTS 4 AND 803. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY MEDTRONIC, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. MEDTRONIC HAS MADE REASONABLE EFFORTS TO OBTAIN MORE COMPLETE INFORMATION AND HAS PROVIDED AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE AN ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY ¿DEFECTS¿ OR HAS ¿MALFUNCTIONED¿. THESE WORDS ARE INCLUDED IN THE FDA 3500A FORM AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REGULATORY REPORTING. MEDTRONIC OBJECTS TO THE USE OF THESE WORDS AND OTHERS LIKE THEM BECAUSE OF THE LACK OF DEFINITION AND THE CONNOTATIONS IMPLIED BY THESE TERMS. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT. IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. | STIMULATOR, SPINAL-CORD, TOTALLY IMPLANTED FOR PAIN RELIEF | MEDTRONIC MED REL MEDTRONIC PUERTO RICO | 37703 |  | LGW | 20140714 |  | No | HEALTH PROFESSIONAL | Stimulator, Spinal-Cord, Totally Implanted For Pain Relief
Malfunction | 20200423 | ['Obstruction of Flow', 'Date/Time-Related Software Problem', 'Battery Problem', 'Power Problem'] | INFORMATION RECEIVED BY MEDTRONIC INDICATED THAT THE CLOCK OF THE INSULIN PUMP WAS SLOW. THE CUSTOMER STATED THAT THE INSULIN PUMP HAD UNEXPLAINABLE NO DELIVERY ALARMS AND INSULIN PUMP HAD REJECTED THE BRAND NEW BATTERIES. THE BATTERY ONLY LASTING 7-10 DAYS. NO HARM REQUIRING MEDICAL INTERVENTION WAS REPORTED. THE INSULIN PUMP WILL NOT BE RETURNED FOR ANALYSIS. |  | (B)(4). CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. DEVICE PASSED DISPLACEMENT TEST, SELF TEST, OFF NO POWER TEST, A21 ERROR TEST AND ALL OPERATING CURRENTS TESTED WITHIN THE SPECIFICATION. NO FAILED BATTERY TEST ALARM OR CHARGE/BATTERY ANOMALY WERE NOTED. DEVICE PASSED REWIND TEST, BASIC OCCLUSION TEST, PRIME/A33 TEST, EXCESSIVE NO DELIVERY TEST AND OCCLUSION TEST. NO UNEXPECTED NO DELIVERY ALARM NOTED. DEVICE PROGRAMMED WITH THE CURRENT TIME AND DATE AND MONITORED FOR SEVERAL DAYS. THE TIME AND DATE IS FUNCTIONING PROPERLY. (B)(4). | ARTIFICIAL PANCREAS DEVICE SYSTEM, THRESHOLD SUSPEND | MEDTRONIC PUERTO RICO OPERATIONS CO. | MMT-751LNAS | A5751LNASJ | OZO |  | 20200605 | Device was returned to manufacturer | LAY USER/PATIENT | Automated Insulin Dosing , Threshold Suspend
Malfunction | 20200423 | ['Battery Problem', 'Power Problem'] | INFORMATION RECEIVED BY MEDTRONIC INDICATED THAT THE INSULIN PUMP HAD POWER ERROR DETECTED ALARM. IT WAS REPORTED THAT THE CUSTOMER INSERTED NEW BATTERY AND RECEIVED POWER LOSS ALARM. NO HARM REQUIRING MEDICAL INTERVENTION WAS REPORTED. THE INSULIN PUMP WILL BE RETURNED FOR ANALYSIS. |  | (B)(4). CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. THE DEVICE WILL BE RETURNED FOR ANALYSIS AND FURTHER INFORMATION WILL FOLLOW ONCE THE ANALYSIS HAS BEEN COMPLETED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. THE INSULIN PUMP INVOLVED IN THIS EVENT IS THE 640G INSULIN INFUSION PUMP, WHICH IS NOT MARKETED IN THE UNITED STATES. HOWEVER, THE DEVICE IS SIMILAR TO THE PARADIGM REAL-TIME INSULIN INFUSION PUMP, WHICH IS MARKETED IN THE UNITED STATES. DEVICE RECEIVED WITH BLANK DISPLAY DUE TO CORRODED BATTERY TUBE SPRING CONTACT AND CORRODED BATTERY TUBE. UNABLE TO PERFORM FUNCTIONAL TESTING INCLUDING THE OPERATING CURRENTS AND SELF TEST OR VERIFY ERROR 35 ALARM AND UNEXPECTED BATTERY POWER LOSS ANOMALY DUE TO BLANK DISPLAY. | PUMP, INFUSION, INSULIN, TO BE USED WITH INVASIVE GLUCOSE SENSOR | MEDTRONIC PUERTO RICO OPERATIONS CO. | MMT-1711B | HG10EH4 | OYC |  | 20200528 | Device was returned to manufacturer | LAY USER/PATIENT | Pump, Infusion, Insulin, To Be Used With Invasive Glucose Sensor
Malfunction | 20200423 | ['Device Alarm System', 'Device Difficult to Program or Calibrate', 'Battery Problem', 'Power Problem'] | INFORMATION RECEIVED BY MEDTRONIC INDICATED THAT THE INSULIN PUMP BUTTONS WERE LESS RESPONSIVE, WAS HARDER TO PRESS AND HAD TO BE PRESSED SEVERAL TIMES TO WORK, HAD BAD BATTERY ERRORS, HAD BEEN REPLACING BATTERIES MORE OFTEN THAN NORMAL AND COULD NOT HEAR THE INSULIN PUMP AT TIMES. NO HARM REQUIRING MEDICAL INTERVENTION WAS REPORTED. THE INSULIN PUMP WILL NOT BE RETURNED FOR ANALYSIS. |  | (B)(4). CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. | ARTIFICIAL PANCREAS DEVICE SYSTEM, THRESHOLD SUSPEND | MEDTRONIC PUERTO RICO OPERATIONS CO. | MMT-751NAS | A4751NASJ | OZO |  |  | No | LAY USER/PATIENT | Automated Insulin Dosing , Threshold Suspend
Malfunction | 20200408 | ['Battery Problem'] | IT WAS REPORTED THAT THE BATTERY EXHIBITED CRITICAL BATTERY ALARMS WHEN THE BATTERY¿S CAPACITY WAS AT 25%. THE BATTERY WAS REMOVED FROM SERVICE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. | IT WAS FURTHER REPORTED THAT THE BATTERY REMAINS IN USE. | A SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR DEVICE EVALUATION. PRODUCT EVENT SUMMARY: THE BATTERY WAS NOT RETURNED FOR EVALUATION. LOG FILE ANALYSIS REVEALED THAT TWO (2) CRITICAL BATTERY ALARMS INVOLVING THE BATTERY WERE LOGGED ON 08-APR-2020 AT 20:52:48 AND ON 12-APR-2020 AT 19:40:28. THE DATA POINT PRIOR TO EACH CRITICAL BATTERY ALARM REVEALED THAT THE BATTERY WAS AT 24% RELATIVE STATE OF CHARGE (RSOC). DURING THE CRITICAL BATTERY ALARMS, THE BATTERY DEPLETED FROM 24% RSOC TO 0% RSOC. AS A RESULT, THE REPORTED EVENT WAS CONFIRMED. DURING THESE INSTANCES, THE BATTERY WAS NOT THE ACTIVE BATTERY. GIVEN THAT THE CAPACITY DROPPED TO 0% RSOC MAY BE INDICATIVE OF AN ESTIMATION ERROR. HOWEVER, ANALYSIS COULD NOT BE PERFORMED SINCE THE DEVICE WAS NOT RETURNED. THE MOST LIKELY ROOT CAUSE OF THE REPORTED EVENT CAN BE ATTRIBUTED TO AN ESTIMATION ERROR, WHICH RESULTED IN THE BATTERY'S CAPACITY DROPPING BELOW 10% RSOC, TRIGGERING CRITICAL BATTERY ALARMS. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. ### MEDTRONIC IS SUBMITTING THIS REPORT TO COMPLY WITH FDA REPORTING REGULATIONS UNDER 21 CFR PARTS 4 AND 803. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY MEDTRONIC, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. MEDTRONIC HAS MADE REASONABLE EFFORTS TO OBTAIN MORE COMPLETE INFORMATION AND HAS PROVIDED AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE AN ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY ¿DEFECTS¿ OR HAS ¿MALFUNCTIONED¿. THESE WORDS ARE INCLUDED IN THE FDA 3500A FORM AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REGULATORY REPORTING. MEDTRONIC OBJECTS TO THE USE OF THESE WORDS AND OTHERS LIKE THEM BECAUSE OF THE LACK OF DEFINITION AND THE CONNOTATIONS IMPLIED BY THESE TERMS. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. A GOOD FAITH EFFORT WILL BE MADE TO OBTAIN THE APPLICABLE INFORMATION RELEVANT TO THE REPORT. IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANS PLANTATION AND DESTINATION THERAPY IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY END-STAGE LEFT VENTRICULAR HEART FAILURE. THE SYSTEM IS DESIGNED FOR IN-HOSPITAL AND OUT-OF-HOSPITAL SETTINGS. CONCOMITANT OF MEDICAL PRODUCTS: DTBA1D1 ICD IMPLANTED (B)(6) 2015; 419588 LEAD IMPLANTED (B)(6); 694765 LEAD IMPLANTED (B)(6) 2010; 407652 LEAD IMPLANTED (B)(6) 2010. INVESTIGATION OF THIS EVENT IS PENDING AND A SUPPLEMENTAL REPORT WILL BE SENT UPON ITS COMPLETION. IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. A SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR A CORRECTION TO DESCRIBE EVENT OR PROBLEM REGARDING THE DEVICE DISPOSITION. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC IS SUBMITTING THIS REPORT TO COMPLY WITH FDA REPORTING REGULATIONS UNDER 21 CFR PARTS 4 AND 803. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY MEDTRONIC, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. MEDTRONIC HAS MADE REASONABLE EFFORTS TO OBTAIN MORE COMPLETE INFORMATION AND HAS PROVIDED AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE AN ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY ¿DEFECTS¿ OR HAS ¿MALFUNCTIONED¿. THESE WORDS ARE INCLUDED IN THE FDA 3500A FORM AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REGULATORY REPORTING. MEDTRONIC OBJECTS TO THE USE OF THESE WORDS AND OTHERS LIKE THEM BECAUSE OF THE LACK OF DEFINITION AND THE CONNOTATIONS IMPLIED BY THESE TERMS. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. A GOOD FAITH EFFORT WILL BE MADE TO OBTAIN THE APPLICABLE INFORMATION RELEVANT TO THE REPORT. IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. | VENTRICULAR (ASSIST) BYPASS | HEARTWARE, INC. | 1650DE |  | DSQ | 20190731 |  | No | LAY USER/PATIENT | Ventricular (Assist) Bypass
Malfunction | 20200425 | ['Battery Problem'] | INFORMATION WAS RECEIVED FROM THE CONSUMER REGARDING A PATIENT WITH AN IMPLANTABLE NEUROSTIMULATOR (INS). THE PATIENT REPORTED THAT THE CONTROLLER WAS CHARGED TO 100% AND WHEN IT WAS PLUGGED IN THE RECHARGER (RTM) TO CHARGE INS AND THE CONTROLLER SCREEN WENT BLANK AND THE GREEN LIGHT STOPPED FLASHING. THE CONTROLLER BATTERY WAS REMOVED, AND THE CONTROLLER WAS PLUGGED INTO 2 DIFFERENT OUTLETS, BUT THE SCREEN STAYED BLANK. THE PATIENT REPORTED THAT THE INS WAS DEAD, AND THEY HAVE LOT OF PAIN IN THE LEG. |  | CONCOMITANT MEDICAL PRODUCTS: PRODUCT ID: 97745, SERIAL#: (B)(4), PRODUCT TYPE: PROGRAMMER, PATIENT. IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. | STIMULATOR, SPINAL-CORD, TOTALLY IMPLANTED FOR PAIN RELIEF | MEDTRONIC PUERTO RICO OPERATIONS CO. | 97715 |  | LGW | 20190528 |  | No | HEALTH PROFESSIONAL | Stimulator, Spinal-Cord, Totally Implanted For Pain Relief
Injury | 20200414 | ['Premature Discharge of Battery', 'Battery Problem'] | FOLLOWING THE BATTERY PERFORMANCE ALERT (BPA) ADVISORY, A BPA WAS RECEIVED BY THE CLINICIAN AND THE DEVICE WAS EXPLANTED. THE PATIENT CONDITION WAS STABLE. |  | THE RESULTS/METHOD AND CONCLUSION CODES ALONG WITH INVESTIGATION RESULTS WILL BE PROVIDED IN THE FINAL REPORT. THE DEVICE IS INCLUDED IN THE BATTERY PERFORMANCE ALERT ADVISORY ISSUED BY ABBOTT ON 28 AUGUST 2017. PREMATURE BATTERY DEPLETION WAS CONFIRMED BY ANALYSIS. A DEVICE EVALUATION WAS PERFORMED, AND NO SOURCES OF HIGH CURRENTS WERE NOTED. THE CAUSE OF THE PREMATURE BATTERY DEPLETION WAS CONSISTENT WITH LI CLUSTER FORMATION. FROM THESE ANALYSES, IN THE ABSENCE OF A HIGH CURRENT DRAW, IT IS PROBABLE THAT A LITHIUM CLUSTER-INDUCED SHORT CIRCUIT CAUSED PREMATURE BATTERY DEPLETION. LI CLUSTERS ARE A KNOWN DEPLETION MECHANISM FOR THESE ADVISORY PRODUCTS THAT HAVE BEEN INVESTIGATED AND ASSOCIATED WITH A FIELD ACTION IN OCTOBER 2016. THE RESULTS OF THE INVESTIGATION ARE INCONCLUSIVE SINCE THE DEVICE WAS NOT RETURNED FOR ANALYSIS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED. | IMPLANTABLE CARDIOVERTER DEFIBRILLATOR | ST. JUDE MEDICAL, INC.(CRM-SUNNYVALE) | CD2357-40Q | 4388537 | NVZ | 20151231 |  | Device was returned to manufacturer | HEALTH PROFESSIONAL | Pulse Generator, Permanent, Implantable
Malfunction | 20200422 | ['Failure to Charge', 'Battery Problem'] | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S INTERNAL BATTERY WAS REPLACED TO ADDRESS THE ISSUE. |  |  | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC | 1040007 |  | CBK |  | 20200420 | Device was returned to manufacturer | HEALTH PROFESSIONAL | Ventilator, Continuous, Facility Use
Malfunction | 20200422 | ['Failure to Charge', 'Battery Problem'] | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S INTERNAL BATTERY WAS REPLACED TO ADDRESS THE ISSUE. |  |  | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC | 1040007 |  | CBK |  | 20200421 | Device was returned to manufacturer | HEALTH PROFESSIONAL | Ventilator, Continuous, Facility Use
Malfunction |  | ['Component Missing', 'Battery Problem'] | INFORMATION WAS RECEIVED INDICATING THAT A SMITHS MEDICAL MEDFUSION PUMP'S BATTERY MALFUNCTIONED. THE PUMP ALSO NEEDED A BATTERY DOOR AND A POWER CORD. THE ISSUE WAS DISCOVERED DURING A SERVICE CHECK AND THERE WAS NO PATIENT INVOLVEMENT. |  | ONE MEDFUSION PUMP WAS RECEIVED CRACKED AND CHIPPED OUT ON BOTH FRONT CORNERS OF THE TOP CASE, A CRACK ON THE RIGHT PLUNGER CASE, A MISSING BATTERY DOOR AND BATTERY. THE EVENT HISTORY LOG WAS REVIEWED AND THERE WAS A BATTERY COMMUNICATION TIMEOUT. THE PUMP WAS UNABLE TO BE TESTED SINCE THE BATTERY WAS NOT RETURNED WITH THE PUMP. THE INTERCONNECT BOARD WAS FOUND WITH CONTAMINATION AND THAT WAS THE CAUSE OF THE REPORTED ISSUE. THE ISSUE WAS USER INDUCED. THE INTERCONNECT BOARD WAS REPLACED AND A BATTERY WAS ADDED TO RESOLVE THE ISSUE. | PUMP, INFUSION | SMITHS MEDICAL ASD, INC. | 4000 |  | FRN |  | 20200413 | Device was returned to manufacturer | I | Pump, Infusion
Malfunction | 20200423 | ['Device Alarm System', 'Obstruction of Flow', 'Battery Problem'] | INFORMATION RECEIVED BY MEDTRONIC INDICATED THAT THE INSULIN PUMP WAS BEEPING FOR NO REASON AND HAD TO RESET OFTEN FOR NO REASON. THE CUSTOMER ALSO STATED THAT THE INSULIN PUMP HAD UNEXPLAINED NO DELIVERY ALARMS AND BATTERY LIFE HAD DIMINISHED. NO HARM REQUIRING MEDICAL INTERVENTION WAS REPORTED. THE INSULIN PUMP WILL NOT BE RETURNED FOR ANALYSIS. |  | (B)(4). CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. | ARTIFICIAL PANCREAS DEVICE SYSTEM, THRESHOLD SUSPEND | MEDTRONIC PUERTO RICO OPERATIONS CO. | MMT-751NAS | A3751NASJ | OZO |  |  | No | LAY USER/PATIENT | Automated Insulin Dosing , Threshold Suspend
Malfunction |  | ['Battery Problem'] | IT WAS REPORTED TO PHILIPS THAT THE BATTERY (DOM 19014-0065-P) FOR THE DEVICE WAS NO LONGER CHARGING. THERE WAS NO PATIENT INVOLVEMENT. THE LI-ION BATTERY PACK WAS RETURNED TO THE FAILURE ANALYSIS LAB FOR EVALUATION. THE BATTERY WAS VISUALLY INSPECTED FOR ANY MECHANICAL DAMAGE AND/OR CONTAMINATION, AND NO ANOMALIES WERE FOUND. AN INITIAL BATTERY CAPACITY TEST RESULTED IN NONE OF THE LED INDICATORS ILLUMINATING, INDICATING A DEPLETED OR FAULTY BATTERY. UPON EVALUATING THE RETURNED BATTERY, IT WAS FOUND THAT THE BATTERY WAS ABLE TO CHARGE IN BOTH THE MRX HEARTSTART AND CADEX CHARGER. FURTHERMORE, IT WAS VERIFIED DURING FUNCTIONAL TESTING THAT MANUAL SHOCKS COULD BE SUCCESSFULLY ADMINISTERED WHEN THE BATTERY WAS INSTALLED IN A DEVICE. HOWEVER, ALTHOUGH THE COMPONENT WAS ABLE TO PROPERLY CHARGE, THE BATTERY MODE CF FLAG WAS FOUND TO BE SET UPON RECEIPT. DUE TO THE ABNORMALITY SURROUNDING THE GAS GAUGE REGISTER STATUS FLAG, THE COMPONENT WAS DETERMINED TO BE POTENTIALLY FAULTY AND WAS SCHEDULED TO BE RETURNED TO THE VENDOR. UPON RESPONSE FROM THE VENDOR, IT WAS CLARIFIED THAT A SET BATTERY MODE CF FLAG IS NOT INDICATIVE OF A COMPONENT MALFUNCTION, BUT INSTEAD IS AN INDICATION THAT THE BATTERY NEEDS TO BE CALIBRATED TO OBTAIN A MORE ACCURATE READING. THE BATTERY WAS SUCCESSFULLY CALIBRATED AND THE FLAG RETURNED TO NORMAL. AS SUCH, THERE WAS NO SIGN OF A COMPONENT FAILURE AND THE BATTERY WAS DEEMED TO BE FULLY FUNCTION. NO FURTHER EVALUATION WAS REQUESTED. | IT WAS REPORTED TO PHILIPS THAT THE BATTERY (DOM 19014-0065-P) FOR THE DEVICE WAS NO LONGER CHARGING. THERE WAS NO PATIENT INVOLVEMENT. | SUBMISSION OF A REPORT DOES NOT CONSTITUTE AN ADMISSION THAT MEDICAL PERSONNEL, USER FACILITY, IMPORTER, DISTRIBUTOR, MANUFACTURER, OR PRODUCT CAUSED OR CONTRIBUTED TO THE EVENT. | DEFIBRILLATOR | PHILIPS MEDICAL SYSTEMS | M3536A |  | MKJ |  | 20201006 | Device was returned to manufacturer | HEALTH PROFESSIONAL | Automated External Defibrillators (Non-Wearable)
Malfunction | 20200309 | ['Device Alarm System', 'Battery Problem'] | INFORMATION WAS RECEIVED INDICATING THAT SMITHS MEDICAL MEDFUSION 4000 PUMP HAS MALFUNCTIONED. IT WAS REPORTED THAT THE PUMP WAS BEING USED INFUSE MEDICATION INTERMITTENTLY WHEN IT GAVE A NEAR EMPTY ALERT. THE REGISTERED NURSE WENT TO PUT FLUSH ON MEDICATION AND GOT A BIOMED ERROR. THERE WERE NO ADVERSE EVENTS REPORTED. |  |  | PUMP, INFUSION |  |  |  | FRN |  |  | No | HEALTH PROFESSIONAL | Pump, Infusion
Malfunction | 20200101 | ['Battery Problem', 'Insufficient Information'] | ADDITIONAL INFORMATION WAS RECEIVED AND PATIENT STATED THAT THE PROBLEM WITH BATTERY NOT KEEPING CHARGE HAVE NOT BEEN RESOLVED. NO FURTHER COMPLICATIONS NOTED OR ANTICIPATED | INFORMATION WAS RECEIVED FROM A CONSUMER (CON) REGARDING A PATIENT WHO WAS IMPLANTED WITH A NEUROSTIMULATOR (INS) FOR COMPLEX REGULAR PAIN SYNDROME TYPE II. IT WAS REPORTED THAT THE PATIENT WAS HAVING PROBLEMS AND ISSUES WITH THE BATTERY/IMPLANT ITSELF. NO SYMPTOMS OR FURTHER COMPLICATIONS WERE REPORTED OR ANTICIPATED. | MEDTRONIC IS SUBMITTING THIS REPORT TO COMPLY WITH FDA REPORTING REGULATIONS UNDER 21 CFR PARTS 4 AND 803. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY MEDTRONIC, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. MEDTRONIC HAS MADE REASONABLE EFFORTS TO OBTAIN MORE COMPLETE INFORMATION AND HAS PROVIDED AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE AN ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY ¿DEFECTS¿ OR HAS ¿MALFUNCTIONED¿. THESE WORDS ARE INCLUDED IN THE FDA 3500A FORM AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REGULATORY REPORTING. MEDTRONIC OBJECTS TO THE USE OF THESE WORDS AND OTHERS LIKE THEM BECAUSE OF THE LACK OF DEFINITION AND THE CONNOTATIONS IMPLIED BY THESE TERMS. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. A GOOD FAITH EFFORT WILL BE MADE TO OBTAIN THE APPLICABLE INFORMATION RELEVANT TO THE REPORT. IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. | STIMULATOR, SPINAL-CORD, TOTALLY IMPLANTED FOR PAIN RELIEF | MEDTRONIC MED REL MEDTRONIC PUERTO RICO | 37714 |  | LGW | 20141014 |  | No | HEALTH PROFESSIONAL | Stimulator, Spinal-Cord, Totally Implanted For Pain Relief
Malfunction | 20200220 | ['Application Program Problem', 'Battery Problem'] | IT WAS REPORTED THAT THE SMITHS MEDICAL CADD MS3 PUMP HAD LOST PROGRAMMING DUE TO NO BATTERY. REPORTER STATED THAT IT WAS UNKNOWN IF THE REPORTED EVENT OCCURRED WHILE IN USE WITH A PATIENT. |  | DEVICE EVALUATION: ONE CADD MS3 PUMP WAS RETURNED FOR ANALYSIS DUE TO LOSS IN PROGRAMMING. FUNCTIONAL AND VISUAL TESTING WAS PERFORMED. THE ISSUE WAS NO DUPLICATED. TESTING WAS PERFORMED AND THE DEVICE DID NOT LOSE ANY PROGRAM. THE DEVICE RAN THE PROGRAM FOR AN EXTENDED PERIOD OF TIME WITH NO ISSUES OCCURRING. THEREFORE, NO PROBLEM FOUND WITH THE ISSUE. | PUMP, INFUSION | ST PAUL | 7400 |  | FRN |  | 20200406 | Device was returned to manufacturer | I | Pump, Infusion
Malfunction | 20200302 | ['Application Program Problem', 'Battery Problem'] | IT WAS REPORTED THAT THE SMITHS MEDICAL CADD MS3 PUMP HAD LOST PROGRAMMING DUE TO NO BATTERY. REPORTER STATED THAT IT WAS UNKNOWN IF THE REPORTED EVENT OCCURRED WHILE IN USE WITH A PATIENT. |  | DEVICE EVALUATION: ONE CADD MS3 PUMP WAS RETURNED FOR ANALYSIS DUE TO LOSS IN PROGRAMMING. FUNCTIONAL AND VISUAL TESTING WAS PERFORMED. THE ISSUE WAS NO DUPLICATED. TESTING WAS PERFORMED AND THE DEVICE DID NOT LOSE ANY PROGRAM. THE DEVICE RAN THE PROGRAM FOR AN EXTENDED PERIOD OF TIME WITH NO ISSUES OCCURRING. THEREFORE, NO PROBLEM FOUND WITH THE ISSUE. | PUMP, INFUSION | ST PAUL | 7400 |  | FRN |  | 20200406 | Device was returned to manufacturer | I | Pump, Infusion
Malfunction | 20200324 | ['Application Program Problem', 'Battery Problem'] | IT WAS REPORTED THAT THE SMITHS MEDICAL CADD MS3 PUMP HAD LOST PROGRAMMING DUE TO NOT CHARGING THE BATTERY. REPORTER STATED THAT IT WAS UNKNOWN IF THE REPORTED EVENT OCCURRED WHILE IN USE WITH A PATIENT. |  | DEVICE EVALUATION: ONE CADD MS3 PUMP WAS RETURNED FOR ANALYSIS DUE TO LOSS IN PROGRAMMING. FUNCTIONAL AND VISUAL TESTING WAS PERFORMED. THE ISSUE WAS NO DUPLICATED. TESTING WAS PERFORMED AND THE DEVICE DID NOT LOSE ANY PROGRAM. THE DEVICE RAN THE PROGRAM FOR AN EXTENDED PERIOD OF TIME WITH NO ISSUES OCCURRING. THEREFORE, NO PROBLEM FOUND WITH THE ISSUE. | PUMP, INFUSION | ST PAUL | 7400 |  | FRN |  | 20200406 | Device was returned to manufacturer | I | Pump, Infusion
Malfunction | 20200317 | ['Application Program Problem', 'Battery Problem'] | IT WAS REPORTED THAT THE SMITHS MEDICAL CADD MS3 PUMP HAD LOST PROGRAMMING DUE TO NO BATTERY CHANGE. REPORTER STATED THAT THERE WAS NO PATIENT INVOLVEMENT IN THE REPORTED EVENT. |  | DEVICE EVALUATION: ONE CADD MS3 PUMP WAS RETURNED FOR ANALYSIS DUE TO LOSS IN PROGRAMMING. FUNCTIONAL AND VISUAL TESTING WAS PERFORMED. THE ISSUE WAS NO DUPLICATED. TESTING WAS PERFORMED AND THE DEVICE DID NOT LOSE ANY PROGRAM. THE DEVICE RAN THE PROGRAM FOR AN EXTENDED PERIOD OF TIME WITH NO ISSUES OCCURRING. THEREFORE, NO PROBLEM FOUND WITH THE ISSUE. | PUMP, INFUSION | ST PAUL | 7400 |  | FRN |  | 20200406 | Device was returned to manufacturer | I | Pump, Infusion
Malfunction | 20200303 | ['Application Program Problem', 'Battery Problem'] | IT WAS REPORTED THAT THE SMITHS MEDICAL CADD MS3 PUMP HAD LOST PROGRAMMING DUE TO BATTERY CHANGE. REPORTER STATED THAT IT WAS UNKNOWN IF THE REPORTED EVENT OCCURRED WHILE IN USE WITH A PATIENT. |  | DEVICE EVALUATION: ONE CADD MS3 PUMP WAS RETURNED FOR ANALYSIS DUE TO LOSS IN PROGRAMMING. FUNCTIONAL AND VISUAL TESTING WAS PERFORMED. THE ISSUE WAS NOT DUPLICATED. TESTING WAS PERFORMED AND THE DEVICE DID NOT LOSE ANY PROGRAM. THE DEVICE RAN THE PROGRAM FOR AN EXTENDED PERIOD OF TIME WITH NO ISSUES OCCURRING. THEREFORE, NO PROBLEM FOUND WITH THE ISSUE. | PUMP, INFUSION | ST PAUL | 7400 |  | FRN |  | 20200406 | Device was returned to manufacturer | I | Pump, Infusion
Malfunction |  | ['Battery Problem', 'Charging Problem'] | ADDITIONAL INFORMATION WAS RECEIVED AND IT WAS REPORTED THAT THE CAUSE OF HAVING TO RECHARGE MORE FREQUENTLY IS THE DEVICE IS ALMOST EIGHT YEARS OLD. THE PATIENT IS GOING TO GET A NEW "SLIM" INS IN (B)(6) 2020. THE PATIENT LOWERED THE INS DOWN TO DECREASE RECHARGE FREQUENCY. THE PATIENT PROVIDED A WEIGHT OF 172 POUNDS. NO FURTHER COMPLICATIONS WERE REPORTED. NO PATIENT SYMPTOMS WERE REPORTED. | INFORMATION WAS RECEIVED FROM A PATIENT REGARDING THEIR IMPLANTABLE NEUROSTIMULATOR (INS). INFORMATION WAS REPORTED THAT THE PATIENT STATED THAT HE NEEDS TO HAVE HIS INS REPLACED BECAUSE HE HAS TO CHARGE MORE FREQUENTLY. THE PATIENT FURTHER CLARIFIED THAT HE NEEDS TO CHARGE HIS INS EVERY WEEK NOW AND DUE TO THIS HE HAS BEEN LETTING HIS INS GO DEAD MORE. THE PATIENT STATED THIS STARTED ABOUT 4 OR 5 MONTHS AGO. THE PATIENT STATED THEIR HEALTHCARE PROVIDER (HCP) TOLD THE PATIENT THAT HE WOULD NEED TO GO BACK TO (B)(6) TO HAVE THE INS REPLACED SINCE THAT IS WHERE THE INS WAS IMPLANTED. NO FURTHER COMPLICATIONS WERE REPORTED. NO PATIENT SYMPTOMS WERE REPORTED. | MEDTRONIC IS SUBMITTING THIS REPORT TO COMPLY WITH FDA REPORTING REGULATIONS UNDER 21 CFR PARTS 4 AND 803. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY MEDTRONIC, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. MEDTRONIC HAS MADE REASONABLE EFFORTS TO OBTAIN MORE COMPLETE INFORMATION AND HAS PROVIDED AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE AN ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY ¿DEFECTS¿ OR HAS ¿MALFUNCTIONED¿. THESE WORDS ARE INCLUDED IN THE FDA 3500A FORM AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REGULATORY REPORTING. MEDTRONIC OBJECTS TO THE USE OF THESE WORDS AND OTHERS LIKE THEM BECAUSE OF THE LACK OF DEFINITION AND THE CONNOTATIONS IMPLIED BY THESE TERMS. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. A GOOD FAITH EFFORT WILL BE MADE TO OBTAIN THE APPLICABLE INFORMATION RELEVANT TO THE REPORT. IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. | STIMULATOR, SPINAL-CORD, TOTALLY IMPLANTED FOR PAIN RELIEF | MEDTRONIC MED REL MEDTRONIC PUERTO RICO | 37713 |  | LGW | 20120528 |  | No | HEALTH PROFESSIONAL | Stimulator, Spinal-Cord, Totally Implanted For Pain Relief
Malfunction | 20200423 | ['Device Difficult to Program or Calibrate', 'Battery Problem', 'Connection Problem', 'Material Integrity Problem', 'Power Problem', 'Patient Device Interaction Problem'] | CUSTOMER REPORTED VIA PHONE CALL THE INSULIN PUMP BUTTON WAS NOT RESPONSIVE, RESERVOIR WAS NOT SECURE, SCRATCHES ON THE SCREEN MADE IT HARD TO READ, BATTERY CAP WAS CRACKED, PUMP HAS TO BE REWOUND MORE THAN ONCE, AND BATTERY LIFE WAS DIMINISHED. CUSTOMER BLOOD GLUCOSE READING WAS UNKNOWN. TROUBLESHOOTING WAS NOT PERFORMED. THE INSULIN PUMP WILL NOT BE RETURNING FOR ANALYSIS. |  | (B)(4). CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. | PUMP, INFUSION, INSULIN, TO BE USED WITH INVASIVE GLUCOSE SENSOR | MEDTRONIC PUERTO RICO OPERATIONS CO. | MMT-723NAS | B5723NASJ | OYC |  |  | No | LAY USER/PATIENT | Pump, Infusion, Insulin, To Be Used With Invasive Glucose Sensor
Malfunction |  | ['Device Alarm System', 'Battery Problem'] | INFORMATION WAS RECEIVED INDICATING THAT A SMITHS MEDICAL MEDFUSION PUMP HAD A BATTERY ERROR. THE ISSUE WAS DISCOVERED DURING A SERVICE CHECK AND THERE WAS NO PATIENT INVOLVEMENT. | INFORMATION WAS RECEIVED INDICATING THAT A SMITHS MEDICAL MEDFUSION PUMP HAD A BATTERY ERROR. THE ISSUE WAS DISCOVERED DURING A SERVICE CHECK AND THERE WAS NO PATIENT INVOLVEMENT. | RETURNED DEVICE WAS RECEIVED WITH MISSING BATTERY. BROKEN EAR CLIP, TOP CASE CRACKED ON CORNERS, TUBING GUIDES AND HANDLE, CRACKED BOTTOM CASE BY L-BRACKET, AND RIGHT PLUNGER CASE.. EVIDENCE OF REPORTED PROBLEM IN EVENT LOG WAS FOUND. DURING THE EVALUATION OF THE DEVICE BATTERY COMMUNICATION TIMEOUT AT POWER UP. BATTERY VALUES ALL ZERO SINCE BATTERY IS MISSING. THE CUSTOMER REPORTED CONDITION WAS CONFIRMED. PROBLEM SOURCE WAS TRACED TO USER INTERFACE. A MANUFACTURING DHR REVIEW WAS NOT PERFORMED BECAUSE THE RESULTS OF THE COMPLAINT INVESTIGATION DO NOT INDICATE A PROBLEM WITH THE MANUFACTURE OF THE DEVICE. NO CAUSES OR POTENTIAL CAUSES OF THE CUSTOMER'S REPORTED PROBLEM WERE FOUND DURING THE REVIEW OF SERVICE AND REPAIR RECORDS. | PUMP, INFUSION | SMITHS MEDICAL ASD, INC. | 4000 |  | FRN |  | 20200413 | Device was returned to manufacturer | I | Pump, Infusion
Injury | 20200412 | ['Battery Problem'] | CUSTOMER REPORTED BEING UNABLE TO TEST USING HIS ADC FREESTYLE LIBRE READER DUE TO THE READER NOT TURNING ON AND ONLY TURNING ON WHEN CONNECTED TO A DATA CABLE. CUSTOMER EXPERIENCED CONFUSION AND SELF-PRESENTED TO A HEALTHCARE PROFESSIONAL WHO OBTAINED A READING OF "GREATER THAN 600" ON AN HCP DEVICE. THE CUSTOMER WAS DIAGNOSED WITH HYPERGLYCEMIA AND ADMINISTERED INSULIN (DOSE/TYPE UNKNOWN) AS TREATMENT. THERE WAS NO REPORT OF DEATH OR PERMANENT INJURY ASSOCIATED WITH THIS EVENT. |  | AT THIS TIME PRODUCT HAS NOT YET BEEN RETURNED. AN EXTENDED INVESTIGATION HAS BEEN PERFORMED FOR THE REPORTED COMPLAINT AND THERE WAS NO INDICATION THAT THE PRODUCT DID NOT MEET SPECIFICATION. DHRS FOR THE FREESTYLE LIBRE READER WAS REVIEWED AND THE DHRS SHOWED THE FREESTYLE LIBRE READER PASSED ALL TESTS PRIOR TO RELEASE. IF THE PRODUCT IS RETURNED, A PHYSICAL INVESTIGATION WILL BE PERFORMED AND A FOLLOW-UP REPORT SUBMITTED. | FLASH GLUCOSE MONITORING SYSTEM | ABBOTT DIABETES CARE INC | 71937-01 |  | PZE |  |  | No | LAY USER/PATIENT | Sensor, Glucose, Invasive, Non-Adjunctive, Factory-Calibrated, User-Initiated
Malfunction | 20191029 | ['Battery Problem', 'Charging Problem'] | INFORMATION WAS RECEIVED FROM A CONSUMER REGARDING A PATIENT WHO WAS IMPLANTED WITH A NEUROSTIMULATOR (INS). IT WAS REPORTED THAT THE PATIENT'S NEW IMPLANTS TAKE A LOT LONGER TO KEEP CHARGE THAN THE OLD IMPLANTS. THE PATIENT STATED THAT THEY CHARGED THE OLD INS ONCE A MONTH. NOW THEY HAVE TO CHARGE EVERY DAY BECAUSE NEITHER INS HAS KEPT A CHARGE SINCE IMPLANT. THE PATIENT IS ALWAYS ACTIVE AND HAS THE THERAPY ON ALL THE TIME. PATIENT SERVICES REVIEWED SETTING AND PROGRAMMING AND THE PATIENT WOULD DISCUSS WITH A HEALTHCARE PROVIDER (HCP) AND MANUFACTURER REPRESENTATIVE (REP). NO SYMPTOMS WERE REPORTED. NO FURTHER COMPLICATIONS WERE REPORTED OR ANTICIPATED. | ADDITIONAL INFORMATION WAS RECEIVED FROM A CONSUMER REGARDING THE PATIENT. IT WAS REPORTED THAT THERE WERE NO CIRCUMSTANCES THAT LED TO HAVING TO CHARGE EVERY DAY. THE PATIENT INDICATED THAT SHE LIKED HER OLD IMPLANTABLE NEUROSTIMULATOR BETTER. THE ISSUE WITH HAVING TO CHARGE THE IMPLANTABLE NEUROSTIMULATOR EVERY DAY HAS BEEN RESOLVED. THERE WERE NO PATIENT SYMPTOMS OR COMPLICATIONS REPORTED. | IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. | STIMULATOR, SPINAL-CORD, TOTALLY IMPLANTED FOR PAIN RELIEF | MEDTRONIC PUERTO RICO OPERATIONS CO. | 97715 |  | LGW | 20200814 |  | No | HEALTH PROFESSIONAL | Stimulator, Spinal-Cord, Totally Implanted For Pain Relief
Malfunction | 20191029 | ['Battery Problem', 'Charging Problem'] | INFORMATION WAS RECEIVED FROM A CONSUMER REGARDING A PATIENT WHO WAS IMPLANTED WITH A NEUROSTIMULATOR (INS). IT WAS REPORTED THAT THE PATIENT'S NEW IMPLANTS TAKE A LOT LONGER TO KEEP CHARGE THAN THE OLD IMPLANTS. THE PATIENT STATED THAT THEY CHARGED THE OLD INS ONCE A MONTH. NOW THEY HAVE TO CHARGE EVERY DAY BECAUSE NEITHER INS HAS KEPT A CHARGE SINCE IMPLANT. THE PATIENT IS ALWAYS ACTIVE AND HAS THE THERAPY ON ALL THE TIME. PATIENT SERVICES REVIEWED SETTING AND PROGRAMMING AND THE PATIENT WOULD DISCUSS WITH A HEALTHCARE PROVIDER (HCP) AND MANUFACTURER REPRESENTATIVE (REP). NO SYMPTOMS WERE REPORTED. NO FURTHER COMPLICATIONS WERE REPORTED OR ANTICIPATED. REFER TO MANUFACTURER REPORT #3004209178-2020-07791 | ADDITIONAL INFORMATION WAS RECEIVED FROM A CONSUMER REGARDING THE PATIENT. IT WAS REPORTED THAT THERE WERE NO CIRCUMSTANCES THAT LED TO HAVING TO CHARGE EVERY DAY. THE PATIENT INDICATED THAT SHE LIKED HER OLD IMPLANTABLE NEUROSTIMULATOR BETTER. THE ISSUE WITH HAVING TO CHARGE THE IMPLANTABLE NEUROSTIMULATOR EVERY DAY HAS BEEN RESOLVED. THERE WERE NO PATIENT SYMPTOMS OR COMPLICATIONS REPORTED. | IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. | STIMULATOR, SPINAL-CORD, TOTALLY IMPLANTED FOR PAIN RELIEF | MEDTRONIC PUERTO RICO OPERATIONS CO. | 97715 |  | LGW | 20200814 |  | No | HEALTH PROFESSIONAL | Stimulator, Spinal-Cord, Totally Implanted For Pain Relief
Injury |  | ['Battery Problem'] | INFORMATION WAS RECEIVED FROM A HEALTHCARE PROFESSIONAL (HCP) REGARDING A PATIENT WHO WAS IMPLANTED WITH AN IMPLANTABLE NEUROSTIMULATOR (INS). IT WAS REPORTED THAT THE PATIENT HAD DEEP BRAIN STIMULATION (DBS) BATTERY DEPLETION AND LOSS OF TREATMENT, REQUIRING A BATTERY REPLACEMENT. THE REPLACEMENT DATE WAS (B)(6) 2020. NO FURTHER COMPLICATIONS WERE REPORTED/ANTICIPATED. |  | IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. | STIMULATOR, ELECTRICAL, IMPLANTED, FOR PARKINSONIAN TREMOR | MEDTRONIC MED REL MEDTRONIC PUERTO RICO | 37612 |  | MHY | 20110428 |  | No | HEALTH PROFESSIONAL | Stimulator, Electrical, Implanted, For Parkinsonian Tremor
Malfunction |  | ['Inappropriate/Inadequate Shock/Stimulation', 'Improper or Incorrect Procedure or Method', 'Battery Problem', 'Insufficient Information'] | ADDITIONAL INFORMATION WAS RECEIVED FROM THE PATIENT REPORTING THAT THEY WERE TOLD TO CALL BACK TO PROVIDE CLARIFICATION ON WHAT ¿WASN¿T WORKING¿. THE PATIENT STATED THAT THEIR INS WAS IN OVER DISCHARGE. THEY CLARIFIED THAT THEY HAD FORGOT TO CHARGE IT AND THAT THIS WAS THIRD OVER DISCHARGE, SO NOW IT WAS JUST IMPLANTED AND DEAD AS THEY WERE TOLD IT COULD NO LONGER BE JUMP STARTED. THE PATIENT STATED THAT THE DATE WAS ABOUT THREE YEARS AGO NOW. THEY STATED THAT THEY WANTED TO HAVE THE DEVICE REMOVED SO THAT THEY COULD HAVE MRIS. THE PATIENT ALSO INQUIRED IF THERE WERE NEWER DEVICES THAT DID NOT ALLOW OVER DISCHARGES TO HAPPEN. INFORMATION WAS REVIEWED WITH THE PATIENT. A LIST OF PHYSICIAN¿S WAS SENT TO THE PATIENT AS THEY INDICATED THAT THEY DID NOT HAVE A MANAGING HCP. NO FURTHER COMPLICATIONS WERE REPORTED/ANTICIPATED. | INFORMATION WAS RECEIVED FROM A HEALTHCARE PROVIDER (HCP) REGARDING A PATIENT WHO WAS IMPLANTED WITH A NEUROSTIMULATOR (INS). THE CALLER REQUESTED MRI COMPATIBILITY INFORMATION, AND TECHNICAL SERVICES REVIEWED THAT AN MRI OF THE PATIENT'S KNEE WAS NOT ADVISED. IT WAS REPORTED THAT THE INS WAS OFF AND "DOESN'T WORK". THE HCP WAS UNSURE WHAT THE PATIENT MEANT BY "DOESN'T WORK". IT WAS RECOMMENDED THAT THE PATIENT FOLLOW UP WITH THEIR MANAGING PHYSICIAN. NO SYMPTOMS WERE REPORTED. NO FURTHER COMPLICATIONS WERE REPORTED OR ANTICIPATED. ADDITIONAL INFORMATION RECEIVED FROM THE CONSUMER REPORTED THAT NO FURTHER STEPS OR ACTIONS WERE TAKEN AS THEY HAD BEEN TOLD THE IMPLANT COULDN¿T BE CHARGED AGAIN. THE PATIENT STATED THAT AN EXPLANT DATE HAD YET TO BE ANNOUNCED. PT REPORTED THAT THE DEVICE HADN'T WORKED FOR A VERY LONG TIME AND THAT THEY WERE ABOUT TO HAVE A KNEE REPLACEMENT SURGERY AND THEY WOULD LIKE TO HAVE THE DEVICE REMOVED. PATIENT SAID THAT THE IMPLANTING HEALTHCARE PROVIDER WAS NO LONGER PRACTICING WHERE THEY LIVED AND THAT THE HEALTHCARE PROVIDER WHO REPLACED THE IMPLANTING HEALTHCARE PROVIDER WOULDN'T DO THE DEVICE REMOVAL. WHEN ASKED FOR THE EVENT DATE, PATIENT SAID THAT IT WAS LIKE FIVE YEARS AFTER THE DEVICE WAS IMPLANTED. ADDITIONAL INFORMATION WAS RECEIVED FROM THE PATIENT. THEY REPORTED THAT THEIR INS HAD NOT WORKED FOR A NUMBER OF YEARS. THEY COULDN'T FIND A HEALTHCARE PROVIDER (HCP) TO REMOVE THEIR DEVICE BECAUSE IT BEING LOCATED TO THEIR KNEE AND NOT THEIR SPINE. RECENTLY, THE DEVICE WAS GIVING THEM ALL TYPES OF PAIN AND THEY NEEDED SOMEONE TO TAKE IT OUT SINCE THE PAIN WAS CONSTANT. EVERY ONCE IN A WHILE IT WILL MALFUNCTION SENDING A SHOCK TO THE LEG OR UP THE BUTT TO TRIGGER THEIR SCIATICA; PT HAD TO GO TO THE EMERGENCY ROOM FOR IT. THE PATIENT WOULD CONTINUE TO CONTACT DOCTORS FROM THE PHYSICIAN LISTINGS THEY WERE SENT. | MEDTRONIC IS SUBMITTING THIS REPORT TO COMPLY WITH FDA REPORTING REGULATIONS UNDER 21 CFR PARTS 4 AND 803. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY MEDTRONIC, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. MEDTRONIC HAS MADE REASONABLE EFFORTS TO OBTAIN MORE COMPLETE INFORMATION AND HAS PROVIDED AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE AN ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY ¿DEFECTS¿ OR HAS ¿MALFUNCTIONED¿. THESE WORDS ARE INCLUDED IN THE FDA 3500A FORM AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REGULATORY REPORTING. MEDTRONIC OBJECTS TO THE USE OF THESE WORDS AND OTHERS LIKE THEM BECAUSE OF THE LACK OF DEFINITION AND THE CONNOTATIONS IMPLIED BY THESE TERMS. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. A GOOD FAITH EFFORT WILL BE MADE TO OBTAIN THE APPLICABLE INFORMATION RELEVANT TO THE REPORT. IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. MEDTRONIC IS SUBMITTING THIS REPORT TO COMPLY WITH FDA REPORTING REGULATIONS UNDER 21 CFR PARTS 4 AND 803. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY MEDTRONIC, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. MEDTRONIC HAS MADE REASONABLE EFFORTS TO OBTAIN MORE COMPLETE INFORMATION AND HAS PROVIDED AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE AN ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY ¿DEFECTS¿ OR HAS ¿MALFUNCTIONED¿. THESE WORDS ARE INCLUDED IN THE FDA 3500A FORM AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REGULATORY REPORTING. MEDTRONIC OBJECTS TO THE USE OF THESE WORDS AND OTHERS LIKE THEM BECAUSE OF THE LACK OF DEFINITION AND THE CONNOTATIONS IMPLIED BY THESE TERMS. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. A GOOD FAITH EFFORT WILL BE MADE TO OBTAIN THE APPLICABLE INFORMATION RELEVANT TO THE REPORT. IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. | STIMULATOR, SPINAL-CORD, TOTALLY IMPLANTED FOR | MEDTRONIC MED REL MEDTRONIC PUERTO RICO | 37714 |  | LGW | 20130814 |  | No | HEALTH PROFESSIONAL | Stimulator, Spinal-Cord, Totally Implanted For Pain Relief
Malfunction |  | ['Failure to Interrogate', 'Migration or Expulsion of Device', 'Inappropriate/Inadequate Shock/Stimulation', 'Battery Problem', 'Charging Problem', 'Communication or Transmission Problem'] | ADDITIONAL INFORMATION RECEIVED FROM THE PATIENT INDICATED THAT THEY GOT A BAD ELECTRIC SHOCK WHEN THEY PUT ON THE RECHARGER. THEY NOTED THAT BEFORE THIS ISSUE, THEY WOULD GET SMALL ZAPS EVERY NOW AND THEN FOR THE PAST COUPLE OF MONTHS. THEY MENTIONED THAT NOW THEY GOT ALL 8 COUPLING BARS, WHICH BEGAN AFTER THE BAD ELECTRIC SHOCK. THEY STATED THAT THEY GET SHOCKED WHEN TURNING ON STIMULATION. THE PATIENT ADDED THAT THEY TRY MOVING THE ANTENNA ¿LIKE 200 TIMES.¿ THEY ALSO STATED THAT THEY THINK THE INS MOVED FROM THEIR SIDE TO THEIR LEFT HIP, AND ALMOST TO THEIR SPINE. THE PATIENT NOTED THAT THEY DON¿T FEEL IT HAS MOVED ANY FURTHER FOR MONTHS. THEY MENTIONED THAT THEY RECHARGER HAS NOT BEEN CONNECTING WITH THE INS. THEY STATED THAT THEY ARE HAVING PAIN ALL THROUGH THEIR BACK AND LEFT HIP. THEY STATED THAT THEY CAN'T EVEN WALK ACROSS THE PARKING LOT WITHOUT BENDING OVER AND CRYING. THE PATIENT ADDED THAT THEY FEEL LIKE THEY ARE LOCKED UP, STIFF, AND CAN¿T MOVE. THEY MENTIONED THAT THEY CAN¿T SIT LONG AND GET NUMBNESS THROUGHOUT THEIR HAND, ARMS, AND LEGS. THE PATIENT INDICATED THAT THIS STARTED 2-3 MONTHS AGO. NO FURTHER COMPLICATIONS WERE REPORTED OR ANTICIPATED. | INFORMATION WAS RECEIVED FROM THE CONSUMER REGARDING A PATIENT WHO WAS IMPLANTED WITH A NEUROSTIMULATOR FOR SPINAL CORD STIMULATION. IT WAS REPORTED THAT THE INS WAS PLACED ON THE LEFT SIDE OF HER BODY AND THE INS MOVED AND WAS GOING INTO HER SPINE. PATIENT WAS TEMPTED TO GO TO THE HOSPITAL TO GET A SHOT OF SOMETHING TO HELP WITH THE PAIN BUT SHE WAS AVOIDING GOING BECAUSE OF THE COVID-19. PATIENT STATED IT WAS DIFFICULT FOR HER TO STAND OR WALK. PATIENT STATED THIS HAD BEEN GOING ON FOR A LONG TIME AND SHE HAD BEEN IN PAIN FOR MANY MONTHS. PATIENT ALSO REPORTED SEEING ERI ON HER CONTROLLER. NO FURTHER COMPLICATIONS WERE REPORTED/ANTICIPATED. ADDITIONAL INFORMATION WAS RECEIVED FROM PATIENT. IT WAS REPORTED THAT SHE HAD AN ISSUE WITH RECHARGING THE INS ON AND OFF FOR AWHILE. PATIENT STATED SHE WOULD HAVE FULL COUPLING THEN IT WOULD GO TO ZERO. PATIENT REPORTED INS HAD MOVED. PATIENT STATED SHE HAD EXCRUCIATING PAIN ON THE LEFT SIDE NEAR THE INS. PATIENT STATED HER STIM WAS ON ITS LAST LEG AND NEEDED TO BE REPLACED. PATIENT STATED SHE HAS SEEN ERI FROM SOMETIME. | IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. MEDTRONIC IS SUBMITTING THIS REPORT TO COMPLY WITH FDA REPORTING REGULATIONS UNDER 21 CFR PARTS 4 AND 803. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY MEDTRONIC, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. MEDTRONIC HAS MADE REASONABLE EFFORTS TO OBTAIN MORE COMPLETE INFORMATION AND HAS PROVIDED AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE AN ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY ¿DEFECTS¿ OR HAS ¿MALFUNCTIONED¿. THESE WORDS ARE INCLUDED IN THE FDA 3500A FORM AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REGULATORY REPORTING. MEDTRONIC OBJECTS TO THE USE OF THESE WORDS AND OTHERS LIKE THEM BECAUSE OF THE LACK OF DEFINITION AND THE CONNOTATIONS IMPLIED BY THESE TERMS. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. A GOOD FAITH EFFORT WILL BE MADE TO OBTAIN THE APPLICABLE INFORMATION RELEVANT TO THE REPORT. IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. | STIMULATOR, SPINAL-CORD, TOTALLY IMPLANTED FOR PAIN RELIEF | MEDTRONIC MED REL MEDTRONIC PUERTO RICO | 37714 |  | LGW | 20121228 |  | No | HEALTH PROFESSIONAL | Stimulator, Spinal-Cord, Totally Implanted For Pain Relief
Malfunction |  | ['Device Alarm System', 'Battery Problem'] | ORACLE RO 1058663: BATTERY DEFECTIVE ERROR. ADDITIONAL INFORMATION RECEIVED BY SMITHS MEDICAL ON 09-APR-2020 ATTACHED IN COMPLAINT OBJECT: ALL OF THESE PUMPS WERE FOUND TO BE BROKEN WHILE DOING A SERVICE CHECK ON THEM. NO PATIENT INVOLVEMENT. | INFORMATION WAS RECEIVED INDICATING THAT A SMITHS MEDICAL MEDFUSION PUMP HAD A BATTERY DEFECTIVE ERROR. THE ISSUE WAS DISCOVERED DURING A SERVICE CHECK AND THERE WAS NO PATIENT INVOLVEMENT. | OTHER, OTHER TEXT: RETURNED DEVICE WAS RECEIVED WITH MISSING BATTERY. BROKEN EAR CLIP, TOP CASE CRACKED ON CORNERS, TUBING GUIDES AND HANDLE, CRACKED BOTTOM CASE BY L-BRACKET, AND RIGHT PLUNGER CASE.. EVIDENCE OF REPORTED PROBLEM IN EVENT LOG WAS FOUND. DURING THE EVALUATION OF THE DEVICE AFTER CHECKING ALARM HISTORY, THE BATTERY WAS UNPLUGGED CAUSING THE BATTERY ERRORS IN HISTORY. THE BATTERY WAS FROM 2019. THE CUSTOMER REPORTED CONDITION WAS CONFIRMED. PROBLEM SOURCE WAS TRACED TO USER INTERFACE. A MANUFACTURING DHR REVIEW WAS NOT PERFORMED BECAUSE THE RESULTS OF THE COMPLAINT INVESTIGATION DO NOT INDICATE A PROBLEM WITH THE MANUFACTURE OF THE DEVICE. NO CAUSES OR POTENTIAL CAUSES OF THE CUSTOMER'S REPORTED PROBLEM WERE FOUND DURING THE REVIEW OF SERVICE AND REPAIR RECORDS. | PUMP, INFUSION | SMITHS MEDICAL ASD, INC. | 4000 |  | FRN |  | 20200413 | Device was returned to manufacturer | I | Pump, Infusion
Malfunction |  | ['Battery Problem'] | A CUSTOMER REPORTED THAT THE DEVICE'S BATTERY FAILED. IT WAS REPORTED TO PHILIPS THAT THE ALLEGED FAILURE WAS OBSERVED WHILE THE DEVICE WAS IN USE ON A PATIENT, HOWEVER, NO ADVERSE PATIENT IMPACT WAS REPORTED BY THE CUSTOMER. | A CUSTOMER REPORTED THAT THE DEVICE BATTERY FAILED DURING OPCHECK. ADDITIONAL INFORMATION RECEIVED CLARIFIED THAT THE ISSUE TOOK PLACE DURING TESTING. THERE WAS NO PATIENT INVOLVEMENT. | A FOLLOW-UP REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETED. | ALS DEFIBRILLATOR MONITOR | PHILIPS MEDICAL SYSTEMS | 861290 |  | MKJ |  |  | Yes | HEALTH PROFESSIONAL | Automated External Defibrillators (Non-Wearable)
Malfunction |  | ['Battery Problem'] | THE CUSTOMER REPORTED BATTERY INDICATOR FLASHING. THERE WAS NO PATIENT INVOLVEMENT. THE MANUFACTURER¿S FIELD SERVICE ENGINEER (FSE) CONFIRMED THE REPORTED BATTERY ISSUE. THE MANUFACTURER¿S FIELD SERVICE ENGINEER (FSE) REPLACED THE BATTERY TO ADDRESS THE REPORTED PROBLEM. THE UNIT WAS CHECKED OVERALL, RUN IN TESTED, CLEANED AND FUNCTIONALLY TESTED AND NO ABNORMALITY WAS CONFIRMED. |  | DATE OF EVENT: (B)(6) 2020. DATE OF REPORT: 28APR2020. | VENTILATOR, CONTINUOUS, MINIMAL VENTILATORY SUPPORT,FACILITY USE | RESPIRONICS CALIFORNIA, INC | V60 |  | MNT |  |  | Yes | HEALTH PROFESSIONAL | Ventilator, Continuous, Minimal Ventilatory Support, Facility Use
Malfunction |  | ['Battery Problem'] | THE CUSTOMER REPORTED BATTERY FAILURE. THERE WAS NO PATIENT INVOLVEMENT. THE MANUFACTURER'S FIELD SERVICE ENGINEER (FSE) CONFIRMED THE REPORTED ISSUE. THE FSE REPLACED THE DEFECTIVE BATTERY TO ADDRESS THE REPORTED PROBLEM. THE UNIT SUCCESSFULLY PASSED THE REQUIRED PERFORMANCE VERIFICATION TEST. |  | DATE OF EVENT: (B)(6) 2020. DATE OF REPORT: 28APR2020. | VENTILATOR, CONTINUOUS, MINIMAL VENTILATORY SUPPORT,FACILITY USE | RESPIRONICS CALIFORNIA, INC | V60 |  | MNT |  |  | Yes | HEALTH PROFESSIONAL | Ventilator, Continuous, Minimal Ventilatory Support, Facility Use
Malfunction | 20190806 | ['Battery Problem'] | IT WAS REPORTED THAT THE PUMP BATTERY INTERMITTENTLY DEPLETED QUICKLY. THERE WAS NO REPORTED IMPACT TO THE CUSTOMER¿S BLOOD GLUCOSE LEVEL. THE CUSTOMER CONTINUED TO USE THE PUMP FOR INSULIN THERAPY. |  | THE DEVICE HAS BEEN RECEIVED FOR EVALUATION; HOWEVER, DEVICE EVALUATION IS NOT YET COMPLETE. A SUPPLEMENTAL REPORT WILL BE FILED UPON COMPLETION OF THE EVALUATION. | AUTOMATED INSULIN DOSING, THRESHOLD SUSPEND | TANDEM DIABETES CARE | 1000096 |  | OZO |  | 20200427 | Device was returned to manufacturer | LAY USER/PATIENT | Automated Insulin Dosing , Threshold Suspend
Malfunction | 20200424 | ['Battery Problem', 'Charging Problem', 'Insufficient Information'] | INFORMATION WAS RECEIVED FROM A CONSUMER REGARDING AN IMPLANTABLE NEUROSTIMULATOR (INS) FOR FAILED BACK SURGERY SYNDROME AND SPINAL PAIN. IT WAS REPORTED THAT THE PATIENT¿S INS BATTERY HAS ¿RUN DOWN¿ BECAUSE THEY HAD NOT CHARGED THE INS IN A YEAR. THE PATIENT NOTICED REPOSITION ANTENNA SCREEN ON THE RECHARGER AND WAS NOT ABLE TO CHARGE THE INS. THE PATIENT WAS REDIRECTED TO THEIR HEALTHCARE PROVIDER (HCP). NO SYMPTOMS WERE REPORTED. NO FURTHER COMPLICATIONS WERE REPORTED/ANTICIPATED. |  | IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. | STIMULATOR, SPINAL-CORD, TOTALLY IMPLANTED FOR PAIN RELIEF | MEDTRONIC MED REL MEDTRONIC PUERTO RICO | 97714 |  | LGW | 20141128 |  | No | HEALTH PROFESSIONAL | Stimulator, Spinal-Cord, Totally Implanted For Pain Relief
Malfunction | 20200424 | ['Mechanical Problem', 'Device Difficult to Program or Calibrate', 'Battery Problem'] | INFORMATION RECEIVED BY MEDTRONIC INDICATED THAT THE DOWN ARROW WAS NOT WORKING. CUSTOMER STATED THAT THE INSULIN PUMP WAS LOUDER DURING REWIND. INSULIN PUMP REJECTED NEW BATTERIES. INSULIN SMELLS WHEN RESERVOIR TAKES OUT. NO HARM REQUIRING MEDICAL INTERVENTION WAS REPORTED. THE INSULIN PUMP WILL NOT BE RETURNED FOR ANALYSIS. |  | (B)(4). CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. | ARTIFICIAL PANCREAS DEVICE SYSTEM, THRESHOLD SUSPEND | MEDTRONIC PUERTO RICO OPERATIONS CO. | MMT-751LNAS | A6751LNASJ | OZO |  |  | No | LAY USER/PATIENT | Automated Insulin Dosing , Threshold Suspend
Injury | 20200420 | ['Battery Problem'] | INFORMATION RECEIVED BY MEDTRONIC INDICATED THAT THE PATIENT HAD HIGH BLOOD GLUCOSE LEVELS, DIABETIC KETOACIDOSIS AND HOSPITALIZATION. CUSTOMER'S BLOOD GLUCOSE LEVEL WAS OVER 600 MG/DL. TROUBLESHOOTING WAS PERFORMED. CUSTOMER TREATED THEIR HIGH BLOOD GLUCOSE VIA IV FLUID AND MEDICAL DAILY INSULIN. CUSTOMER ADVISED THE INSULIN PUMP WAS TAKEN OFF OF THE PATIENT BY THE HOSPITAL AND THE BATTERY DIED WHICH RESULTED IN THE INSULIN PUMP BEING OFF. CUSTOMER WAS ADVISED TO CHANGE OUT THEIR ENTIRE INFUSION SET, RESERVOIR, INSULIN AND TREAT THEIR BLOOD GLUCOSE VIA HEALTHCARE PROVIDER'S INSTRUCTIONS. THE INSULIN PUMP WILL NOT BE RETURNED FOR ANALYSIS. |  | (B)(4). CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. DEVICE PASSED THE FUNCTIONAL TEST, INCLUDING THE DISPLACEMENT TEST, REWIND, BASIC OCCLUSION TEST, OCCLUSION TEST, PRIME/A33 TEST, EXCESSIVE NO DELIVERY TEST AND DELIVERY ACCURACY TEST AT 0.08730 INCHES. THE STOP CURRENT AND RUN CURRENT MEASUREMENT TESTS ARE WITHIN SPECIFICATION. DEVICE ALSO PASSED SELF TEST, OFF NO POWER ALARM TEST AND A21 ERROR TEST. DEVICE POWERED UP PROPERLY AFTER THE TEST BATTERY WAS INSTALLED. DEVICE WAS MONITORED FOR 4 DAYS. NO BLANK DISPLAY NOTED. DEVICE UPLOADED PROPERLY USING CARE LINK. DEVICE HAD CRACKED BATTERY TUBE THREADS, CRACKED CASE AT THE RESERVOIR TUBE WINDOW CORNER AND CRACKED RESERVOIR TUBE LIP. THE TEST P-CAP AND RESERVOIR DOES LOCK IN PLACE IN THE RESERVOIR COMPARTMENT. (B)(4). | PUMP, INFUSION, INSULIN, TO BE USED WITH INVASIVE GLUCOSE SENSOR | MEDTRONIC PUERTO RICO OPERATIONS CO. | MMT-523RNAP | A1523RNAPJ | OYC |  | 20200611 | Device was returned to manufacturer | LAY USER/PATIENT | Pump, Infusion, Insulin, To Be Used With Invasive Glucose Sensor
Malfunction | 20200313 | ['Battery Problem'] | INFORMATION WAS RECEIVED INDICATING THAT SMITHS MEDICAL MEDFUSION 4000 PUMP HAS MALFUNCTIONED. IT WAS REPORTED THAT THE PUMP WAS BEING USED INFUSE MEDICATION INTERMITTENTLY WHEN IT STOPPED WHILE INFUSING ANTIBIOTIC (FLAGYL) DESPITE BEING PLUGGED IN. WHEN THE PUMP WAS UNPLUGGED AND IT ALARMED BATTERY LOW. THE REGISTERED NURSE RESPONDED TO THE ALARM AND PLUGGED THE PUMP INTO THE WALL AND THE PUMP STOPPED ALARMING. THE NURSE REMAINED IN THE ROOM WITH THE PATIENT AND A FEW MINUTES LATER THE PUMP BEEPED LOW BATTERY AGAIN. THE ISSUE PERSISTED AND EVENTUALLY THE PUMP TURNED OFF. THERE WERE NO ADVERSE EVENTS REPORTED. |  |  | PUMP, INFUSION | ST PAUL | 4000-0105-51 |  | FRN |  |  | No | HEALTH PROFESSIONAL | Pump, Infusion
Malfunction | 20200313 | ['Battery Problem'] | INFORMATION WAS RECEIVED INDICATING THAT SMITHS MEDICAL MEDFUSION 4000 PUMP HAS MALFUNCTIONED. IT WAS REPORTED THAT THE DEVICE'S BATTERY DEPLETED. IT WAS ALSO INDICATED THAT THE PUMP'S BASE WAS NOT RESPONDING. THERE WERE NO ADVERSE EVENTS REPORTED. |  |  |  | ST PAUL | 4000-0105-51 |  | FRN |  |  | No | HEALTH PROFESSIONAL | Pump, Infusion
Malfunction | 20200424 | ['Device Difficult to Program or Calibrate', 'Obstruction of Flow', 'Battery Problem', 'Material Integrity Problem', 'Patient Device Interaction Problem'] | INFORMATION RECEIVED BY MEDTRONIC INDICATED THAT THE BATTERY LIFE WAS DIMINISHED, HAD TO CHANGE BEFORE THE WEEK WAS UP, BUTTONS WERE RUBBED, BATTERY CAP WAS WORN, BUTTONS WERE HARDER TO PRESS AND HAD UNEXPLAINED NO DELIVERY ALARMS. NO HARM REQUIRING MEDICAL INTERVENTION WAS REPORTED. THE INSULIN PUMP WILL BE RETURNED FOR ANALYSIS. |  | (B)(4). CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. | ARTIFICIAL PANCREAS DEVICE SYSTEM, THRESHOLD SUSPEND | MEDTRONIC PUERTO RICO OPERATIONS CO. | MMT-751NAS | A4751NASJ | OZO |  |  | No | LAY USER/PATIENT | Automated Insulin Dosing , Threshold Suspend
Injury | 20200423 | ['Delayed Charge Time', 'Battery Problem'] | IT WAS FURTHER REPORTED THAT THE DEVICE WAS EXPLANTED AND REPLACED. | IT WAS FURTHER REPORTED THAT THE DEVICE WILL BE EXPLANTED AND REPLACED. IT WAS REPORTED THAT THE PHYSICIAN WAS WORRIED ABOUT RECOMMENDED REPLACEMENT TIME (RRT) INDICATOR WHICH DID NOT TRIGGER IN THE BEGINNING OF (B)(6) WHEN THE 2.62 BATTERY VOLTAGE WAS REACHED AND STILL WAS NOT TRIGGERED IN APRIL DURING NORMAL FOLLOW-UP. IT WAS EXPLAINED THAT THERE NEEDS TO BE THREE CONSECUTIVE MEASUREMENTS IN ORDER TO TRIGGER RRT BUT THE PHYSICIAN FELT THERE IS AN ISSUE WITH THE DEVICE IN ITS CAPACITY TO CORRECTLY DISCRIMINATE THE RESIDUAL LONGEVITY. IT WAS ALSO NOTED THAT THE CHARGE TIME WAS ABOVE FIFTEEN SECONDS WHICH IS TOO LONG. THE ICD REMAINS IN USE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. | MEDTRONIC IS SUBMITTING THIS REPORT TO COMPLY WITH FDA REPORTING REGULATIONS UNDER 21 CFR PARTS 4 AND 803. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY MEDTRONIC, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. MEDTRONIC HAS MADE REASONABLE EFFORTS TO OBTAIN MORE COMPLETE INFORMATION AND HAS PROVIDED AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE AN ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY ¿DEFECTS¿ OR HAS ¿MALFUNCTIONED¿. THESE WORDS ARE INCLUDED IN THE FDA 3500A FORM AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REGULATORY REPORTING. MEDTRONIC OBJECTS TO THE USE OF THESE WORDS AND OTHERS LIKE THEM BECAUSE OF THE LACK OF DEFINITION AND THE CONNOTATIONS IMPLIED BY THESE TERMS. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. A GOOD FAITH EFFORT WILL BE MADE TO OBTAIN THE APPLICABLE INFORMATION RELEVANT TO THE REPORT. IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. PRODUCT EVENT SUMMARY: THE DEVICE WAS NOT RETURNED FOR ANALYSIS. HOWEVER, PERFORMANCE DATA COLLECTED FROM THE DEVICE WAS RECEIVED AND ANALYZED. ANALYSIS OF THE DEVICE MEMORY WAS PERFORMED AND NO ANOMALIES WERE FOUND. IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. | DEFIBRILLATOR, AUTOMATIC IMPLANTABLE CARDIOVERTER | MEDTRONIC EUROPE SARL | D354DRM |  | LWS | 20150128 |  | No | LAY USER/PATIENT | Implantable Cardioverter Defibrillator (Non-Crt)
Malfunction | 20200407 | ['Battery Problem'] | IT WAS REPORTED THAT THE PUMP BATTERY WAS DEPLETING QUICKLY. THE CUSTOMER CONTINUED TO USE THE PUMP FOR INSULIN THERAPY. THE CUSTOMER¿S BLOOD GLUCOSE WAS NOT IMPACTED. |  | NO PRODUCT WAS RETURNED FOR EVALUATION. SHOULD NEW RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. | PUMP, INFUSION, INSULIN, TO BE USED WITH INVASIVE GLUCOSE SENSOR | TANDEM DIABETES CARE | 1002717 |  | OYC |  |  | No | LAY USER/PATIENT | Pump, Infusion, Insulin, To Be Used With Invasive Glucose Sensor
Malfunction |  | ['Failure to Deliver Energy', 'Battery Problem', 'Charging Problem', 'Insufficient Information'] | INFORMATION WAS RECEIVED FROM A CONSUMER CONCERNING PATIENT WITH AN IMPLANTABLE NEUROSTIMULATOR (INS). IT WAS REPORTED THAT THERE WAS A SUSPECTED OVERDISCHARGE. THE PATIENT HAD BACK SURGERY IN (B)(6) 2019 AND HAD INS TURNED OFF. IN (B)(6) 2019, THE PATIENT TRIED CHARGING THE IMPLANT AND IT WOULD NOT CHARGE. THE PATIENT SPOKE TO THE MANUFACTURER REPRESENTATIVE (REP) AND THEY TOLD HER WHAT TO DO TO GET IT CHARGED UP BUT SHE COULD NEVER GET IT. THE PATIENT REPORTED SHE NEEDS AN MRI AND NEEDS TO HAVE THE INS CHARGED SO SHE CAN CHECK MRI COMPATIBILITY. THE PATIENT WAS REDIRECTED TO THEIR HEALTHCARE PROVIDER (HCP). NO PATIENT SYMPTOMS WERE REPORTED. NO FURTHER COMPLICATIONS WERE REPORTED/ANTICIPATED. ADDITIONAL INFORMATION WAS RECEIVED FROM THE PATIENT. IT WAS REPORTED THAT THE STIMULATOR DEVICE HAS A DEAD BATTERY DUE TO NO USE FOLLOWING SURGERY, JUMP START TRIED TWICE AND UNSUCCESSFUL. APPARENTLY LEAD WIRES WERE NEVER CHANGED AND GETTING STIM TO WORK AND GO TO MRI MODE UNSUCCESSFUL. NO NEW PLANS AT THIS TIME. NO FURTHER COMPLICATIONS WERE REPORTED. |  | IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. | STIMULATOR, SPINAL-CORD, TOTALLY IMPLANTED FOR PAIN RELIEF | MEDTRONIC PUERTO RICO OPERATIONS CO. | 97714 |  | LGW | 20161128 |  | No | HEALTH PROFESSIONAL | Stimulator, Spinal-Cord, Totally Implanted For Pain Relief
Malfunction | 20200424 | ['Circuit Failure', 'Battery Problem'] | INFORMATION RECEIVED BY MEDTRONIC INDICATED THAT THE INSULIN PUMP HAD RECEIVED REPLACE BATTERY ALARM FREQUENTLY. THE CUSTOMER STATED THAT THE INSULIN PUMP RECEIVED A PUMP ERROR ALARM TWICE DURING THE SELF-TEST. NO HARM REQUIRING MEDICAL INTERVENTION WAS REPORTED. THE INSULIN PUMP WILL NOT BE RETURNED FOR ANALYSIS. |  | (B)(4). CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. THE INSULIN PUMP INVOLVED IN THIS EVENT IS THE 640G INSULIN INFUSION PUMP, WHICH IS NOT MARKETED IN THE UNITED STATES. HOWEVER, THE DEVICE IS SIMILAR TO THE PARADIGM REAL-TIME INSULIN INFUSION PUMP, WHICH IS MARKETED IN THE UNITED STATES. | PUMP, INFUSION, INSULIN, TO BE USED WITH INVASIVE GLUCOSE SENSOR | MEDTRONIC PUERTO RICO OPERATIONS CO. | MMT-1710K | HG0QB1B | OYC |  |  | No | LAY USER/PATIENT | Pump, Infusion, Insulin, To Be Used With Invasive Glucose Sensor
Injury | 20200330 | ['Battery Problem', 'Impedance Problem', 'Insufficient Information'] | THE HEALTHCARE PROFESSIONAL (HCP) REPORTED THROUGH A MANUFACTURER REPRESENTATIVE THAT THE IMPLANTABLE NEUROSTIMULATOR (INS) IMPEDANCE WAS ¿ABNORMAL AND THE REMAINING BATTERY POWER DID NOT CHANGE.¿ IMPEDANCE RESULTS INDICATED THE #4 ELECTRODE WAS 1549 OHMS AND ALL OTHER ELECTRODES WERE DISPLAYED AS 3099 OHMS. THE INS STATUS WAS ¿OK¿ AT 2.6V WITH A CAPACITY USED OF 40-85%; ¿THERE WAS NO CHANGE FOR ABOUT TWO YEARS.¿ THE PATIENT¿S PHYSICIAN OBSERVED THE CAUSE OF THE ISSUE TO BE A PRODUCT MALFUNCTION. THE INS WAS REPLACED AS A RESULT OF THE EVENT AND THE ISSUE WAS RESOLVED. IT WAS NOTED THE DISEASE THE PATIENT WAS PRIMARILY/ORIGINALLY TREATED FOR WAS A CONSCIOUSNESS DISORDER. THERE WERE NO FURTHER COMPLICATIONS REPORTED OR ANTICIPATED. |  | IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. DEVICE EVALUATION: ANALYSIS OF THE IMPLANTABLE NEUROSTIMULATOR (INS) FOUND THE INS WAS FUNCTIONALLY OKAY WITH INSIGNIFICANT ANOMALIES. THE INS PASSED FUNCTIONAL TESTING. PLEASE NOTE THAT METHOD CODE 4114 AND RESULT CODE 3221 NO LONGER APPLY TO THIS REPORT. MEDTRONIC IS SUBMITTING THIS REPORT TO COMPLY WITH FDA REPORTING REGULATIONS UNDER 21 CFR PARTS 4 AND 803. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY MEDTRONIC, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. MEDTRONIC HAS MADE REASONABLE EFFORTS TO OBTAIN MORE COMPLETE INFORMATION AND HAS PROVIDED AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE AN ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY ¿DEFECTS¿ OR HAS ¿MALFUNCTIONED¿. THESE WORDS ARE INCLUDED IN THE FDA 3500A FORM AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REGULATORY REPORTING. MEDTRONIC OBJECTS TO THE USE OF THESE WORDS AND OTHERS LIKE THEM BECAUSE OF THE LACK OF DEFINITION AND THE CONNOTATIONS IMPLIED BY THESE TERMS. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. A GOOD FAITH EFFORT WILL BE MADE TO OBTAIN THE APPLICABLE INFORMATION RELEVANT TO THE REPORT. IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. | STIMULATOR, SPINAL-CORD, TOTALLY IMPLANTED FOR PAIN RELIEF | MEDTRONIC EUROPE SARL | 7427V |  | LGW | 20131214 | 20200519 | Device was returned to manufacturer | HEALTH PROFESSIONAL | Stimulator, Spinal-Cord, Totally Implanted For Pain Relief
Malfunction | 20200424 | ['Device Difficult to Program or Calibrate', 'Battery Problem', 'Patient Device Interaction Problem'] | INFORMATION RECEIVED BY MEDTRONIC INDICATED THAT THE ACT BUTTON WAS STICKING. CUSTOMER STATED THAT THE INSULIN PUMP REJECTED NEW BATTERIES. CUSTOMER STATED THAT THE RUBBER BASE WAS MISSING. NO HARM REQUIRING MEDICAL INTERVENTION WAS REPORTED. THE INSULIN PUMP WILL NOT BE RETURNED FOR ANALYSIS. |  | (B)(4). CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. | PUMP, INFUSION, INSULIN, TO BE USED WITH INVASIVE GLUCOSE SENSOR | MEDTRONIC PUERTO RICO OPERATIONS CO. | MMT-523RNAS | A1523RNASJ | OYC |  |  | No | LAY USER/PATIENT | Pump, Infusion, Insulin, To Be Used With Invasive Glucose Sensor
Injury | 20200413 | ['Battery Problem'] | FOLLOWING THE BATTERY PERFORMANCE ALERT (BPA) ADVISORY, A BPA WAS RECEIVED BY THE CLINICIAN AND THE DEVICE WAS EXPLANTED. THE PATIENT WAS IN STABLE CONDITION. |  | PREMATURE BATTERY DEPLETION WAS CONFIRMED BY ANALYSIS. A DEVICE EVALUATION WAS PERFORMED, AND NO SOURCES OF HIGH CURRENT WERE NOTED. THE CAUSE OF THE PREMATURE BATTERY DEPLETION WAS CONSISTENT WITH LI CLUSTER FORMATION. FROM THESE ANALYSES, IN THE ABSENCE OF HIGH CURRENT DRAW, IT IS PROBABLE THAT THE PREMATURE BATTERY DEPLETION WAS CAUSED BY A LITHIUM CLUSTER INDUCED SHORT CIRCUIT. LI CLUSTERS ARE A KNOWN DEPLETION MECHANISM FOR THESE ADVISORY PRODUCTS THAT HAS BEEN INVESTIGATED AND ASSOCIATED WITH A FIELD ACTION IN OCTOBER 2016. THE DEVICE IS INCLUDED IN THE BATTERY PERFORMANCE ALERT ADVISORY ISSUED BY ABBOTT ON 28 AUGUST 2017. THE RESULTS, METHOD, AND CONCLUSION CODES ALONG WITH INVESTIGATION RESULTS WILL BE PROVIDED IN THE FINAL REPORT. | IMPLANTABLE CARDIOVERTER DEFIBRILLATOR | ABBOTT | CD3367-40QC | 4540474 | NIK | 20160331 | 20200609 | Device was returned to manufacturer | HEALTH PROFESSIONAL | Defibrillator, Automatic Implantable Cardioverter, With Cardiac Resynchronization (Crt-D)
Malfunction | 20200317 | ['Failure to Calibrate', 'Battery Problem', 'Failure of Device to Self-Test', 'Protective Measures Problem'] | IT WAS REPORTED TO RESMED THAT AN ASTRAL DEVICE FAILED TO COMPLETE ITS INTERNAL SELF-TEST AND DISPLAYED AN INTERNAL BATTERY DEGRADED WARNING ALARM. THERE WAS NO PATIENT HARM OR SERIOUS INJURY REPORTED AS A RESULT OF THIS INCIDENT. | IT WAS REPORTED TO RESMED THAT AN ASTRAL DEVICE FAILED TO COMPLETE ITS INTERNAL SELF-TEST. THERE WAS NO PATIENT HARM OR SERIOUS INJURY REPORTED AS A RESULT OF THIS INCIDENT. | THE ASTRAL DEVICE WAS RETURNED TO RESMED FOR AN INVESTIGATION. REVIEW OF THE DEVICE DATA LOGS CONFIRMED AN INTERNAL BATTERY DEGRADED WARNING ALARM AND REVEALED AN ERROR MESSAGE (SF180) RELATED TO A BATTERY CHARGER FAULT. THE INTERNAL BATTERY WAS REPLACED TO ADDRESS THE ISSUE. BASED ON ALL AVAILABLE EVIDENCE AND COMPLAINT INVESTIGATIONS OF A SIMILAR NATURE, THE INVESTIGATION DETERMINED THAT THE REPORTED INTERNAL BATTERY DEGRADED WARNING ALARM AND SF180 WERE DUE TO AN ISOLATED COMPONENT FAILURE WITHIN THE DEVICE BATTERY ASSEMBLY WHILE THE FAILURE TO COMPLETE ITS INTERNAL SELF-TEST WAS DUE TO A DEFECTIVE NRV. RESMED'S RISK ANALYSIS FOR THIS FAILURE MODE CONCLUDES THAT THE RISK IS ACCEPTABLE. RESMED REFERENCE#: (B)(4). THE DEVICE WAS RETURNED TO RESMED AND AN EVALUATION CONFIRMED THE COMPLAINT. THE NON-RETURN VALVE (NRV) ASSEMBLY WAS REPLACED TO ADDRESS THIS ISSUE. THE DEVICE WAS SERVICED AND FULLY TESTED BEFORE IT WAS RETURNED TO THE CUSTOMER. RESMED REFERENCE#: (B)(4). |  | RESMED LTD | 27093 |  | CBK |  | 20200406 | Device was returned to manufacturer | LAY USER/PATIENT | Ventilator, Continuous, Facility Use
Malfunction | 20200313 | ['Battery Problem'] | INFORMATION WAS RECEIVED INDICATING THAT SMITHS MEDICAL MEDFUSION 4000 PUMP HAS MALFUNCTIONED. IT WAS REPORTED THAT THE DEVICE'S BATTERY DEPLETED. IT WAS ALSO INDICATED THAT THE PUMP'S BASE WAS NOT RESPONDING. THERE WERE NO ADVERSE EVENTS REPORTED. |  |  |  | ST PAUL | 4000-0105-51 |  | FRN |  |  | No | HEALTH PROFESSIONAL | Pump, Infusion
Malfunction | 20200317 | ['Failure to Calibrate', 'Battery Problem', 'Failure of Device to Self-Test', 'Protective Measures Problem'] | IT WAS REPORTED TO RESMED THAT AN ASTRAL DEVICE FAILED TO COMPLETE ITS INTERNAL SELF-TEST AND DISPLAYED AN INTERNAL BATTERY DEGRADED WARNING ALARM. THERE WAS NO PATIENT HARM OR SERIOUS INJURY REPORTED AS A RESULT OF THIS INCIDENT. | IT WAS REPORTED TO RESMED THAT AN ASTRAL DEVICE FAILED TO COMPLETE ITS INTERNAL SELF-TEST. THERE WAS NO PATIENT HARM OR SERIOUS INJURY REPORTED AS A RESULT OF THIS INCIDENT. | THE ASTRAL DEVICE WAS RETURNED TO RESMED FOR AN INVESTIGATION. REVIEW OF THE DEVICE DATA LOGS CONFIRMED AN INTERNAL BATTERY DEGRADED WARNING ALARM AND REVEALED AN ERROR MESSAGE (SF180) RELATED TO A BATTERY CHARGER FAULT. THE INTERNAL BATTERY WAS REPLACED TO ADDRESS THE ISSUE. BASED ON ALL AVAILABLE EVIDENCE AND COMPLAINT INVESTIGATIONS OF A SIMILAR NATURE, THE INVESTIGATION DETERMINED THAT THE REPORTED INTERNAL BATTERY DEGRADED WARNING ALARM AND SF180 WERE DUE TO AN ISOLATED COMPONENT FAILURE WITHIN THE DEVICE BATTERY ASSEMBLY WHILE THE FAILURE TO COMPLETE ITS INTERNAL SELF-TEST WAS DUE TO A DEFECTIVE NRV. RESMED'S RISK ANALYSIS FOR THIS FAILURE MODE CONCLUDES THAT THE RISK IS ACCEPTABLE. RESMED REFERENCE#:(B)(4). THE DEVICE WAS RETURNED TO RESMED AND AN EVALUATION CONFIRMED THE COMPLAINT. THE NON-RETURN VALVE (NRV) ASSEMBLY WAS REPLACED TO ADDRESS THIS ISSUE. THE DEVICE WAS SERVICED AND FULLY TESTED BEFORE IT WAS RETURNED TO THE CUSTOMER. (B)(4). |  | RESMED LTD | 27093 |  | CBK |  | 20200406 | Device was returned to manufacturer | LAY USER/PATIENT | Ventilator, Continuous, Facility Use
Malfunction |  | ['Battery Problem'] | PER TM - THE SCANNER IS NOT HOLDING A CHARGE. ON (B)(6) 2020 - SAMPLE EVALUATION CONFIRMED BATTERY ABRUPT SHUTDOWN. |  | THE DEVICE WAS RETURNED TO THE SERVICE FACILITY FOR EVALUATION. DURING EVALUATION, THE REPORTED ISSUE OF SCANNER IS NOT HOLDING A CHARGE WAS CONFIRMED. THE SCANNER WAS INSPECTED IT SHUTS DOWN PREMATURELY WHEN THE AC POWER REMOVED, THE CAUSE OF THE ISSUE IS DUE TO AN INTERNAL LI-ION BATTERY FAILURE. THE DEVICE WAS SERVICED, TESTED AND RETURNED TO THE CUSTOMER. A HISTORY REVIEW OF SERIAL NUMBER (B)(4) SHOWED ONE OTHER SIMILAR COMPLAINT FROM THIS SERIAL NUMBER. THE PREVIOUS COMPLAINT FOR THIS SERIAL NUMBER (B)(4) EVALUATION WAS CONFIRMED (REFERENCE: MDR NUMBER 3006260740- 2020-00772). | SYSTEM, IMAGING, PULSED ECHO, ULTRASONIC | BARD ACCESS SYSTEMS | 9770600 | N/A | IYO |  | 20191126 | Device was returned to manufacturer | HEALTH PROFESSIONAL | System, Imaging, Pulsed Echo, Ultrasonic
Malfunction | 20200518 | ['Battery Problem', 'Therapeutic or Diagnostic Output Failure'] | THE RECIPIENT REPORTEDLY EXPERIENCED INTERMITTENCIES FOLLOWED BY LOSS OF SOUND. EXTERNAL EQUIPMENT WAS EXCHANGED AND PROGRAMMING ADJUSTMENTS WERE MADE, HOWEVER, THE RECIPIENT THEN EXPERIENCED SOUND QUALITY ISSUES, AND DECREASED BATTER LIFE. REVISION SURGERY WILL BE SCHEDULED. |  | DISCLAIMER: ADVANCED BIONICS DOES NOT INTEND THAT THIS REPORT BE ANY ADMISSION OF LIABILITY, FAULT OR PRODUCT DEFECT. THE RECIPIENT'S DEVICE WAS EXPLANTED. THE RECIPIENT WAS REIMPLANTED WITH ANOTHER ADVANCED BIONICS COCHLEAR DEVICE. DISCLAIMER: ADVANCED BIONICS DOES NOT INTEND THAT THIS REPORT BE ANY ADMISSION OF LIABILITY, FAULT OR PRODUCT DEFECT. ADVANCED BIONICS CONSIDERS THE INVESTIGATION INTO THIS REPORTABLE EVENT AS CLOSED. THE EXTERNAL VISUAL INSPECTION REVEALED THAT THE ELECTRODE WAS SEVERED PRIOR TO RECEIPT, AS WELL AS TOOL DAMAGE TO THE TOP AND BOTTOM COVER. ALL OF THESE ANOMALIES ARE BELIEVED TO HAVE OCCURRED DURING REVISION SURGERY. THE DEVICE PASSED THE PHOTOGRAPHIC IMAGING INSPECTION. SYSTEM LOCK COULD ONLY BE OBTAINED AT CERTAIN SPACING. THE ELECTRODE CONDITION PREVENTED SOME ELECTRICAL TESTS FROM BEING PERFORMED. THE DEVICE PASSED SOME OF THE ELECTRICAL TESTS PERFORMED. THE DEVICE FAILED THE RESIDUAL GAS ANALYSIS TEST. THIS DEVICE HAD MOISTURE THAT EXCEEDED THE RESIDUAL GAS TEST LIMIT. BASED ON ASSESSMENT OF THE RESIDUAL GAS ANALYSIS TEST DATA, IT IS BELIEVED THAT THIS DEVICE WAS NOT HERMETIC. THIS ULTIMATELY CAUSED THE DEVICE TO CEASE FUNCTIONING. CORRECTIVE ACTIONS WERE IMPLEMENTED. THIS IS THE FINAL REPORT. DISCLAIMER: ADVANCED BIONICS DOES NOT INTEND THAT THIS REPORT BE ANY ADMISSION OF LIABILITY, FAULT OR PRODUCT DEFECT. DISCLAIMER: ADVANCED BIONICS DOES NOT INTEND THAT THIS REPORT BE ANY ADMISSION OF LIABILITY, FAULT OR PRODUCT DEFECT. | COCHLEAR IMPLANT | ADVANCED BIONICS, LLC | CI-1400-02H | NA | MCM | 20060601 | 20200618 | Device was returned to manufacturer | LAY USER/PATIENT | Implant, Cochlear
Malfunction |  | ['Battery Problem'] | THE CUSTOMER REPORTED BATTERY FAILURE. THE DEVICE WAS NOT IN CLINICAL USE AT THE TIME THE ISSUE WAS DISCOVERED. THERE WAS NO PATIENT OR USER HARM REPORTED. |  | DATE OF EVENT: (B)(6) 2020. DATE OF REPORT: 28APR2020. G4: 20JUL2020 B4: (B)(6) 2020 THE KEY MARKET CONTACT WAS CONTACTED AND STATED THAT THE CUSTOMER REPLACED THE BATTERY ON THEIR OWN TO RESOLVE THE ISSUE. THE CUSTOMER DISPOSED OF THE DEFECTIVE BATTERY AND IT WILL NOT BE RETURNED. SUBMISSION OF A REPORT DOES NOT CONSTITUTE AN ADMISSION THAT MEDICAL PERSONNEL, USER FACILITY, IMPORTER, DISTRIBUTOR, MANUFACTURER, OR PRODUCT CAUSED OR CONTRIBUTED TO THE EVENT. | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS CALIFORNIA, INC | V200 |  | CBK |  |  | Yes | I | Ventilator, Continuous, Facility Use
Malfunction |  | ['Battery Problem'] | THE CUSTOMER REPORTED BATTERY FAILURE. THE CUSTOMER INDICATED THE BUILT-IN ACCUMULATOR BATTERY IS OUT OF ORDER. IT IS UNKNOWN IF THE DEVICE DID HAVE PATIENT INVOLVEMENT AT THE TIME THE ISSUE WAS DISCOVERED, HOWEVER, THERE WAS NO PATIENT OR USER HARM REPORTED. |  | G4: 06JUL2020 B4: (B)(6)2020 THE KEY MARKET CONFIRMED THAT THE CUSTOMER JUST REQUESTED THE BATTERY AS A PREVENTIVE MEASURE. THERE'S IS NO ALLEGED MALFUNCTION. SUBMISSION OF A REPORT DOES NOT CONSTITUTE AN ADMISSION THAT MEDICAL PERSONNEL, USER FACILITY, IMPORTER, DISTRIBUTOR, MANUFACTURER, OR PRODUCT CAUSED OR CONTRIBUTED TO THE EVENT. DATE OF EVENT: (B)(6) 2020. DATE OF REPORT: 28APR2020. | VENTILATOR, CONTINUOUS, MINIMAL VENTILATORY SUPPORT,FACILITY USE | RESPIRONICS CALIFORNIA, INC | V60 |  | MNT |  |  | Yes | I | Ventilator, Continuous, Minimal Ventilatory Support, Facility Use
Injury | 20200409 | ['Battery Problem'] | FOLLOWING THE BATTERY PERFORMANCE ALERT (BPA) ADVISORY, A BPA WAS RECEIVED BY THE CLINICIAN AND THE DEVICE WAS EXPLANTED. THE PATIENT WAS DISCHARGED. |  | THE DEVICE IS INCLUDED IN THE BATTERY PERFORMANCE ALERT ADVISORY ISSUED BY ABBOTT ON 28 AUGUST 2017. PREMATURE BATTERY DEPLETION WAS CONFIRMED BY ANALYSIS. NO SOURCES OF HIGH CURRENT WERE NOTED. THE CAUSE OF THE PREMATURE BATTERY DEPLETION WAS CONSISTENT WITH LITHIUM (LI) CLUSTER FORMATION. FROM THESE ANALYSES, IN THE ABSENCE OF HIGH CURRENT DRAW, IT IS PROBABLE THAT THE PREMATURE BATTERY DEPLETION WAS CAUSED BY A LITHIUM CLUSTER INDUCED SHORT CIRCUIT. LI CLUSTERS ARE A KNOWN DEPLETION MECHANISM FOR THESE ADVISORY PRODUCTS THAT HAS BEEN INVESTIGATED AND ASSOCIATED WITH A FIELD ACTION IN OCTOBER 2016. | IMPLANTABLE CARDIOVERTER DEFIBRILLATOR | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | CD1357-40Q | 4523602 | LWS | 20160331 | 20200429 | Device was returned to manufacturer | HEALTH PROFESSIONAL | Implantable Cardioverter Defibrillator (Non-Crt)
Malfunction | 20200424 | ['Mechanical Problem', 'Device Difficult to Program or Calibrate', 'Obstruction of Flow', 'Battery Problem', 'Power Problem'] | INFORMATION RECEIVED BY MEDTRONIC INDICATED THAT THE INSULIN PUMP HAD REJECTED NEW BATTERY, BATTERY LIFE WAS 6 DAYS, HAD UNEXPLAINED NO DELIVERY ALARMS, WAS SLOW TO REWIND, WAS GRINDING DURING REWIND AND THE BUTTONS HAD TO PRESSED MORE THAN ONCE. NO HARM REQUIRING MEDICAL INTERVENTION WAS REPORTED. TROUBLESHOOTING WAS DECLINED. THE INSULIN PUMP WILL NOT BE RETURNED FOR ANALYSIS. |  | (B)(4). CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. | PUMP, INFUSION, INSULIN, TO BE USED WITH INVASIVE GLUCOSE SENSOR | MEDTRONIC PUERTO RICO OPERATIONS CO. | MMT-723LNAS | A6723LNASJ | OYC |  |  | No | LAY USER/PATIENT | Pump, Infusion, Insulin, To Be Used With Invasive Glucose Sensor
Malfunction |  | ['Failure to Power Up', 'Battery Problem'] | THE CUSTOMER REPORTED A VENTILATOR THAT WOULD NOT TURN ON DUE TO A FAILED BATTERY. THIS EVENT DID NOT OCCUR DURING CLINICAL USE. THERE WAS NO ALLEGATION OF HARM ASSOCIATED WITH THIS REPORT. |  | DATE OF EVENT: (B)(6) 2020. DATE OF REPORT: 28APR2020. G4: (B)(6) 2021. B4: (B)(6) 2021 H11: B5: IT WAS REPORTED TO PHILIPS THAT THE DEVICE WAS UNABLE TO TURN ON. H10: WHEN THE FIELD SERVICE ENGINEER (FSE) WENT ONSITE TO IMPLEMENT THE POWER MANAGEMENT BOARD'S FIELD CHANGE ORDER, THE CUSTOMER TOLD FSE THAT THE DEVICE HAD BEEN REMOVED FROM SERVICE. THE PHILIPS FIELD SERVICE ENGINEER (FSE) WAS UNABLE TO EVALUATE AND CONFIRM THE REPORTED ISSUE. NO EVALUATION AND REPAIR WERE PERFORMED TO DETERMINE THE ROOT CAUSE SUBMISSION OF A REPORT DOES NOT CONSTITUTE AN ADMISSION THAT MEDICAL PERSONNEL, USER FACILITY, IMPORTER, DISTRIBUTOR, MANUFACTURER, OR PRODUCT CAUSED OR CONTRIBUTED TO THE EVENT. | VENTILATOR, CONTINUOUS, MINIMAL VENTILATORY SUPPORT,FACILITY USE |  | V60 |  | MNT |  |  | Yes | HEALTH PROFESSIONAL | Ventilator, Continuous, Minimal Ventilatory Support, Facility Use
Malfunction |  | ['Battery Problem'] | THE CUSTOMER REPORTED THAT THE BUILT-IN BATTERY IS FAULTY. THE CUSTOMER CONTACTED PRODUCT SUPPORT AND REQUESTED FOR SERVICE REPAIR. IT IS UNKNOWN IF THE UNIT WAS IN USE ON A PATIENT. |  | G4: 21APR2020 B4: (B)(6)2020. PER TECHNICAL SUPPORT THE UNIT WAS IN USE ON PATIENT, BUT THERE'S NO PATIENT HARM REPORTED. SUBMISSION OF A REPORT DOES NOT CONSTITUTE AN ADMISSION THAT MEDICAL PERSONNEL, USER FACILITY, IMPORTER, DISTRIBUTOR, MANUFACTURER, OR PRODUCT CAUSED OR CONTRIBUTED TO THE EVENT. DATE OF EVENT: (B)(6) 2020. DATE OF REPORT: 28APR2020. G4: 06JUL2020 B4: (B)(6)2020 PER TECHNICAL SUPPORT THE CUSTOMER DIDN'T BUY PART TO REPAIR THE UNIT. THE CUSTOMER ALSO DID NOT ACCEPT THE QUOTE FOR REPAIR. THE UNIT IS NOT UNDER WARRANTY AND THE CUSTOMER DID NOT RESPONSE AND THE CASE WAS CLOSED OUT BY TECHNICAL SUPPORT DUE TO LOCK OF RESPONSE FROM CUSTOMER. A SUPPLEMENTAL REPORT WILL BE SUBMITTED IF NEW INFORMATION IS RECEIVED. SUBMISSION OF A REPORT DOES NOT CONSTITUTE AN ADMISSION THAT MEDICAL PERSONNEL, USER FACILITY, IMPORTER, DISTRIBUTOR, MANUFACTURER, OR PRODUCT CAUSED OR CONTRIBUTED TO THE EVENT. | VENTILATOR, CONTINUOUS, MINIMAL VENTILATORY SUPPORT,FACILITY USE | RESPIRONICS CALIFORNIA, INC | V60 |  | MNT |  |  | Yes | I | Ventilator, Continuous, Minimal Ventilatory Support, Facility Use
Malfunction | 20200424 | ['Mechanical Problem', 'Device Difficult to Program or Calibrate', 'Battery Problem'] | INFORMATION RECEIVED BY MEDTRONIC INDICATED THAT THE INSULIN PUMP HAD STUCK BUTTON. CUSTOMER REPORTED THAT THE BUTTON NEED TO PRESSED MULTIPLE TIME. CUSTOMER REPORTED THAT THE BATTERY LIFE SOME TIME LESS. THE INSULIN PUMP WAS SLOWER AND MADE MORE NOISE REWIND. NO HARM REQUIRING MEDICAL INTERVENTION WAS REPORTED. THE DEVICE WILL NOT BE RETURNED FOR ANALYSIS. |  | (B)(4). CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. | ARTIFICIAL PANCREAS DEVICE SYSTEM, THRESHOLD SUSPEND | MEDTRONIC PUERTO RICO OPERATIONS CO. | MMT-751LNAS | A4751LNASJ | OZO |  |  | No | LAY USER/PATIENT | Automated Insulin Dosing , Threshold Suspend
Malfunction | 20200424 | ['Device Difficult to Program or Calibrate', 'Obstruction of Flow', 'Battery Problem'] | INFORMATION RECEIVED BY MEDTRONIC INDICATED THAT THE INSULIN PUMP BUTTONS HAD PRESS HARDER. CUSTOMER'S SPOUSE STATED THE BATTERY LIFE WAS NOT LASTING AND INSULIN PUMP GOT LOT CLOSER TO OUT OF INSULIN BEFORE WARNING. NO HARM REQUIRING MEDICAL INTERVENTION WAS REPORTED. TROUBLESHOOTING WAS DECLINED. THE INSULIN PUMP WILL NOT BE RETURNED FOR ANALYSIS. |  | (B)(4). CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. ALL BUTTONS FUNCTIONING PROPERLY. NO BUTTON ERROR ALARM DURING TESTING. NO MOISTURE OR CORRODED KEYPAD FOUND. NO FLATTED KEYPAD. CONNECTOR WAS INSPECTED AND LOCKED ON LCD BOARD. DEVICE PASSED DISPLACEMENT TEST, SELF TEST, OFF NO POWER TEST AND ALL OPERATING CURRENTS TESTED WITHIN THE SPECIFICATION. NO CHARGE OR BATTERY ANOMALY WERE NOTED. DEVICE PASSED REWIND TEST, BASIC OCCLUSION TEST, PRIME OR A33 TEST, EXCESSIVE NO DELIVERY TEST AND OCCLUSION TEST. NO UNEXPECTED NO DELIVERY ALARM NOTED. (B)(4). | ARTIFICIAL PANCREAS DEVICE SYSTEM, THRESHOLD SUSPEND | MEDTRONIC PUERTO RICO OPERATIONS CO. | MMT-751NAS | A6751NASJ | OZO |  | 20200701 | Device was returned to manufacturer | LAY USER/PATIENT | Automated Insulin Dosing , Threshold Suspend
Malfunction | 20200424 | ['Insufficient Flow or Under Infusion', 'Battery Problem'] | INFORMATION RECEIVED BY MEDTRONIC INDICATED THAT THE INSULIN PUMP WAS NOT GIVING INSULIN PROPERLY AND ALSO BATTERIES DYING QUICKER. NO HARM REQUIRING MEDICAL INTERVENTION WAS REPORTED. THE INSULIN PUMP WILL NOT BE RETURNED FOR ANALYSIS. |  | (B)(4). CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. | PUMP, INFUSION, INSULIN, TO BE USED WITH INVASIVE GLUCOSE SENSOR | MEDTRONIC PUERTO RICO OPERATIONS CO. | MMT-723NAH | B5723NAHJ | OYC |  |  | No | LAY USER/PATIENT | Pump, Infusion, Insulin, To Be Used With Invasive Glucose Sensor
Malfunction | 20200407 | ['Battery Problem'] | IT WAS REPORTED THAT THE SYSTEM CONTROLLER SHOWED A "REPLACE BACKUP BATTERY" ALARM, ALTHOUGH THE BATTERY WAS NOT DUE TO BE REPLACED. THE SYSTEM CONTROLLER WAS EXCHANGED. NO FURTHER INFORMATION WAS PROVIDED. |  | NO FURTHER INFORMATION WAS PROVIDED. THE PATIENT REMAINS ONGOING ON LVAS SUPPORT. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETED. SECTION D10, H3: ADDITIONAL INFORMATION. SECTION D4, H4: CORRECTION. MANUFACTURER'S INVESTIGATION CONCLUSIONS: THE REPORTED EVENT OF A BACKUP BATTERY FAULT ALARM WAS CONFIRMED VIA THE LOG FILE DOWNLOADED FROM THE RETURNED SYSTEM CONTROLLER (SERIAL NUMBER: (B)(6)); HOWEVER, THE ALARM WAS NOT REPRODUCED DURING TESTING. THE LOG FILE CONTAINED APPROXIMATELY 5 DAYS OF DATA (30MAR2020 ¿ 03APR2020 PER THE TIMESTAMP). THE LOG FILE CAPTURED BACKUP BATTERY FAULT ALARMS DUE TO BACKUP BATTERY LOAD TEST FAILURES ON 03APR2020 FROM 20:03:42 TO 23:11:31, WHEN THE CONTROLLER WAS POWERED DOWN AFTER BEING EXCHANGED. THE ALARM BRIEFLY CLEARED MULTIPLE TIMES WHEN ADDITIONAL LOAD TESTS WERE PERFORMED; HOWEVER, THE ALARM RETURNED EACH TIME AS THE BACKUP BATTERY FAILED THE LOAD TEST EACH TIME. THE BACKUP BATTERY FAILED THE LOAD TESTS DUE TO THE LOADED VOLTAGE BEING UNDER THE ACCEPTABLE THRESHOLD COMPARED TO THE UNLOADED VOLTAGE. THE ALARMS DID NOT AFFECT THE CONTROLLER¿S ABILITY TO OPERATE THE PUMP AT THE SET SPEED. THE DRIVELINE WAS DISCONNECTED ON 03APR2020 AT 23:11:16 TO EXCHANGE THE SYSTEM CONTROLLER. NO OTHER NOTABLE ALARMS WERE ACTIVE IN THE LOG FILE. THE PUMP MAINTAINED A SPEED ABOVE OR EQUAL TO THE LOW SPEED LIMIT THROUGHOUT THE LOG FILE WHILE THE DRIVELINE WAS CONNECTED. THE RETURNED SYSTEM CONTROLLER PASSED FUNCTIONAL TESTING AND SUCCESSFULLY OPERATED IN A MOCK CIRCULATORY LOOP FOR AN EXTENDED PERIOD OF TIME WITHOUT ANY ISSUES OR ATYPICAL ALARMS PRODUCED. DURING TESTING, MULTIPLE LOAD TESTS WERE PERFORMED AND THE CONTROLLER PASSED EACH TIME WITHOUT ISSUE. THE ROOT CAUSE OF THE REPORTED EVENT COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS ANALYSIS. THERE WAS AN INCIDENTAL FINDING OF DIM PUMP RUNNING LEDS. DEVICE HISTORY RECORDS WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QA SPECIFICATIONS. (B)(6) WAS SHIPPED TO THE CUSTOMER ON 09NOV2017. THE HEARTMATE III PATIENT HANDBOOK AND HEARTMATE III INSTRUCTIONS FOR USE (IFU) COVER ALL ALARMS (VISUAL AND AUDIBLE), INCLUDING THE BACKUP BATTERY FAULT ALARM CONDITIONS, AND THE ACTIONS TO TAKE IF THE ALARMS CANNOT BE RESOLVED. THE HEARTMATE III IFU EXPLAINS HOW TO REPLACE THE SYSTEM CONTROLLER BACKUP BATTERY, AND HOW TO REPLACE THE SYSTEM CONTROLLER. SECTION 5 ¿SURGICAL PROCEDURES¿, EXPLAINS HOW TO INSTALL THE BACKUP BATTERY IN THE SYSTEM CONTROLLER. THE HEARTMATE III PATIENT HANDBOOK AND HEARTMATE III IFU CAUTION USERS TO CALL THEIR HOSPITAL CONTACTS IF THEY THINK, FOR ANY REASON, ANY PORTION OF THEIR EQUIPMENT IS NOT FUNCTIONING AS USUAL, IS BROKEN, OR THEY ARE UNCOMFORTABLE WITH THE OPERATION OF THE EQUIPMENT. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. | VENTRICULAR (ASSISST) BYPASS | THORATEC CORPORATION | 106531INT | 6203435 | DSQ | 20181231 | 20200528 | Device was returned to manufacturer | LAY USER/PATIENT | Ventricular (Assist) Bypass
Malfunction | 20200315 | ['Battery Problem'] | INFORMATION WAS RECEIVED INDICATING THAT SMITHS MEDICAL MEDFUSION 4000 PUMP HAS MALFUNCTIONED. WHEN MOVING THE INFANT TO A NEW BED, ALL PUMPS WERE UNPLUGGED; ONE SYRINGE PUMP THAT HAD BEEN PLUGGED IN AND RUNNING IMMEDIATELY ALARMED "DEPLETED BATTERY. THE SYRINGE WAS MOVED TO THE OTHER PUMP AND FEEDINGS CONTINUED. THERE WERE NO ADVERSE EVENTS REPORTED. |  |  | PUMP, INFUSION | ST PAUL | 4000-0105-51 |  | FRN |  |  | No | HEALTH PROFESSIONAL | Pump, Infusion
Malfunction | 20200318 | ['Battery Problem'] | INFORMATION WAS RECEIVED INDICATING THAT SMITHS MEDICAL MEDFUSION 4000 PUMP HAS MALFUNCTIONED. THE PUMP WAS USED TO INFUSE LIPIDS IN 60 ML SYRINGE. THE PUMP WAS UNPLUGGED FROM THE WALL IN ORDER TO TRANSPORT THE PATIENT. THE PUMP IMMEDIATELY STARTED ALARMING "SYSTEM FAILURE BATTERY DEPLETED." THE PUMP'S HISTORY REVEALED THAT THERE WAS A BATTERY COMMUNICATION TIMEOUT AND A BATTERY DEPLETED ALERT. THERE WERE NO ADVERSE EVENTS REPORTED. |  |  | PUMP, INFUSION | ST PAUL | 4000-0105-51 |  | FRN |  |  | No | HEALTH PROFESSIONAL | Pump, Infusion
Malfunction | 20200309 | ['Battery Problem'] | INFORMATION WAS RECEIVED INDICATING THAT SMITHS MEDICAL MEDFUSION 4000 PUMP HAS MALFUNCTIONED. IT WAS REPORTED THAT THE PUMP WAS ALARMING BATTERY COMMUNICATION TIMEOUT AND IT WAS SWITCHED BY ANOTHER PUMP. THE DEVICE'S HISTORY SHOWED A SYSTEM ADVISORY: BASE TASK DEBILITATED. IT DISPLAYED ?REBOOTING DUE TO EXCEPTION", THEN LATER GAVE BATTERY COMMUNICATION TIMEOUT AND INVALID SYRINGE SIZE MESSAGES. THERE WERE NO ADVERSE EVENTS REPORTED. |  |  | PUMP, INFUSION | ST PAUL | 4000-0106-01 |  | FRN |  |  | No | HEALTH PROFESSIONAL | Pump, Infusion
Malfunction | 20200319 | ['Battery Problem'] | INFORMATION WAS RECEIVED INDICATING THAT SMITHS MEDICAL MEDFUSION 4000 PUMP HAS MALFUNCTIONED. IT WAS REPORTED THAT THE NURSE REPORTED A LOW BATTERY ALARM. SINCE SHE DID NOT HAVE ENOUGH CORDS FOR ALL SYRINGE PUMPS, SHE MOVED THE CORD FROM HER MILRINONE PUMP TO THE PUMP IN NEED. IMMEDIATELY THE MILRINONE PUMP ALARMED BIOMED AND SHUT OFF. THE DEVICE'S HISTORY SHOWED THAT THE PUMP DIED UPON PLUGGING IN. THERE WERE NO ADVERSE EVENTS REPORTED. |  | E4 IS UNKNOWN. NO INFORMATION HAS BEEN PROVIDED TO DATE. THE PUMP WAS NOT RETURNED TO THE MANUFACTURER BUT WAS INVESTIGATED AND REPAIRED AT THE CUSTOMER FACILITY BY A FIELD SERVICE TECHNICIAN. A VISUAL INSPECTION OF THE PUMP FOUND THAT THE FACTORY SEAL IS MISSING. THE BATTERY/POWER CORD WAS WORKING AS INTENDED. DURING FUNCTIONAL TESTING, NO ALARMS WERE PRESENT. THE REPORTED FAILURE WAS NOT CONFIRMED. THE ROOT CAUSE COULD NOT BE ESTABLISHED. A DEVICE HISTORY RECORD (DHR) REVIEW WAS CONDUCTED WHICH INDICATED ALL INSPECTIONS WERE COMPLETED AND NO ISSUES WERE NOTED DURING MANUFACTURING. THIS REMEDIATION MDR WAS GENERATED UNDER PROTOCOL B10009406, AS A RESULT OF WARNING LETTER CMS#(B)(4). E4 IS UNKNOWN. NO INFORMATION HAS BEEN PROVIDED TO DATE. THE PUMP WAS NOT RETURNED TO THE MANUFACTURER BUT WAS INVESTIGATED AND REPAIRED AT THE CUSTOMER FACILITY BY A FIELD SERVICE TECHNICIAN. A VISUAL INSPECTION OF THE PUMP FOUND THAT THE FACTORY SEAL IS MISSING. THE BATTERY/POWER CORD WAS WORKING AS INTENDED. DURING FUNCTIONAL TESTING, NO ALARMS WERE PRESENT. THE REPORTED FAILURE WAS NOT CONFIRMED. THE ROOT CAUSE COULD NOT BE ESTABLISHED. A DEVICE HISTORY RECORD (DHR) REVIEW WAS CONDUCTED WHICH INDICATED ALL INSPECTIONS WERE COMPLETED AND NO ISSUES WERE NOTED DURING MANUFACTURING. THIS REMEDIATION MDR WAS GENERATED UNDER PROTOCOL B10009406, AS A RESULT OF WARNING LETTER CMS# 617147. | PUMP, INFUSION | ST PAUL | 4000-0105-51 |  | FRN |  |  | No | HEALTH PROFESSIONAL | Pump, Infusion
Malfunction | 20200424 | ['Battery Problem', 'Power Problem'] | INFORMATION RECEIVED BY MEDTRONIC INDICATED THAT INSULIN PUMP HAD REPLACE BATTERY ALARM. THE CUSTOMER REPORTED NEW BATTERY DID NOT RESOLVE THE ISSUE. INSULIN PUMP HAD LOW BATTERY ALARM BEFORE REPLACE BATTERY ALARM. NO HARM REQUIRING MEDICAL INTERVENTION WAS REPORTED. THE INSULIN PUMP WILL BE RETURNED FOR ANALYSIS. |  | (B)(4). CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. THE DEVICE WILL BE RETURNED FOR ANALYSIS AND FURTHER INFORMATION WILL FOLLOW ONCE THE ANALYSIS HAS BEEN COMPLETED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. DEVICE PASSED THE DISPLACEMENT TEST. SLEEP CURRENT MEASUREMENT AND ACTIVE CURRENT MEASUREMENT WITHIN SPECIFICATIONS. THE POWER MANAGEMENT TOOL CONFIRMED POWER ERROR 25, LOW BATTERY ALERT, POWER LOSS ALARM, REPLACE BATTERY NOW ALARM WAS TRIGGERED WHEN THE BACKUP BATTERY LOADED VOLTAGE WAS LESS THAN 3.5V FOR 4 CONSECUTIVE HOURS DUE TO CONNECTOR RESISTANCE ELECTRICAL BOARD. AFTER DISCONNECTING AND RECONNECTING THE INTERNAL BATTERY CONNECTOR ON ELECTRICAL BOARD, DEVICE WAS MONITORED AND FUNCTIONED PROPERLY. DEVICE RECEIVED CRACKED BATTERY TUBE THREADS, FADING SERIAL NUMBER LABEL AND CRACKED CASE AT BATTERY TUBE SIDE. | ARTIFICIAL PANCREAS DEVICE SYSTEM, SINGLE HORMONAL CONTROL | MEDTRONIC PUERTO RICO OPERATIONS CO. | MMT-1780KPK | HG3WNEM | OZP |  | 20200508 | Device was returned to manufacturer | LAY USER/PATIENT | Automated Insulin Dosing Device System, Single Hormonal Control
Malfunction | 20200424 | ['Device Difficult to Program or Calibrate', 'Obstruction of Flow', 'Battery Problem'] | INFORMATION RECEIVED BY MEDTRONIC INDICATED THAT THE ACT BUTTON WAS HARD TO PUSH. CUSTOMER STATED THAT THE INSULIN PUMP HAD NO DELIVERY ALARMS. NO HARM REQUIRING MEDICAL INTERVENTION WAS REPORTED. THE INSULIN PUMP WILL NOT BE RETURNED FOR ANALYSIS. |  | (B)(4). CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. DEVICE RECEIVED WITH ALL BUTTONS RESPONDING PROPERLY. NO DAMAGE ON KEYPAD ASSEMBLY. NO UNLOCKED LCD KEYPAD CONNECTOR. DEVICE RECEIVED WITH ALL OPERATING CURRENTS WITHIN SPECIFICATION AND PASSED FUNCTIONAL TESTING INCLUDING THE REWIND, BASIC OCCLUSION TEST, OCCLUSION TEST, PRIME OR A33 TEST, EXCESSIVE NO DELIVERY TEST AND DISPLACEMENT TEST. NO EXCESSIVE NO DELIVERY OR OCCLUSION ALARM NOTED. NO UNEXPECTED LOW BATTERY ALARM ANOMALIES NOTED. (B)(4). | PUMP, INFUSION, INSULIN, TO BE USED WITH INVASIVE GLUCOSE SENSOR | MEDTRONIC PUERTO RICO OPERATIONS CO. | MMT-723LNAS | A7723LNASJ | OYC |  | 20200805 | Device was returned to manufacturer | LAY USER/PATIENT | Pump, Infusion, Insulin, To Be Used With Invasive Glucose Sensor
Malfunction | 20200425 | ['Battery Problem', 'Power Problem'] | INFORMATION RECEIVED BY MEDTRONIC INDICATED THAT THE INSULIN PUMP HAD REPLACE BATTERY ALARM. NO HARM REQUIRING MEDICAL INTERVENTION WAS REPORTED. THE INSULIN PUMP WILL BE RETURNED FOR ANALYSIS. |  | (B)(4). CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. THE DEVICE WILL BE RETURNED FOR ANALYSIS AND FURTHER INFORMATION WILL FOLLOW ONCE THE ANALYSIS HAS BEEN COMPLETED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. DEVICE POWERED UP PROPERLY AFTER BATTERY INSTALLATION. NO BLANK DISPLAY NOTED. DEVICE PASSED THE SELF TEST, SLEEP CURRENT MEASUREMENT, ACTIVE CURRENT MEASUREMENT, AND THE DISPLACEMENT TEST. THE ADAPT TOOL INDICATED ABNORMAL BEHAVIOR OF THE LITHIUM BACKUP BATTERY LOADED VOLTAGE AS SHOWN ON THE POWER MANAGEMENT GRAPH AND CONFIRMED LOW BATTERY AND POWER LOSS ALARMS DUE TO CONNECTOR RESISTANCE ELECTRICAL BOARD. AFTER DISCONNECTING AND RECONNECTING THE INTERNAL BATTERY CONNECTOR ON ELECTRICAL BOARD, DEVICE WAS MONITORED AND FUNCTIONED PROPERLY. NO UNEXPECTED REPLACE BATTERY NOW OR BATTERY FAILED ALARMS NOTED. DEVICE WAS RECEIVED WITHOUT THE ORIGINAL BATTERY CAP. UNABLE TO VERIFY BATTERY CAP DAMAGE DUE TO MISSING BATTERY CAP. | ARTIFICIAL PANCREAS DEVICE SYSTEM, THRESHOLD SUSPEND | MEDTRONIC PUERTO RICO OPERATIONS CO. | MMT-1715K | HG1A39R | OZO |  | 20200507 | Device was returned to manufacturer | LAY USER/PATIENT | Automated Insulin Dosing , Threshold Suspend
Malfunction | 20200210 | ['Battery Problem'] | INFORMATION RECEIVED BY MEDTRONIC INDICATED THAT THE INSULIN PUMP HAD BATTERY ERROR. NO HARM REQUIRING MEDICAL INTERVENTION WAS REPORTED. THE INSULIN PUMP WILL BE RETURNED FOR ANALYSIS. |  | UNIT RECEIVED WITH MISSING RETAINER, MISSING RESERVOIR TUBE O-RING AND BROKEN RESERVOIR TUBE LIP. UNABLE TO PERFORM DISPLACEMENT TEST OR LOCK RESERVOIR IN PLACE DUE TO MISSING RETAINER. UNIT PASSED THE SLEEP CURRENT MEASUREMENT AND ACTIVE CURRENT MEASUREMENT TEST. NO LOW BATTERY ALARMS OR UNEXPECTED BATTERY POWER LOSS NOTED DURING TESTING. UNIT FUNCTIONING PROPERLY. THE POWER MANAGEMENT PARAMETERS GRAPH CONFIRMED THE UNLOADED VOLTAGE (UNLOADED VLITH) AND LOADED VOLTAGE (LOADED VLITH) WAS WITHIN SPEC RANGE. PUMP RECEIVED WITH NORMAL OPERATING CURRENTS. (B)(4). THIS REPORT IS PART OF A RETROSPECTIVE REVIEW AND REMEDIATION EFFORTS IN RESPONSE TO A WARNING LETTER. (B)(4). MEDTRONIC, INC. (MEDTRONIC) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY MEDTRONIC, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. MEDTRONIC HAS MADE REASONABLE EFFORTS TO OBTAIN MORE COMPLETE INFORMATION IN THE TIME ALLOTTED AND HAS PROVIDED AS MUCH INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE AN ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY "DEFECTS" OR HAS "MALFUNCTIONED". THESE WORDS ARE INCLUDED IN THE FDA 3500A FORM AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. MEDTRONIC OBJECTS TO THE USE OF THESE WORDS AND OTHERS LIKE IT BECAUSE OF THE LACK OF DEFINITION AND THE CONNOTATIONS IMPLIED BY THESE TERMS. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT. | PUMP, INFUSION, INSULIN, TO BE USED WITH INVASIVE GLUCOSE SENSOR | MEDTRONIC PUERTO RICO OPERATIONS CO. | MMT-1712P | HG0XFWR | OYC |  | 20200310 | Device was returned to manufacturer | LAY USER/PATIENT | Pump, Infusion, Insulin, To Be Used With Invasive Glucose Sensor
Malfunction | 20200210 | ['Battery Problem', 'Power Problem'] | INFORMATION RECEIVED BY MEDTRONIC INDICATED THAT THE BATTERY OF INSULIN PUMP DID NOT LAST LONGER. THE CUSTOMER WAS ALSO REPORTED THAT BATTERY COMPARTMENT WAS NEITHER DAMAGED NOR CORRODED, BATTERY CAP WAS NEITHER MISSING NOR DAMAGED, AND THE SPRING WAS NEITHER DAMAGED NOR CORRODED. NO HARM REQUIRING MEDICAL INTERVENTION WAS REPORTED. THE INSULIN PUMP WILL BE RETURNED FOR ANALYSIS. |  | THE INSULIN PUMP PASSED THE DISPLACEMENT TEST, ACTIVE CURRENT TEST AND SELF TEST. PUMP FAILED THE SLEEP CURRENT TEST DUE TO MOISTURE DAMAGE ON THE ELECTRONIC ASSEMBLY. CHARGE, BATTERY LASTS LESS THAN EXPECTED CONFIRMED. FAILED BATTERY TEST CONFIRMED. PUMP RECEIVED WITH CORRODED BATTERY TUBE. (B)(4). THIS REPORT IS PART OF A RETROSPECTIVE REVIEW AND REMEDIATION EFFORTS IN RESPONSE TO A WARNING LETTER. UPDATED H9: Z-0956-2020. MEDTRONIC, INC. (MEDTRONIC) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY MEDTRONIC, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. MEDTRONIC HAS MADE REASONABLE EFFORTS TO OBTAIN MORE COMPLETE INFORMATION IN THE TIME ALLOTTED AND HAS PROVIDED AS MUCH INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE AN ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY "DEFECTS" OR HAS "MALFUNCTIONED". THESE WORDS ARE INCLUDED IN THE FDA 3500A FORM AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. MEDTRONIC OBJECTS TO THE USE OF THESE WORDS AND OTHERS LIKE IT BECAUSE OF THE LACK OF DEFINITION AND THE CONNOTATIONS IMPLIED BY THESE TERMS. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT. | ARTIFICIAL PANCREAS DEVICE SYSTEM, SINGLE HORMONAL CONTROL | MEDTRONIC PUERTO RICO OPERATIONS CO. | MMT-1780KPK | HG2FH8F | OZP | 20210404 | 20200309 | Device was returned to manufacturer | LAY USER/PATIENT | Automated Insulin Dosing Device System, Single Hormonal Control
Malfunction | 20200212 | ['Battery Problem', 'Power Problem'] | INFORMATION RECEIVED BY MEDTRONIC INDICATED THAT THE INSULIN PUMP HAD REPLACE BATTERY NOW ALARM. CUSTOMER DID RECEIVED LOW BATTERY ALERT ALARM PRIOR TO REPLACE BATTERY NOW ALARM. THE CUSTOMER STATES THE BATTERY WAS NOT PREVIOUSLY USED. NO HARM REQUIRING MEDICAL INTERVENTION WAS REPORTED. THE INSULIN PUMP WILL BE RETURNED FOR ANALYSIS. |  | INSULIN PUMP PASSED SLEEP CURRENT MEASUREMENT, ACTIVE CURRENT MEASUREMENT AND SELF TEST. INSULIN PUMP WAS MONITORED AND TESTED WITH NO UNEXPECTED BATTERY POWER LOSS AND LOW BATTERY ALERT NOTED. PUMP RECEIVED WITH DETACHED, MISSING RESERVOIR RETAINER RING, BROKEN RESERVOIR TUBE LIP AND MISSING RESERVOIR TUBE LIP O-RING NOTED. THE TEST RESERVOIR DOES NOT STAY LOCKED IN PLACE DUE TO DETACHED, MISSING RETAINER. UNABLE TO PERFORM DISPLACEMENT TEST DUE TO DETACHED, MISSING RETAINER. (B)(4). THIS REPORT IS PART OF A RETROSPECTIVE REVIEW AND REMEDIATION EFFORTS IN RESPONSE TO A WARNING LETTER. UPDATED H9: Z-0956-2020. MEDTRONIC, INC. (MEDTRONIC) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY MEDTRONIC, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. MEDTRONIC HAS MADE REASONABLE EFFORTS TO OBTAIN MORE COMPLETE INFORMATION IN THE TIME ALLOTTED AND HAS PROVIDED AS MUCH INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE AN ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY "DEFECTS" OR HAS "MALFUNCTIONED". THESE WORDS ARE INCLUDED IN THE FDA 3500A FORM AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. MEDTRONIC OBJECTS TO THE USE OF THESE WORDS AND OTHERS LIKE IT BECAUSE OF THE LACK OF DEFINITION AND THE CONNOTATIONS IMPLIED BY THESE TERMS. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT. | ARTIFICIAL PANCREAS DEVICE SYSTEM, SINGLE HORMONAL CONTROL | MEDTRONIC PUERTO RICO OPERATIONS CO. | MMT-1780KPK | HG234RF | OZP |  | 20200305 | Device was returned to manufacturer | LAY USER/PATIENT | Automated Insulin Dosing Device System, Single Hormonal Control
Malfunction | 20200206 | ['Battery Problem', 'Power Problem'] | INFORMATION RECEIVED BY MEDTRONIC INDICATED THAT THE INSULIN PUMP RECEIVED REPLACE BATTERY ALERT. CUSTOMER STATED THEY DID RECEIVE LOW BATTERY ALERT PRIOR TO REPLACE BATTERY ALERT LESS THAN 8 HOURS. CUSTOMER STATED BATTERY WAS NOT PREVIOUSLY USED. CUSTOMER STATED BATTERY CAP WAS NOT LOOSE, CRACKED OR DAMAGED. CUSTOMER STATED THIS WAS THE SECOND OCCURRENCE OF OFF NO POWER WITH A LOW BATTERY WARNING. NO HARM REQUIRING MEDICAL INTERVENTION WAS REPORTED. THE INSULIN PUMP WILL BE RETURNED FOR ANALYSIS. |  | INSULIN PUMP PASSED THE SELF TEST, DISPLACEMENT TEST AND ACTIVE CURRENT MEASUREMENT. HOWEVER, THE PUMP FAILED THE SLEEP CURRENT MEASUREMENT DUE TO A PROBLEM ISOLATED ON THE ELECTRONIC ASSEMBLY. (B)(4). | ARTIFICIAL PANCREAS DEVICE SYSTEM, SINGLE HORMONAL CONTROL | MEDTRONIC PUERTO RICO OPERATIONS CO. | MMT-1780KPK | HG3ZGGW | OZP |  | 20200305 | Device was returned to manufacturer | LAY USER/PATIENT | Automated Insulin Dosing Device System, Single Hormonal Control
Malfunction | 20200425 | ['Battery Problem'] | INFORMATION RECEIVED BY MEDTRONIC INDICATED THAT THE INSULIN PUMP HAD A POWER ERROR DETECTED ALARM. THE CUSTOMER STATED THAT POWER ERROR DETECTED RECURRED WITHIN A WEEK OF TROUBLESHOOT PREVIOUS OCCURRENCE. NO HARM REQUIRING MEDICAL INTERVENTION WAS REPORTED. THE INSULIN PUMP WILL BE RETURNED FOR ANALYSIS. |  | (B)(4). . CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. THE DEVICE WILL BE RETURNED FOR ANALYSIS AND FURTHER INFORMATION WILL FOLLOW ONCE THE ANALYSIS HAS BEEN COMPLETED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. DEVICE PASSED THE SELF TEST, SLEEP CURRENT MEASUREMENT, ACTIVE CURRENT MEASUREMENT, AND THE DISPLACEMENT TEST. DEVICE TRACE DOWNLOAD ANALYSIS CONFIRMED THE INSULIN PUMP ALARMED POWER ERROR 25. THE ADAPT TOOL CONFIRMED POWER ERROR 25 ALARMS WERE TRIGGERED WHEN THE BATTERY LOADED VOLTAGE (LOADED VLITH) WAS LESS THAN 3.5V FOR 4 CONSECUTIVE HOURS DUE TO CONNECTOR RESISTANCE ELECTRICAL BOARD. AFTER DISCONNECTING AND RECONNECTING THE INTERNAL BATTERY CONNECTOR ON ELECTRICAL BOARD, DEVICE WAS MONITORED AND FUNCTIONED PROPERLY. | ARTIFICIAL PANCREAS DEVICE SYSTEM, THRESHOLD SUSPEND | MEDTRONIC PUERTO RICO OPERATIONS CO. | MMT-1715K | HG1HN9X | OZO |  | 20200504 | Device was returned to manufacturer | LAY USER/PATIENT | Automated Insulin Dosing , Threshold Suspend
Malfunction | 20200424 | ['Display or Visual Feedback Problem', 'Obstruction of Flow', 'Battery Problem'] | THE CUSTOMER REPORTED VIA PHONE CALL THAT THEIR INSULIN PUMP SCREEN WAS BLURRY, HARD TO READ AND THE PUMP WAS SQUIRTING OUT INSULIN WHILE PRIMING. THE CUSTOMER¿S BLOOD GLUCOSE LEVEL WAS UNKNOWN AT THE TIME OF THE INCIDENT. TROUBLESHOOTING WAS NOT PERFORMED. THE PUMP WILL NOT BE RETURNING FOR ANALYSIS. |  | CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. | PUMP, INFUSION, INSULIN, TO BE USED WITH INVASIVE GLUCOSE SENSOR | MEDTRONIC MINIMED | MMT-XXX |  | OYC |  |  | No | LAY USER/PATIENT | Pump, Infusion, Insulin, To Be Used With Invasive Glucose Sensor
Malfunction | 20200219 | ['Battery Problem'] | INFORMATION RECEIVED BY MEDTRONIC INDICATED THAT THE INSULIN PUMP BATTERY ISSUES. IT WAS REPORTED THAT THE LOW BATTERY ISSUE WAS RECURRING AFTER THE TROUBLESHOOTING. NO HARM REQUIRING MEDICAL INTERVENTION WAS REPORTED. THE INSULIN PUMP WILL BE RETURNED FOR ANALYSIS. |  | UNIT RECEIVED WITH HIGH SLEEP CURRENT MEASUREMENT TEST. PROBLEM ISOLATED TO ELECTRONIC ASSEMBLY. (B)(4). | ARTIFICIAL PANCREAS DEVICE SYSTEM, SINGLE HORMONAL CONTROL | MEDTRONIC PUERTO RICO OPERATIONS CO. | MMT-1780KPK | HG435GE | OZP |  | 20200309 | Device was returned to manufacturer | LAY USER/PATIENT | Automated Insulin Dosing Device System, Single Hormonal Control
Malfunction | 20191201 | ['Battery Problem'] | INFORMATION RECEIVED BY MEDTRONIC INDICATED THAT THE INSULIN PUMP HAD POWER ERROR DETECTED ALARM. THE CUSTOMER WAS ABLE TO CLEAR AND REWIND THE INSULIN PUMP. NOTIFICATION LIGHT DID NOT IMMEDIATELY STOP FLASHING. NO HARM REQUIRING MEDICAL INTERVENTION WAS REPORTED. THE INSULIN PUMP WILL BE RETURNED FOR ANALYSIS. |  | INSULIN PUMP TRACE DOWNLOAD ANALYSIS CONFIRMED THE PUMP ALARMED POWER ERROR DETECTED ALARM. THE POWER MANAGEMENT TOOL INDICATED ABNORMAL BEHAVIOR OF THE LITHIUM BACKUP BATTERY LOADED VLITH VOLTAGE AS SHOWN ON THE POWER MANAGEMENT GRAPH AND CONFIRMED POWER ERROR DETECTED ALARM DUE TO CONNECTOR RESISTANCE J6 /PCB 1. AFTER DISCONNECTING AND RECONNECTING THE INTERNAL BATTERY CONNECTOR ON J6/PCB 1, THE PUMP WAS MONITORED AND FUNCTIONED PROPERLY. (B)(4). THIS REPORT IS PART OF A RETROSPECTIVE REVIEW AND REMEDIATION EFFORTS IN RESPONSE TO A WARNING LETTER. UPDATED H9: Z-0955-2020. MEDTRONIC, INC. (MEDTRONIC) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY MEDTRONIC, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. MEDTRONIC HAS MADE REASONABLE EFFORTS TO OBTAIN MORE COMPLETE INFORMATION IN THE TIME ALLOTTED AND HAS PROVIDED AS MUCH INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE AN ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY "DEFECTS" OR HAS "MALFUNCTIONED". THESE WORDS ARE INCLUDED IN THE FDA 3500A FORM AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. MEDTRONIC OBJECTS TO THE USE OF THESE WORDS AND OTHERS LIKE IT BECAUSE OF THE LACK OF DEFINITION AND THE CONNOTATIONS IMPLIED BY THESE TERMS. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT. | ARTIFICIAL PANCREAS DEVICE SYSTEM, THRESHOLD SUSPEND | MEDTRONIC PUERTO RICO OPERATIONS CO. | MMT-1715K | HG2GU33 | OZO | 20210419 | 20200306 | Device was returned to manufacturer | LAY USER/PATIENT | Automated Insulin Dosing , Threshold Suspend
Malfunction | 20200303 | ['Battery Problem', 'Power Problem'] | INFORMATION RECEIVED BY MEDTRONIC INDICATED THAT THE INSULIN PUMP HAD REPLACE BATTERY NOW WITHOUT LOW BATTERY ALERTS. NO HARM REQUIRING MEDICAL INTERVENTION WAS REPORTED. THE INSULIN PUMP WAS RETURNED FOR ANALYSIS. |  | UNIT RECEIVED WITH BLANK DISPLAY DUE TO CORRODED ELECTRONICS ASSEMBLY. MOTOR AND FORCE SENSOR WAS ALSO CORRODED. UNABLE TO PERFORM DISPLACEMENT TEST, SLEEP CURRENT MEASUREMENT TEST, ACTIVE CURRENT MEASUREMENT TEST AND SELF TEST DUE TO BLANK DISPLAY. UNABLE TO VERIFY POWER LOSS ALARM OR LOW BATTERY ALERT DUE TO BLANK DISPLAY. (B)(4). THIS REPORT IS PART OF A RETROSPECTIVE REVIEW AND REMEDIATION EFFORTS IN RESPONSE TO A WARNING LETTER. UPDATED H9: Z-0955-2020. MEDTRONIC, INC. (MEDTRONIC) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY MEDTRONIC, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. MEDTRONIC HAS MADE REASONABLE EFFORTS TO OBTAIN MORE COMPLETE INFORMATION IN THE TIME ALLOTTED AND HAS PROVIDED AS MUCH INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE AN ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY "DEFECTS" OR HAS "MALFUNCTIONED". THESE WORDS ARE INCLUDED IN THE FDA 3500A FORM AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. MEDTRONIC OBJECTS TO THE USE OF THESE WORDS AND OTHERS LIKE IT BECAUSE OF THE LACK OF DEFINITION AND THE CONNOTATIONS IMPLIED BY THESE TERMS. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT. | ARTIFICIAL PANCREAS DEVICE SYSTEM, THRESHOLD SUSPEND | MEDTRONIC PUERTO RICO OPERATIONS CO. | MMT-1715K | HG19KMU | OZO |  | 20200313 | Device was returned to manufacturer | LAY USER/PATIENT | Automated Insulin Dosing , Threshold Suspend
Malfunction | 20200303 | ['Battery Problem', 'No Apparent Adverse Event'] | INFORMATION RECEIVED BY MEDTRONIC INDICATED THAT THE INSULIN PUMP HAD LOW BATTERY ALARM. NO HARM REQUIRING MEDICAL INTERVENTION WAS REPORTED. THE INSULIN PUMP WILL BE RETURNED FOR ANALYSIS. |  | UNIT PASSED DISPLACEMENT TEST, SELF TEST, SLEEP CURRENT MEASUREMENT AND ACTIVE CURRENT MEASUREMENT. HOWEVER, UNIT RECEIVED WITH LONG TRACE DISPLAYS CHARGE, BATTERY LASTS LESS THAN EXPECTED AND PUMP RECEIVED ERROR DETECTED ALARM DUE TO J6 CONNECTOR RESISTANCE ISSUE. (B)(4). | ARTIFICIAL PANCREAS DEVICE SYSTEM, THRESHOLD SUSPEND | MEDTRONIC PUERTO RICO OPERATIONS CO. | MMT-1715KM | HG1JMSN | OZO |  | 20200312 | Device was returned to manufacturer | LAY USER/PATIENT | Automated Insulin Dosing , Threshold Suspend
Malfunction | 20190423 | ['Mechanical Problem', 'Battery Problem', 'Moisture or Humidity Problem'] | INFORMATION RECEIVED BY MEDTRONIC INDICATED THAT THE INSULIN PUMP HAD ISSUE WITH PISTON MISLINED AND HAD MOTOR ERROR. THE CUSTOMER STATED THAT RESERVOIR HAD LEAK. THE CUSTOMER STATED THAT DRIVE SUPPORT CAP WAS NORMAL. THE CUSTOMER STATED THEY WERE ABLE TO CLEAR THE ALARM AND ABLE TO COMPLETE THE REWOUND PROCESS. NO HARM REQUIRING MEDICAL INTERVENTION WAS REPORTED. THE DEVICE WILL NOT BE RETURNED FOR ANALYSIS. |  | (B)(4). CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. | ARTIFICIAL PANCREAS DEVICE SYSTEM, THRESHOLD SUSPEND | MEDTRONIC PUERTO RICO OPERATIONS CO. | MMT-751NAS | A4751NASJ | OZO |  |  | No | LAY USER/PATIENT | Automated Insulin Dosing , Threshold Suspend
Injury | 20200426 | ['Battery Problem', 'Power Problem'] | CUSTOMER REPORTED VIA PHONE CALL THAT THEY EXPERIENCED HIGH BLOOD GLUCOSE. THE CUSTOMER BLOOD GLUCOSE LEVEL WAS 450 MG/DL AT THE TIME OF INCIDENT. CUSTOMER REPORTED THAT THE THEY RECEIVED A LOW BATTERY ALERT PRIOR TO THE REPLACE BATTERY ALERT. CUSTOMER STATED IT WAS LESS THAN 8 HOURS FROM THE LOW BATTERY ALERT TO THE LOW BATTERY ALERT. THE CUSTOMER STATED THE BATTERY WAS NOT PREVIOUSLY USED. CUSTOMER STATED THE BATTERY CAP WAS NOT LOOSE, CRACKED OR DAMAGED. CUSTOMER STATES THIS IS THE SECOND OCCURRENCE OF REPLACE BATTERY ALERT IN LESS THAN 8 HOURS AFTER LOW BATTERY WARNING. THE INSULIN PUMP WILL BE RETURNED FOR ANALYSIS. |  | (B)(4). CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. THE DEVICE WILL BE RETURNED FOR ANALYSIS AND FURTHER INFORMATION WILL FOLLOW ONCE THE ANALYSIS HAS BEEN COMPLETED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. | ARTIFICIAL PANCREAS DEVICE SYSTEM, THRESHOLD SUSPEND | MEDTRONIC PUERTO RICO OPERATIONS CO. | MMT-1715K | HG2JHJP | OZO | 20210513 |  | No | LAY USER/PATIENT | Automated Insulin Dosing , Threshold Suspend
Malfunction | 20200424 | ['Battery Problem'] | THE CUSTOMER REPORTED VIA PHONE CALL THE INSULIN PUMP HAD ALARM GENERATED WHEN A POWER SYSTEM ERROR (BACK UP BATTERY FAILS) DETECTED AND THIS ERROR DOES NOT PREVENT THE PUMP FROM RUNNING. CUSTOMER CANNOT RECALL BLOOD GLUCOSE AT THE TIME OF INCIDENT. TROUBLESHOOTING WAS PERFORMED. CUSTOMER REPORTS PUMP HAS NOT BEEN EXPOSED TO TEMPERATURES BELOW 41°F (5°C). CUSTOMER CALLED BACK AND THE ISSUE REOCCURRED. THE INSULIN PUMP WILL BE RETURNING FOR ANALYSIS. |  | THIS REPORT IS PART OF A RETROSPECTIVE REVIEW AND REMEDIATION EFFORTS IN RESPONSE TO A WARNING LETTER. UPDATED H9: Z-0955-2020. MEDTRONIC, INC. (MEDTRONIC) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY MEDTRONIC, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. MEDTRONIC HAS MADE REASONABLE EFFORTS TO OBTAIN MORE COMPLETE INFORMATION IN THE TIME ALLOTTED AND HAS PROVIDED AS MUCH INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE AN ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY "DEFECTS" OR HAS "MALFUNCTIONED". THESE WORDS ARE INCLUDED IN THE FDA 3500A FORM AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. MEDTRONIC OBJECTS TO THE USE OF THESE WORDS AND OTHERS LIKE IT BECAUSE OF THE LACK OF DEFINITION AND THE CONNOTATIONS IMPLIED BY THESE TERMS. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT. DEVICE PASSED SLEEP CURRENT MEASUREMENT, ACTIVE CURRENT MEASUREMENT AND SELF TEST. PUMP ERROR 25 FOUND IN HISTORY FILE DUE TO CONNECTOR RESISTANCE ISSUE. DEVICE RECEIVED WITH DETACHED OR MISSING RESERVOIR RETAINER RING AND MISSING RESERVOIR TUBE LIP O-RING NOTED. THE TEST RESERVOIR DOES NOT STAY LOCKED IN PLACE DUE TO DETACHED OR MISSING RETAINER. UNABLE TO PERFORM DISPLACEMENT TEST DUE TO DETACHED OR MISSING RETAINER. (B)(4). CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. THE DEVICE WILL BE RETURNED FOR ANALYSIS AND FURTHER INFORMATION WILL FOLLOW ONCE THE ANALYSIS HAS BEEN COMPLETED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. | ARTIFICIAL PANCREAS DEVICE SYSTEM, THRESHOLD SUSPEND | MEDTRONIC PUERTO RICO OPERATIONS CO. | MMT-1715K | HG1EY42 | OZO |  | 20200430 | Device was returned to manufacturer | LAY USER/PATIENT | Automated Insulin Dosing , Threshold Suspend
Malfunction | 20200324 | ['Device Alarm System', 'Battery Problem'] | INFORMATION WAS RECEIVED INDICATING THAT SMITHS MEDICAL MEDFUSION 4000 PUMP HAS MALFUNCTIONED. THE PUMP QUIT ABRUPTLY. THE DEVICE HISTORY LOG SHOWED THAT THE PUMP DISPLAYED A LOW BATTERY (7%), THEN SYSTEM FAILURE PRIMARY AUDIBLE BACKGROUND TEST AND PROGRESSED TO ACU WATCHDOG ALARM. THERE WERE NO ADVERSE EVENTS REPORTED. |  |  | PUMP, INFUSION | ST PAUL | 4000-0105-51 |  | FRN |  |  | No | HEALTH PROFESSIONAL | Pump, Infusion
Malfunction |  | ['Battery Problem'] | PER BIOMED , BATTERY IS DEAD. UNIT WON'T POWER UP ON BATTERY. (B)(6) 2020 - SAMPLE EVALUATION CONFIRMED BATTERY ABRUPT SHUTDOWN. |  | THE FOLLOWING WERE REVIEWED AS PART OF THIS INVESTIGATION: PATIENT SEVERITY, APPLICABLE PREVIOUS INVESTIGATION(S), SAMPLE(IF AVAILABLE), APPLICABLE FMEA DOCUMENTS, LABELING AND APPLICABLE MANUFACTURE RECORDS. BASED ON A REVIEW OF THIS INFORMATION, THE FOLLOWING WAS CONCLUDED: THE DEVICE WAS RETURNED TO THE SERVICE FACILITY FOR EVALUATION. DURING EVALUATION, THE REPORTED ISSUE OF UNIT WON'T POWER UP ON BATTERY WAS CONFIRMED. THE SCANNER WAS INSPECTED AND FULLY CHARGED THE UNIT ABRUPTLY SHUTS OFF WHEN UNPLUGGED FROM THE AC ADAPTER. THE ROOT CAUSE OF THE REPORTED ISSUE IS DUE TO NO INTERNAL BATTERY INSIDE THE SCANNER, LEADING THE SCANNER TO SHUT OFF WHEN UNPLUGGED. THE DEVICE WAS SERVICED, TESTED AND RETURNED TO THE CUSTOMER. A HISTORY REVIEW OF SERIAL NUMBER (B)(4) SHOWED NO OTHER SIMILAR PRODUCT COMPLAINT(S) FROM THIS SERIAL NUMBER. | SYSTEM, IMAGING, PULSED ECHO, ULTRASONIC | BARD ACCESS SYSTEMS | 9770501 | N/A | IYO |  | 20191127 | Device was returned to manufacturer | HEALTH PROFESSIONAL | System, Imaging, Pulsed Echo, Ultrasonic
Malfunction | 20191211 | ['Battery Problem'] | A NISUS PUMP WAS RETURNED WITH A COMPLAINT ALLEGING THAT THE BATTERY WAS NO LONGER HOLDING A CHARGE BUT WOULD OPERATE FROM THE CHARGER CABLE AS INTENDED. NO INJURIES OR HARM WERE REPORTED. UPON INSPECTION, IT WAS DISCOVERED THAT THE BATTERY HAD EXPERIENCED A SHORT-CIRCUIT ON THE CONTROL BOARD, CAUSING THERMAL DAMAGE TO THE HOUSING INTERIOR. NO FAULT WAS FOUND WITH THE DEVICE OR ITS CIRCUITRY. THE NISUS UNIT FUNCTIONED AS DESIGNED, WHEN POWERED UTILIZING AN EXTERNAL POWER SOURCE. NO ELECTRICAL ANOMALIES OR FAULTS CAN BE CONFIRMED ASSOCIATED WITH THE NISUS DEVICE. ANALYSIS OF THE EVENT BY THE SUPPLIER OF THE ASSOCIATED LITHIUM BATTERY CONFIRMS THE FAULT WAS LIKELY THE RESULT OF A SHORT ON THE BATTERY CONTROL BOARD. THIS EVENT HAS BEEN REPORTED TO THE MANUFACTURER OF THE ASSOCIATED LITHIUM BATTERIES FOR FURTHER ANALYSIS. |  | (B)(4). | NEGATIVE PRESSURE WOUND THERAPY POWERED SUCTION PUMP | CORK MEDICAL, LLC | CMPP-100 |  | OMP | 20230907 | 20191212 | Device was returned to manufacturer | I | Negative Pressure Wound Therapy Powered Suction Pump
Injury | 20200416 | ['Battery Problem'] | IT WAS REPORTED THAT THE PATIENT WAS EXPERIENCING HEART SYMPTOMS AS THERE WAS PREMATURE BATTERY DEPLETION FOR THE PATIENTS IMPLANTABLE PULSE GENERATOR (IPG). THE IPG WAS REMOVED AND REPLACED. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. | ON 2020-05-11: IT WAS FURTHER CLARIFIED THAT THE PATIENT EXPERIENCED PACEMAKER SYNDROME AND CONGESTIVE HEART FAILURE. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. | IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. PRODUCT EVENT SUMMARY: THE DEVICE WAS RETURNED AND ANALYZED. THE DEVICE MET 80% OR GREATER OF EXPECTED LONGEVITY. MEDTRONIC IS SUBMITTING THIS REPORT TO COMPLY WITH FDA REPORTING REGULATIONS UNDER 21 CFR PARTS 4 AND 803. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY MEDTRONIC, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. MEDTRONIC HAS MADE REASONABLE EFFORTS TO OBTAIN MORE COMPLETE INFORMATION AND HAS PROVIDED AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE AN ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY ¿DEFECTS¿ OR HAS ¿MALFUNCTIONED¿. THESE WORDS ARE INCLUDED IN THE FDA 3500A FORM AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REGULATORY REPORTING. MEDTRONIC OBJECTS TO THE USE OF THESE WORDS AND OTHERS LIKE THEM BECAUSE OF THE LACK OF DEFINITION AND THE CONNOTATIONS IMPLIED BY THESE TERMS. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. A GOOD FAITH EFFORT WILL BE MADE TO OBTAIN THE APPLICABLE INFORMATION RELEVANT TO THE REPORT. IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. | PULSE GENERATOR, PERMANENT, IMPLANTABLE | MEDTRONIC PUERTO RICO OPERATIONS CO. | ADDR01 |  | NVZ | 20170328 | 20200427 | Device was returned to manufacturer | LAY USER/PATIENT | Pulse Generator, Permanent, Implantable
Malfunction | 20200120 | ['Battery Problem'] | A NISUS PUMP WAS RETURNED WITH A COMPLAINT ALLEGING THAT THE BATTERY WAS NO LONGER HOLDING A CHARGE BUT WOULD OPERATE FROM THE CHARGER CABLE AS INTENDED. NO INJURIES OR HARM WERE REPORTED. UPON INSPECTION, IT WAS DISCOVERED THAT THE BATTERY HAD EXPERIENCED A SHORT-CIRCUIT ON THE CONTROL BOARD, CAUSING THERMAL DAMAGE TO THE HOUSING INTERIOR. NO FAULT WAS FOUND WITH THE DEVICE OR ITS CIRCUITRY. THE NISUS UNIT FUNCTIONED AS DESIGNED, WHEN POWERED UTILIZING AN EXTERNAL POWER SOURCE. NO ELECTRICAL ANOMALIES OR FAULTS CAN BE CONFIRMED ASSOCIATED WITH THE NISUS DEVICE. ANALYSIS OF THE EVENT BY THE SUPPLIER OF THE ASSOCIATED LITHIUM BATTERY CONFIRMS THE FAULT WAS LIKELY THE RESULT OF A SHORT ON THE BATTERY CONTROL BOARD. THIS EVENT HAS BEEN REPORTED TO THE MANUFACTURER OF THE ASSOCIATED LITHIUM BATTERIES FOR FURTHER ANALYSIS. |  |  | NEGATIVE PRESSURE WOUND THERAPY POWERED SUCTION PUMP | CORK MEDICAL, LLC | CMPP-100 |  | OMP | 20230222 | 20200121 | Device was returned to manufacturer | I | Negative Pressure Wound Therapy Powered Suction Pump
Malfunction | 20200207 | ['Battery Problem'] | A NISUS PUMP WAS RETURNED WITH A COMPLAINT ALLEGING THAT THE BATTERY WAS NO LONGER HOLDING A CHARGE AND WOULD NOT OPERATE FROM THE CHARGER CABLE AS INTENDED. NO INJURIES OR HARM WERE REPORTED. UPON INSPECTION, IT WAS DISCOVERED THAT THE BATTERY HAD EXPERIENCED A SHORT-CIRCUIT ON THE CONTROL BOARD, CAUSING THERMAL DAMAGE TO THE HOUSING INTERIOR. NO FAULT WAS FOUND WITH THE DEVICE OR ITS CIRCUITRY. THE NISUS UNIT FUNCTIONED AS DESIGNED, WHEN POWERED UTILIZING AN EXTERNAL POWER SOURCE. NO ELECTRICAL ANOMALIES OR FAULTS CAN BE CONFIRMED ASSOCIATED WITH THE NISUS DEVICE. ANALYSIS OF THE EVENT BY THE SUPPLIER OF THE ASSOCIATED LITHIUM BATTERY CONFIRMS THE FAULT WAS LIKELY THE RESULT OF A SHORT ON THE BATTERY CONTROL BOARD. THIS EVENT HAS BEEN REPORTED TO THE MANUFACTURER OF THE ASSOCIATED LITHIUM BATTERIES FOR FURTHER ANALYSIS. |  | (B)(4). | NEGATIVE PRESSURE WOUND THERAPY POWERED SUCTION PUMP | CORK MEDICAL, LLC | CMPP-100 |  | OMP | 20240625 | 20200210 | Device was returned to manufacturer | I | Negative Pressure Wound Therapy Powered Suction Pump
Malfunction |  | ['Battery Problem'] | PER TM: THE SCANNER IS NOT HOLDING A CHARGE, ABRUPT SHUT OFF IF THE SCANNER IS UNPLUGGED. |  | THE FOLLOWING WERE REVIEWED AS PART OF THIS INVESTIGATION: PATIENT SEVERITY, APPLICABLE PREVIOUS INVESTIGATION(S), SAMPLE (IF AVAILABLE), APPLICABLE FMEA DOCUMENTS, LABELING, AND APPLICABLE MANUFACTURE RECORDS. BASED ON A REVIEW OF THIS INFORMATION, THE FOLLOWING WAS CONCLUDED: THE DEVICE WAS RETURNED TO THE SERVICE FACILITY FOR EVALUATION. DURING EVALUATION, THE REPORTED ISSUE OF THE SCANNER IS NOT HOLDING A CHARGE, ABRUPT SHUT OFF IF THE SCANNER IS UNPLUGGED IS CONFIRMED. THE ROOT CAUSE OF THE REPORTED ISSUE IS THE INTERNAL BATTERY FAILED. THE DEVICE WAS SERVICED, TESTED AND RETURNED TO THE CUSTOMER. A HISTORY REVIEW OF SERIAL NUMBER (B)(4) SHOWED NO OTHER SIMILAR PRODUCT COMPLAINT(S) FROM THIS SERIAL NUMBER. | SYSTEM, IMAGING, PULSED ECHO, ULTRASONIC | BARD ACCESS SYSTEMS | 9770600Y | NA | IYO |  | 20200414 | Device was returned to manufacturer | HEALTH PROFESSIONAL | System, Imaging, Pulsed Echo, Ultrasonic
Malfunction | 20200416 | ['Failure to Charge', 'Battery Problem'] | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S INTERNAL BATTERY WAS REPLACED TO ADDRESS THE ISSUE. |  |  | VENTILATOR, CONTINUOUS, FACILITY USE | PHILIPS RESPIRONICS INC. | 1054096 |  | CBK |  | 20200416 | Device was returned to manufacturer | HEALTH PROFESSIONAL | Ventilator, Continuous, Facility Use
Malfunction | 20200420 | ['Failure to Charge', 'Battery Problem'] | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, "SERVICE REQUIRED" CODES WERE FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S INTERNAL BATTERY WAS REPLACED TO ADDRESS THE ISSUE. |  |  | VENTILATOR, CONTINUOUS, FACILITY USE | PHILIPS RESPIRONICS INC. | 1054096 |  | CBK |  | 20200407 | Device was returned to manufacturer | HEALTH PROFESSIONAL | Ventilator, Continuous, Facility Use
Malfunction | 20200328 | ['Device Alarm System', 'Battery Problem'] | INFORMATION WAS RECEIVED INDICATING THAT SMITHS MEDICAL MEDFUSION 4000 PUMP HAS MALFUNCTIONED. THE PUMP DISPLAYED A SYSTEM FAILURE, A PRIMARY AUDIBLE ALARM BGND TEST ERROR AND A BIOMED PASSCODE REQUESTED MESSAGE. THE PUMP'S HISTORY SHOWED AN INTERMITTENT VOLUME OVER TIME INFUSION OF 3ML SYRINGE GAVE AN EMPTY ALARM, THEN SYSTEM FAILURE PRIMARY AUDIBLE ALARM BGND, THEN SYSTEM FAILURE ACU WATCHDOG FAILURE. THERE WERE NO ADVERSE EVENTS REPORTED |  |  | PUMP, INFUSION | ST PAUL | 4000-0105-51 |  | FRN |  |  | No | HEALTH PROFESSIONAL | Pump, Infusion
Malfunction | 20200417 | ['Failure to Charge', 'Battery Problem'] | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S INTERNAL BATTERY WAS REPLACED TO ADDRESS THE ISSUE. |  |  | VENTILATOR, CONTINUOUS, FACILITY USE | PHILIPS RESPIRONICS INC. | 1054096 |  | CBK |  | 20200407 | Device was returned to manufacturer | HEALTH PROFESSIONAL | Ventilator, Continuous, Facility Use
Malfunction |  | ['Battery Problem', 'Charging Problem'] | INFORMATION WAS RECEIVED BY A MANUFACTURING REPRESENTATIVE (REP) REGARDING A PATIENT WITH AN IMPLANTABLE NEUROSTIMULATOR (INS). INFORMATION WAS REPORTED THAT THE REP SAID THE PATIENT IS HAVING MORE DIFFICULTY CHARGING THEIR INS. THE REP TALKED TO THE PATIENT'S DAUGHTER TODAY. THE PATIENT/DAUGHTER CLAIMED THAT THE PATIENT IS HAVING TO CHARGE 2-3 TIMES PER DAY. THE REP READ A STRING OF TEXT MESSAGES FROM THE PATIENT/DAUGHTER WHICH PROVIDED THE FOLLOWING INFORMATION: THE IMPLANT IS NOT WORKING PROPERLY AND THE CONTROLLER IS NOT WORKING PROPERLY. A NEW CORD WAS SENT BY PATIENT SERVICES IN THE PAST. IT IS CONFUSING BECAUSE AT ONE POINT THE CHARGER AND THE IMPLANT SHOWED 100% 3-4 TIMES PER DAY TO CHARGE THE INS. SOMETIMES THE IMPLANT CHARGES TO 40% AND IT WON'T CHARGE ANYMORE. OVER THE PAST FEW DAYS THE PATIENT HAD BEEN CHARGING THE INS MORNING AND NIGHT. THE DAUGHTER WANTED TO HAVE ALL OF THE EXTERNAL COMPONENTS REPLACED. THE REP STATED THE PATIENT HAD A HIGH DENSITY AND NON-HIGH DENSITY SET OF PARAMETERS AND HE THOUGHT THE PATIENT WAS NO ON HIGH DENSITY SETTINGS CURRENTLY. THE CALLER PROVIDED THE FOLLOWING PARAMETERS: 6.9MA 500US 50HZ 5 MA 100US 50HZ IT WAS REVIEWED LIKE MORE INFORMATION WOULD BE NEEDED TO DETERMINE THE LIKELY CAUSE OF THE ISSUE, BUT IS DOES NOT SOUNDS LIKE A PROBLEM THAT WOULD BE RESOLVED BY REPLACING THE CONTROLLER AND RECHARGER. NO FURTHER COMPLICATIONS WERE REPORTED. NO PATIENT SYMPTOMS WERE REPORTED. | ADDITIONAL INFORMATION WAS REPORTED THAT THE REP MET WITH THE PATIENT ON TUESDAY, AND SHE REALIZED SHE WAS IN GROUP A WHICH HAS 3 PROGRAMS RUNNING ON MODIFIED HD WITH HIGH SETTINGS. SO IT IS POSSIBLE HER CHARGING REQUIREMENT IS HIGH BECAUSE OF THIS. THE REP REPROGRAMMED THE PATIENT TO, TWO MUCH LOWER PROGRAMS ON GROUP B WITH MORE CONVENTIONAL SETTINGS. THE PATIENT'S COVERAGE IS BETTER THAN IT HAS EVER BEEN ACCORDING TO THE PATIENT. THE REP IS ALSO WORKING WITH THE PATIENT'S DAUGHTER AND SHE WILL UPDATE THE REP ON TUESDAY TO GIVE HER A FULL WEEK TO EVALUATE THE CHARGING BURDEN AND PAIN RELIEF. NO FURTHER INFORMATION WAS REPORTED. ADDITIONAL INFORMATION WAS REPORTED THAT THE PATIENT IS DOING MUCH BETTER NOW AND SHE IS STILL ON MODIFIED HD, BUT USING TWO PROGRAMS INSTEAD OF THREE. SHE IS CHARGING ONCE A DAY, BUT IT HAS BEEN EXPLAINED TO THE PATIENT THAT GIVEN HER PROGRAMS THAT IT IS NORMAL AND SHE IS OKAY WITH THIS. NO FURTHER INFORMATION WAS REPORTED. | IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. MEDTRONIC IS SUBMITTING THIS REPORT TO COMPLY WITH FDA REPORTING REGULATIONS UNDER 21 CFR PARTS 4 AND 803. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY MEDTRONIC, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. MEDTRONIC HAS MADE REASONABLE EFFORTS TO OBTAIN MORE COMPLETE INFORMATION AND HAS PROVIDED AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE AN ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY ¿DEFECTS¿ OR HAS ¿MALFUNCTIONED¿. THESE WORDS ARE INCLUDED IN THE FDA 3500A FORM AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REGULATORY REPORTING. MEDTRONIC OBJECTS TO THE USE OF THESE WORDS AND OTHERS LIKE THEM BECAUSE OF THE LACK OF DEFINITION AND THE CONNOTATIONS IMPLIED BY THESE TERMS. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. A GOOD FAITH EFFORT WILL BE MADE TO OBTAIN THE APPLICABLE INFORMATION RELEVANT TO THE REPORT. IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. | STIMULATOR, SPINAL-CORD, TOTALLY IMPLANTED FOR PAIN RELIEF | MEDTRONIC PUERTO RICO OPERATIONS CO. | 97715 |  | LGW | 20200514 |  | No | HEALTH PROFESSIONAL | Stimulator, Spinal-Cord, Totally Implanted For Pain Relief
Malfunction |  | ['Battery Problem'] | DURING EVALUATION REVIEW IT WAS FOUND THAT THE SR8 RAN ON BATTERY POWER FOR AROUND TEN MINUTES BEFORE SHUTTING DOWN WITH 72% BATTERY LIFE STILL ON THE SCANNER. |  | THE FOLLOWING WERE REVIEWED AS PART OF THIS INVESTIGATION: PATIENT SEVERITY, APPLICABLE PREVIOUS INVESTIGATION(S), SAMPLE (IF AVAILABLE), APPLICABLE FMEA DOCUMENTS, LABELING, AND APPLICABLE MANUFACTURE RECORDS. BASED ON A REVIEW OF THIS INFORMATION, THE FOLLOWING WAS CONCLUDED: THE DEVICE WAS RETURNED TO THE SERVICE FACILITY FOR EVALUATION. DURING EVALUATION, THE REPORTED ISSUE OF ABRUPT SHUTDOWN WAS IDENTIFIED AND CONFIRMED. THE ROOT CAUSE OF THE REPORTED ISSUE WAS IDENTIFIED AS AN INTERNAL BATTERY FAILURE, LEADING THE SCANNER TO ABRUPTLY SHUTTING OFF WHEN RAN ON BATTERY. THE DEVICE WAS SERVICED, TESTED, AND RETURNED TO THE CUSTOMER. A HISTORY REVIEW OF SERIAL NUMBER (B)(4) SHOWED ONE OTHER SIMILAR COMPLAINT FROM THIS SERIAL NUMBER. THE PREVIOUS COMPLAINT FOR THIS SERIAL NUMBER ((B)(4)) EVALUATION WAS CONFIRMED (REFERENCE: MDR NUMBER 3006260740- 2018-02976). | SYSTEM, IMAGING, PULSED ECHO, ULTRASONIC | BARD ACCESS SYSTEMS | 9770600Y | N/A | IYO |  | 20191118 | Device was returned to manufacturer | HEALTH PROFESSIONAL | System, Imaging, Pulsed Echo, Ultrasonic
Malfunction | 20200420 | ['Battery Problem'] | LOW BATTERY. |  |  | THERMOMETER, ELECTRONIC, CLINICAL | WELCH ALLYN, INC. | 901053 |  | FLL |  |  | Yes | I | Thermometer, Electronic, Clinical
Malfunction | 20200415 | ['Failure to Charge', 'Battery Problem'] | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, A "SERVICE REQUIRED" CODE WAS FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S INTERNAL BATTERY WAS REPLACED TO ADDRESS THE ISSUE. |  |  | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS INC | 1054096 |  | CBK |  | 20200415 | Device was returned to manufacturer | HEALTH PROFESSIONAL | Ventilator, Continuous, Facility Use
Malfunction |  | ['Loss of Power', 'Battery Problem'] | THE INDICATOR FOR THE BATTERY SHOWS FULL. THE MACHINE DIES AFTER 3-5 MIN WITHOUT CHARGING. | THE INDICATOR FOR THE BATTERY SHOWS FULL. THE MACHINE DIES AFTER 3-5 MIN WITHOUT CHARGING. | H11:SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE FOLLOWING WERE REVIEWED AS PART OF THIS INVESTIGATION: PATIENT SEVERITY, APPLICABLE PREVIOUS INVESTIGATION(S), SAMPLE (IF AVAILABLE), APPLICABLE FMEA DOCUMENTS, LABELING, AND APPLICABLE MANUFACTURE RECORDS. BASED ON A REVIEW OF THIS INFORMATION, THE FOLLOWING WAS CONCLUDED: THE DEVICE WAS RETURNED TO THE SERVICE FACILITY FOR EVALUATION. DURING EVALUATION, THE REPORTED ISSUE OF THE INDICATOR FOR THE BATTERY SHOWS FULL. THE MACHINE DIES AFTER 3-5 MIN WITHOUT CHARGING IS CONFIRMED. THE ROOT CAUSE OF THE REPORTED ISSUE WAS A FAULTY BATTERY. THE DEVICE WAS SERVICED, TESTED, AND RETURNED TO THE CUSTOMER. A HISTORY REVIEW OF SERIAL NUMBER (B)(6) SHOWED NO OTHER SIMILAR PRODUCT COMPLAINT(S) FROM THIS SERIAL NUMBER. THE DEVICE HAS NOT BEEN RECEIVED BY THE MANUFACTURER FOR EVALUATION. A HISTORY REVIEW OF SERIAL NUMBER (B)(4) SHOWED NO OTHER SIMILAR PRODUCT COMPLAINT(S) FROM THIS SERIAL NUMBER. | SYSTEM, IMAGING, PULSED ECHO, ULTRASONIC | BARD ACCESS SYSTEMS |  | N/A | IYO |  | 20200427 | Device was returned to manufacturer | HEALTH PROFESSIONAL | System, Imaging, Pulsed Echo, Ultrasonic
Malfunction |  | ['Improper or Incorrect Procedure or Method', 'Battery Problem', 'Charging Problem'] | ADDITIONAL INFORMATION WAS RECEIVED FROM THE PATIENT. IT WAS REPORTED THAT PATIENT WAS COMPLETELY SATISFIED WITH THE IMPLANT AND ITS EFFECTIVENESS. AS THEY BEGAN FEELING BETTER, THEY NEGLECTED TO KEEP IT FULLY CHARGED. THE PAIN RESUMED AND WHEN PATIENT TRIED CHARGING THE ELEMENT FOUND THAT IMPOSSIBLE. A COMPANY REP INSPECTED IT AND FOUND THE UNIT WAS DEAD TO THE POINT IT COULD NOT BE CHARGED ANY FURTHER. THE ALTERNATIVE WAS TO HAVE IT REPLACED. PATIENT STATED THEY ARE 91 YEARS OF AGE AND CHOOSE NOT TO HAVE THE REPLACEMENT. PATIENT PROVIDED THEIR WEIGHT INFORMATION. NO FURTHER COMPLICATIONS WERE REPORTED. | INFORMATION WAS RECEIVED FROM A HEALTHCARE PROVIDER (HCP) REGARDING A PATIENT WHO WAS IMPLANTED WITH A NEUROSTIMULATOR (INS). IT WAS REPORTED THAT THE IMPLANT WAS "DEAD". THE CALLER COULD NOT PROVIDE ANY ADDITIONAL INFORMATION. NO SYMPTOMS OR FURTHER COMPLICATIONS WERE REPORTED. | MEDTRONIC IS SUBMITTING THIS REPORT TO COMPLY WITH FDA REPORTING REGULATIONS UNDER 21 CFR PARTS 4 AND 803. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY MEDTRONIC, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. MEDTRONIC HAS MADE REASONABLE EFFORTS TO OBTAIN MORE COMPLETE INFORMATION AND HAS PROVIDED AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE AN ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY ¿DEFECTS¿ OR HAS ¿MALFUNCTIONED¿. THESE WORDS ARE INCLUDED IN THE FDA 3500A FORM AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REGULATORY REPORTING. MEDTRONIC OBJECTS TO THE USE OF THESE WORDS AND OTHERS LIKE THEM BECAUSE OF THE LACK OF DEFINITION AND THE CONNOTATIONS IMPLIED BY THESE TERMS. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. A GOOD FAITH EFFORT WILL BE MADE TO OBTAIN THE APPLICABLE INFORMATION RELEVANT TO THE REPORT. IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. | STIMULATOR, SPINAL-CORD, TOTALLY IMPLANTED FOR PAIN RELIEF | MEDTRONIC MED REL MEDTRONIC PUERTO RICO | 97714 |  | LGW | 20140928 |  | No | HEALTH PROFESSIONAL | Stimulator, Spinal-Cord, Totally Implanted For Pain Relief
Injury | 20200422 | ['Battery Problem'] | FOLLOWING THE BATTERY PERFORMANCE ALERT (BPA) ADVISORY, A BPA WAS RECEIVED BY THE CLINICIAN AND THE DEVICE WAS EXPLANTED. THE PATIENT WAS IN STABLE CONDITION. |  | THE DEVICE IS INCLUDED IN THE BATTERY PERFORMANCE ALERT ADVISORY ISSUED BY ABBOTT ON 28 AUGUST 2017. THE RESULTS, METHOD, AND CONCLUSION CODES ALONG WITH INVESTIGATION RESULTS WILL BE PROVIDED IN THE FINAL REPORT. PREMATURE BATTERY DEPLETION WAS CONFIRMED BY ANALYSIS. NO SOURCES OF HIGH CURRENT WERE NOTED. THE CAUSE OF THE PREMATURE BATTERY DEPLETION WAS CONSISTENT WITH LITHIUM (LI) CLUSTER FORMATION. FROM THESE ANALYSES, IN THE ABSENCE OF HIGH CURRENT DRAW, IT IS PROBABLE THAT THE PREMATURE BATTERY DEPLETION WAS CAUSED BY A LITHIUM CLUSTER INDUCED SHORT CIRCUIT. LI CLUSTERS ARE A KNOWN DEPLETION MECHANISM FOR THESE ADVISORY PRODUCTS THAT HAS BEEN INVESTIGATED AND ASSOCIATED WITH A FIELD ACTION IN OCTOBER 2016. | IMPLANTABLE CARDIOVERTER DEFIBRILLATOR | ABBOTT | CD3371-40QC | 4567144 | NIK | 20160430 | 20200520 | Device was returned to manufacturer | HEALTH PROFESSIONAL | Defibrillator, Automatic Implantable Cardioverter, With Cardiac Resynchronization (Crt-D)
Malfunction | 20200221 | ['Pumping Stopped', 'Battery Problem'] | ADDITIONAL INFORMATION WAS PROVIDED INDICATING THAT THERE WAS NO PATIENT INJURY. | IT WAS REPORTED THAT THE SMITHS MEDICAL CADD MS3 PUMP WAS ALARMING "BATTERY USAGE IS LOW." NO ADVERSE EFFECTS REPORTED. | OTHER, OTHER TEXT: ADDITIONAL INFORMATION: A2,A3,B3,B5,D10, AND H6. OTHER, OTHER TEXT: RETURNED DEVICE WAS RECEIVED IN GOOD PHYSICAL CONDITION. EVIDENCE OF REPORTED PROBLEM IN EVENT LOG WAS FOUND. DURING THE EVALUATION OF THE DEVICE, THE ISSUE COULD NOT BE REPEATED BY ANALYSTS EVEN THOUGH IT WAS FOUND IN THE EVENT HISTORY LOG. FURTHER INVESTIGATION FOUND THAT THE CUSTOMER'S BATTERY WAS LOW AND THAT MAY HAVE PROMPTED THE "ALARMING BATTERY USAGE IS LOW" MESSAGE. THIS IS NORMAL USE OF THE SYSTEM AND SYSTEM'S ALARMS. THE DEVICE HISTORY RECORD WAS REVIEWED AND SHOWED THAT THIS DEVICE MET ALL MANUFACTURING SPECIFICATION FOR PRODUCT RELEASED FOR DISTRIBUTION. NO ISSUES WERE IDENTIFIED THAT WOULD HAVE IMPACTED THIS EVENT. | PUMP, INFUSION | ST PAUL | 21-7411-51 |  | FRN |  | 20200409 | Device was returned to manufacturer | I | Pump, Infusion
Malfunction | 20200408 | ['Failure to Power Up', 'Battery Problem'] | THE CUSTOMER CONTACTED PHYSIO-CONTROL TO REPORT THAT THEIR DEVICE IS NOT FUNCTIONING. IN THIS STATE THE DEVICE MAY NOT BE ABLE TO POWER UP AND DELIVER DEFIBRILLATION THERAPY IF NEEDED. THERE WAS NO PATIENT USE REPORTED WITH THE EVENT. | THE CUSTOMER CONTACTED PHYSIO-CONTROL TO REPORT THAT THEIR DEVICE IS NOT FUNCTIONING. IN THIS STATE THE DEVICE MAY NOT BE ABLE TO POWER UP AND DELIVER DEFIBRILLATION THERAPY IF NEEDED. THERE WAS NO PATIENT USE REPORTED WITH THE EVENT. | PHYSIO-CONTROL EVALUATED THE CUSTOMERS DEVICE AND VERIFIED THAT THE DEVICE WOULD NOT POWER ON. PHYSIO REPLACED THE DEVICE'S BATTERY TO RESOLVE THE REPORTED ISSUE. PROPER DEVICE OPERATION WAS OBSERVED THROUGH FUNCTIONAL AND PERFORMANCE TESTING. THE DEVICE WAS ARCHIVED BY PHYSIO-CONTROL AND THE CUSTOMER WAS PROVIDED A REPLACEMENT DEVICE. CAUSE OF PREMATURE BATTERY DEPLETION COULD NOT BE DETERMINED. PHYSIO-CONTROL CONTINUES TO INVESTIGATE THE REPORTED FAILURE AND WILL SUBMIT A SUPPLEMENTAL REPORT ON THIS EVENT TO THE FDA AS PROVIDED BY 21 CFR 803.56. | AUTOMATED EXTERNAL DEFIBRILLATORS (NON-WEARABLE) | PHYSIO-CONTROL, INC. - 3015876 | CR2 |  | MKJ |  | 20200528 | Device was returned to manufacturer | HEALTH PROFESSIONAL | Automated External Defibrillators (Non-Wearable)
Malfunction | 20200427 | ['Battery Problem'] | INFORMATION WAS RECEIVED FROM THE CONSUMER REGARDING A PATIENT WHO WAS IMPLANTED WITH A NEUROSTIMULATOR FOR SPINAL CORD STIMULATION. IT WAS REPORTED THAT THE PATIENT SAW POR ON THE PP AND INSR. IT WAS REPORTED THE POR WAS AN INFORMATIONAL POR. PATIENT TRIED TO MAKE THE MESSAGE REAPPEAR AND CONFIRMED THE MESSAGE WAS NOW CLEARED. NO FURTHER COMPLICATIONS WERE REPORTED/ANTICIPATED. | ADDITIONAL INFORMATION WAS RECEIVED FROM THE PATIENT. THE PATIENT REPORTED THAT THE CIRCUMSTANCES THAT LED TO THE POR WAS THAT HE CHARGED UP THE STIMULATOR IN THE STOMACH. NO FURTHER COMPLICATIONS WERE REPORTED. | IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. | STIMULATOR, SPINAL-CORD, TOTALLY IMPLANTED FOR PAIN RELIEF | MEDTRONIC MED REL MEDTRONIC PUERTO RICO | 37714 |  | LGW | 20140914 |  | No | HEALTH PROFESSIONAL | Stimulator, Spinal-Cord, Totally Implanted For Pain Relief
Malfunction | 20200229 | ['Battery Problem'] | A US DISTRIBUTOR REPORTED THAT A PATIENT'S WAS RECEIVING BATTERY FAULTS. |  | DEVICE EVALUATION OF BATTERY SN (B)(4) HAS BEEN COMPLETED. THE REPORTED PROBLEM (WILL NOT POWER ON MONITOR) HAS BEEN CONFIRMED. AS RECEIVED, THE BATTERY WAS UNABLE TO CHARGE. UPON EVALUATION, THE BATTERY PACK TRI-CELLS WERE MIS-MATCHED. THE ROOT CAUSE OF THE MIS-MATCHED CELLS WAS UNABLE TO BE POSITIVELY IDENTIFIED. NO ADVERSE EVENT RESULTED FROM THE DEFECTIVE BATTERY. | WEARABLE CARDIOVERTER DEFIBRILLATOR | ZOLL MANUFACTURING CORPORATION | WCD 4000 |  | MVK |  | 20200330 | Device was returned to manufacturer | LAY USER/PATIENT | Wearable Automated External Defibrillator
Malfunction | 20200317 | ['Battery Problem'] | A US DISTRIBUTOR REPORTED THAT A PATIENT WAS EXPERIENCING BATTERY FAULTS. |  | DEVICE EVALUATION OF BATTERY PACK SN (B)(4) HAS BEEN COMPLETED. THE REPORTED PROBLEM (BATTERY FAULTS) WAS CONFIRMED. AS RECEIVED, THE BATTERY PACK WAS UNABLE TO POWER A TEST MONITOR AND CHARGE. UPON EVALUATION, THE NICKEL BUSBAR TABS AT THE P2, P3, AND P4 PADS WERE LOOSE. THE CAUSE OF THE NON-FUNCTIONAL BATTERY PACK IS THE LOOSE BUSBAR PADS. THE ROOT CAUSE OF THE LOOSE BUSBAR CANNOT BE POSITIVELY IDENTIFIED. NO ADVERSE EVENT RESULTED FROM THE DEFECTIVE BATTERY. | WEARABLE CARDIOVERTER DEFIBRILLATOR | ZOLL MANUFACTURING CORPORATION | WCD 4000 |  | MVK |  | 20200330 | Device was returned to manufacturer | LAY USER/PATIENT | Wearable Automated External Defibrillator
Malfunction | 20200322 | ['Battery Problem'] | A US DISTRIBUTOR REPORTED THAT A PATIENT WAS EXPERIENCING BATTERY FAULTS. |  | DEVICE EVALUATION OF BATTERY PACK SN (B)(4) HAS BEEN COMPLETED. THE REPORTED PROBLEM (BATTERY FAULTS) WAS CONFIRMED. AS RECEIVED, THE BATTERY PACK WAS UNABLE TO CHARGE. UPON EVALUATION, THE NICKEL BUSBAR TABS AT THE P1 AND P2 PADS WERE LOOSE. THE CAUSE OF THE NON-FUNCTIONAL BATTERY PACK IS THE LOOSE BUSBAR PADS. THE ROOT CAUSE OF THE LOOSE BUSBAR CANNOT BE POSITIVELY IDENTIFIED. NO ADVERSE EVENT RESULTED FROM THE DEFECTIVE BATTERY PACK. | WEARABLE CARDIOVERTER DEFIBRILLATOR | ZOLL MANUFACTURING CORPORATION | WCD 4000 |  | MVK |  | 20200331 | Device was returned to manufacturer | LAY USER/PATIENT | Wearable Automated External Defibrillator
Malfunction | 20200330 | ['Battery Problem'] | IT WAS REPORTED THAT THE DEVICE WAS FOUND TO BE UNABLE TO START UP CORRECTLY AND PRESENTED A CRITICAL ERROR DISPLAYING THE MESSAGE "4006 FAILURE". NO PATIENT HARMED AND NO INJURY REPORTED BECAUSE THIS WAS FOUND DURING SERVICE AND REPAIR. |  | THE INFORMATION RECEIVED BY THE MANUFACTURER HAS BEEN RE-EVALUATED FOR MDR REPORTING AND IT WAS DETERMINED THAT THIS CASE DOES NOT MEET THE THRESHOLD FOR REPORTING AND IS A NON-REPORTABLE EVENT. IF FURTHER DETAILS ARE PROVIDED CONFIRMING THE OCCURRENCE OF A REPORTABLE EVENT, OUR FILES WILL BE UPDATED ACCORDINGLY AND A FURTHER REPORT SUBMITTED OUTLINING BOTH THE EVENT DETAILS AND OUR INVESTIGATIONS PERFORMED. | NEGATIVE PRESSURE WOUND THERAPY POWERED SUCTION PUMP | SMITH & NEPHEW MEDICAL LTD. |  |  | OMP |  |  | No answer provided | 0 | Negative Pressure Wound Therapy Powered Suction Pump
Injury | 20200421 | ['Battery Problem'] | FOLLOWING THE BATTERY PERFORMANCE ALERT (BPA) ADVISORY, A BPA WAS RECEIVED BY THE CLINICIAN AND THE DEVICE WAS EXPLANTED. THE PATIENT WAS IN STABLE CONDITION. |  | THE DEVICE IS INCLUDED IN THE BATTERY PERFORMANCE ALERT ADVISORY ISSUED BY ABBOTT ON 28 AUGUST 2017. THE RESULTS OF THE INVESTIGATION ARE INCONCLUSIVE SINCE THE DEVICE WAS NOT RETURNED FOR ANALYSIS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED. | IMPLANTABLE CARDIOVERTER DEFIBRILLATOR | ABBOTT | CD3367-40QC | 4641577 | NIK | 20160531 |  | No | HEALTH PROFESSIONAL | Defibrillator, Automatic Implantable Cardioverter, With Cardiac Resynchronization (Crt-D)
Injury | 20200412 | ['Battery Problem'] | FOLLOWING THE BATTERY PERFORMANCE ALERT (BPA) ADVISORY, A BPA WAS RECEIVED BY THE CLINICIAN AND THE DEVICE WAS EXPLANTED. |  | THE RESULTS/METHOD AND CONCLUSION CODES ALONG WITH INVESTIGATION RESULTS WILL BE PROVIDED IN THE FINAL REPORT. PREMATURE BATTERY DEPLETION WAS CONFIRMED BY ANALYSIS. NO SOURCES OF HIGH CURRENT WERE NOTED. THE CAUSE OF THE PREMATURE BATTERY DEPLETION WAS CONSISTENT WITH LITHIUM (LI) CLUSTER FORMATION. FROM THESE ANALYSES, IN THE ABSENCE OF HIGH CURRENT DRAW, IT IS PROBABLE THAT THE PREMATURE BATTERY DEPLETION WAS CAUSED BY A LITHIUM CLUSTER INDUCED SHORT CIRCUIT. LI CLUSTERS ARE A KNOWN DEPLETION MECHANISM FOR THESE ADVISORY PRODUCTS THAT HAS BEEN INVESTIGATED AND ASSOCIATED WITH A FIELD ACTION IN OCTOBER 2016. | IMPLANTABLE CARDIOVERTER DEFIBRILLATOR | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | CD1231-40Q | 3302961 | NVZ | 20120731 | 20200428 | Device was returned to manufacturer | HEALTH PROFESSIONAL | Pulse Generator, Permanent, Implantable
Malfunction | 20200407 | ['Battery Problem'] | IT WAS REPORTED THAT THE PUMP BATTERY WAS DEPLETING QUICKLY. THE CUSTOMER DECLINED FOLLOW-UP FROM TANDEM TECHNICAL SUPPORT REGARDING THE ISSUE. THE CUSTOMER CONTINUED TO USE THE PUMP FOR INSULIN THERAPY. THERE WAS NO REPORTED ADVERSE IMPACT TO THE CUSTOMER¿S BLOOD GLUCOSE LEVEL. |  | NO PRODUCT WAS RETURNED FOR EVALUATION. SHOULD NEW RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. | INSULIN PUMP | TANDEM DIABETES CARE | 004628 |  | LZG |  |  | No | LAY USER/PATIENT | Pump, Infusion, Insulin
Malfunction | 20200407 | ['Battery Problem'] | IT WAS REPORTED THAT THE PUMP BATTERY WAS DEPLETING QUICKLY. THE CUSTOMER DECLINED FOLLOW-UP FROM TANDEM TECHNICAL SUPPORT REGARDING THE ISSUE. THE CUSTOMER CONTINUED TO USE THE PUMP FOR INSULIN THERAPY. THERE WAS NO REPORTED ADVERSE IMPACT TO THE CUSTOMER¿S BLOOD GLUCOSE LEVEL. |  | NO PRODUCT WAS RETURNED FOR EVALUATION. SHOULD NEW RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. | INSULIN PUMP | TANDEM DIABETES CARE | 004628 |  | LZG |  |  | No | LAY USER/PATIENT | Pump, Infusion, Insulin
Malfunction | 20200407 | ['Battery Problem'] | IT WAS REPORTED THAT THE PUMP BATTERY WAS DEPLETING QUICKLY. THE CUSTOMER DECLINED FOLLOW-UP FROM TANDEM TECHNICAL SUPPORT REGARDING THE ISSUE. THE CUSTOMER CONTINUED TO USE THE PUMP FOR INSULIN THERAPY. THERE WAS NO REPORTED ADVERSE IMPACT TO THE CUSTOMER¿S BLOOD GLUCOSE LEVEL. |  | NO PRODUCT WAS RETURNED FOR EVALUATION. SHOULD NEW RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. | INSULIN PUMP | TANDEM DIABETES CARE | 4628-003 |  | OYC |  |  | No | LAY USER/PATIENT | Pump, Infusion, Insulin, To Be Used With Invasive Glucose Sensor
Malfunction | 20200320 | ['Battery Problem'] | INFORMATION WAS RECEIVED INDICATING THAT SMITHS MEDICAL MEDFUSION 4000 PUMP HAS MALFUNCTIONED. THE PUMP WAS USED TO INFUSE EPINEPHRINE. THE PUMP WAS PLUGGED IN, GAVE AND IT DISPLAYED BIOMED ERROR WHEN UNPLUGGED. THE PUMP'S HISTORY SHOWED THAT THE PUMP CHANGED FROM AC TO DC WITH BATTERY CAPACITY AT 95.5 BUT IMMEDIATELY ALARMED SYSTEM FAILURE: BATTERY DEPLETED. THERE WERE NO ADVERSE EVENTS REPORTED. |  |  | PUMP, INFUSION | ST PAUL | 4000-0105-51 |  | FRN |  |  | No | HEALTH PROFESSIONAL | Pump, Infusion
Malfunction | 20200413 | ['Incorrect Measurement', 'Battery Problem'] | IT WAS REPORTED THAT THE PUMP BATTERY GAUGE WAS OBSERVED TO BE FLUCTUATING. IN ADDITION, IT WAS REPORTED THAT THE PUMP BATTERY DEPLETES FASTER THAN EXPECTED. THERE WAS NO REPORTED IMPACT TO THE CUSTOMER'S BLOOD GLUCOSE LEVEL. CUSTOMER DECLINED TO TROUBLESHOOT WITH TANDEM TECHNICAL SUPPORT. |  | NO PRODUCT WAS RETURNED FOR EVALUATION. SHOULD NEW RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. | INSULIN PUMP | TANDEM DIABETES CARE | 4628-003 |  | OYC |  |  | No | LAY USER/PATIENT | Pump, Infusion, Insulin, To Be Used With Invasive Glucose Sensor
Malfunction | 20200422 | ['Failure to Charge', 'Circuit Failure', 'Battery Problem'] | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR ROUTINE PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, "SERVICE REQUIRED" CODES WERE FOUND IN THE VENTILATOR'S DOWNLOADED ERROR LOG. THE DEVICE'S INTERNAL BATTERY AND POWER MANAGEMENT BOARD WERE REPLACED TO ADDRESS THE ISSUES. |  |  | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1040007 |  | CBK |  | 20200421 | Device was returned to manufacturer | HEALTH PROFESSIONAL | Ventilator, Continuous, Facility Use
Malfunction | 20200423 | ['Failure to Charge', 'Battery Problem'] | A VENTILATOR WAS RETURNED TO A THIRD PARTY SERVICE CENTER FOR EVALUATION. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE VENTILATOR AT THE THIRD PARTY SERVICE CENTER, A FAILURE OF THE DEVICE TO CHARGE IT'S INTERNAL BATTERY WAS OBSERVED. THE DEVICE'S INTERNAL BATTERY WAS REPLACED TO ADDRESS THE ISSUE. |  |  | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054260 |  | CBK |  |  | No | HEALTH PROFESSIONAL | Ventilator, Continuous, Facility Use
Malfunction | 20200421 | ['Battery Problem', 'Insufficient Information'] | INFORMATION WAS RECEIVED FROM A CONSUMER REGARDING A PATIENT WHO WAS IMPLANTED WITH AN IMPLANTABLE NEUROSTIMULATOR (INS). IT WAS REPORTED THAT THE PATIENT PROGRAMMER WAS SHOWING ERI. THE ERI WAS CLEARED AND THE LEFT SIDE WAS AT 2.57V. LAST WEEK, THE PATIENT FELT LIKE SHE WAS RUNNING OUT OF STEAM, LIKE HER INS DOES NOT FEEL LIKE ITS WORKING. THE PATIENT'S SPEECH WAS QUIET AND HARD TO UNDERSTAND. THE PATIENT HAD 3 BAD FALLS LAST WEEK. HER TOES WERE CRAMPING AND SHE WALKED ON HER TOES, WHICH SHE SAID WAS HER DYSTONIA DYSKINESIA SYMPTOM. SHE CALLED HER PHYSICIAN BUT HAD NOT HEARD A RESPONSE BACK. NO FURTHER COMPLICATIONS WERE REPORTED/ANTICIPATED. |  | IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. | STIMULATOR, ELECTRICAL, IMPLANTED, FOR PARKINSONIAN TREMOR | MEDTRONIC PUERTO RICO OPERATIONS CO. | 37602 |  | MHY | 20170214 |  | No | HEALTH PROFESSIONAL | Stimulator, Electrical, Implanted, For Parkinsonian Tremor
Malfunction | 20200424 | ['No Display/Image', 'Battery Problem'] | INFORMATION RECEIVED BY MEDTRONIC INDICATED THAT THE INSULIN PUMP TURNS OFF EVERY TIME THEY REPLACES THE BATTERY. NO HARM REQUIRING MEDICAL INTERVENTION WAS REPORTED. THE INSULIN PUMP WILL NOT BE RETURNED FOR ANALYSIS. |  | (B)(4). CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. THE INSULIN PUMP INVOLVED IN THIS EVENT IS THE 640G INSULIN INFUSION PUMP, WHICH IS NOT MARKETED IN THE UNITED STATES. HOWEVER, THE DEVICE IS SIMILAR TO THE PARADIGM REAL-TIME INSULIN INFUSION PUMP, WHICH IS MARKETED IN THE UNITED STATES. | PUMP, INFUSION, INSULIN, TO BE USED WITH INVASIVE GLUCOSE SENSOR | MEDTRONIC PUERTO RICO OPERATIONS CO. | MMT-554LWWL | A3554LWWLJ | OYC |  |  | No | LAY USER/PATIENT | Pump, Infusion, Insulin, To Be Used With Invasive Glucose Sensor
Malfunction | 20190612 | ['Battery Problem'] | THIS DEVICE WAS EXPLANTED DUE TO ERI. THE MANUFACTURERS ANALYSIS REVEALED A BATTERY ISSUE. NO ADVERSE PATIENT EVENTS WERE REPORTED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THIS FILE WILL BE UPDATED. |  | UPON RECEIPT, THE DEVICE INTERROGATION REVEALED THE EOS BATTERY STATUS, DETECTED ON AUGUST 26, 2019. THE DEVICE WAS IMPLANTED FOR 43 MONTHS AND 20 CHARGING CYCLES WERE RECORDED IN THE DEVICES MEMORY. THE ICD WAS SUBJECTED TO AN ELECTRICAL ANALYSIS. FIRST, THE EOS STATUS WAS REMOVED WITH A TECHNICAL PROGRAMMER AND THE ABILITY OF THE DEVICE TO DELIVER THERAPIES WAS VERIFIED. THE ANTI-BRADYCARDIA PACING PULSES PROVED TO BE NORMAL AND IN AMPLITUDE AND FREQUENCY AS PROGRAMMED. A FIBRILLATION SIGNAL WAS APPLIED AND THE DEVICE DETECTED THE FIBRILLATION SIGNAL AS SPECIFIED AND STARTED THE CHARGING OF A DEFIBRILLATION SHOCK. THE SHOCK WAS DELIVERED, BUT THE SPECIFIED ENERGY LEVEL WAS NOT REACHED. SUBSEQUENTLY, THE CURRENT CONSUMPTION OF THE ICD WAS ANALYZED AND FOUND TO BE NORMAL AND AS EXPECTED. HOWEVER, AN INCONSISTENCY BETWEEN CURRENT CONSUMPTION AND BATTERY VOLTAGE WAS NOTED. THEREFORE, THE ICD WAS OPENED AND THE INNER ASSEMBLY WAS INSPECTED. THE VISUAL INSPECTION OF THE INNER ASSEMBLY SHOWED NO ANOMALIES. THE MEASUREMENT OF THE BATTERY VOLTAGE SHOWED A DEPLETED BATTERY. IN A NEXT STEP, THE ELECTRONIC MODULE WAS ATTACHED TO AN EXTERNAL POWER SUPPLY TO CHECK THE FUNCTIONALITY OF THE ELECTRONIC MODULE. THERE WAS NO INDICATION OF A MALFUNCTION. ALL THERAPY FUNCTIONS WERE AVAILABLE AND WORKED AS EXPECTED. THE OVERALL CURRENT CONSUMPTION OF THE MODULE WAS DIRECTLY MEASURED AND PROVED TO BE NORMAL. THE BATTERY WAS SENT TO THE MANUFACTURER FOR FURTHER ANALYSIS. THE MANUFACTURING RECORDS OF THE BATTERY WERE INSPECTED, DOCUMENTING THAT THE BATTERY PARAMETERS WERE WITHIN SPECIFICATION DURING THE BATTERY MANUFACTURING. NO ANOMALIES ASSOCIATED WITH THIS BATTERY WERE DOCUMENTED DURING THE PRODUCTION PROCESS. THE VISUAL INSPECTION OF THE BATTERY DID NOT REVEAL ANY EXTERNAL SIGNS OF DAMAGE. THE VOLTAGE MEASUREMENT CONFIRMED THE BATTERY DEPLETION AND THE MICROCALORIMETRY ANALYSIS SHOWED ELEVATED HEAT DISSIPATION. AT A NEXT STEP, THE BATTERY WAS OPENED FOR DESTRUCTIVE ANALYSIS AND THE INNER ASSEMBLY WAS INSPECTED. THE ANALYSIS IDENTIFIED A DAMAGED INSULATION, WHICH LED TO AN INCREASED INTERNAL SELF-DEPLETION AND, AS A RESULT, TO THE CLINICAL OBSERVATION. IN CONCLUSION, THE PREMATURE BATTERY DEPLETION WAS CAUSED BY INSULATION DAMAGE WITHIN THE BATTERY. THE ELECTRONIC MODULE WORKED AS EXPECTED WITH AN EXTERNAL POWER SUPPLY. | ICD | BIOTRONIK SE & CO. KG | 390090 |  | LWS | 20160229 | 20190806 | Device was returned to manufacturer | HEALTH PROFESSIONAL | Implantable Cardioverter Defibrillator (Non-Crt)
Malfunction | 20200423 | ['Battery Problem'] | AFTER AN IMPLANTATION PERIOD OF APPROX. 73 MONTHS, A BATTERY ERROR MESSAGE WAS REPORTED DURING THE FOLLOW-UP FOR THE PLANNED EXCHANGE OF THE DEVICE. THE PACEMAKER WAS EXPLANTED. |  | THE DEVICE WAS NOT RETURNED FOR ANALYSIS. THE ANALYSIS IS THEREFORE BASED ON THE EXISTING PRODUCTION DOCUMENTS AND THE RETURNED DEVICE DATA. THE MANUFACTURING PROCESS OF THIS DEVICE WAS REVIEWED. THE PRODUCTION DOCUMENTS SHOWED NO ANOMALIES. ALL MANUFACTURING STEPS HAD BEEN CARRIED OUT CORRECTLY. A PERFORMED STANDARD FINAL ELECTRICAL TEST DOCUMENTED PROPER DEVICE BEHAVIOR. THE PROVIDED DEVICE DATA WERE THOROUGHLY ANALYZED. IN PARTICULAR, THE BATTERY VOLTAGE WAS CHECKED REGARDING TO THE CHARGE DRAWN FROM THE BATTERY BASED ON THE DEVICE DATA, AND IT TURNED OUT NOT TO BE AS EXPECTED. BASED ON THE AVAILABLE DEVICE DATA, IT WAS NOT POSSIBLE TO DETERMINE THE CAUSE OF THE CLINICAL OBSERVATION. AN EXAMINATION OF THE PACEMAKER IS THUS NECESSARY FOR A MORE DETAILED ANALYSIS. IT WAS REPORTED THAT THE IMPLANT WAS EXPLANTED BUT IS NOT AVAILABLE FOR ANALYSIS. IF ADDITIONAL RELEVANT INFORMATION SHOULD BECOME AVAILABLE, THE INCIDENT WILL BE UPDATED ACCORDINGLY. | PACEMAKER | BIOTRONIK SE & CO. KG | 371994 |  | NVZ | 20140430 |  | No | HEALTH PROFESSIONAL | Pulse Generator, Permanent, Implantable
Malfunction | 20200427 | ['Mechanical Problem', 'Obstruction of Flow', 'Battery Problem', 'Power Problem', 'Insufficient Information', 'Patient Device Interaction Problem'] | INFORMATION RECEIVED BY MEDTRONIC INDICATED THAT INSULIN PUMP HAD MOTOR ERROR ALARM AND AND UNEXPECTED ERROR CODE ALARM. CUSTOMER REPORTED THAT INSULIN PUMP RECEIVED UNEXPECTED NO DELIVERY ALARM, SCREEN WAS SCRATCHED AND INSULIN PUMP REJECTED NEW BATTERIES. NO HARM REQUIRING MEDICAL INTERVENTION WAS REPORTED. THE INSULIN PUMP WILL NOT BE RETURNED FOR ANALYSIS. |  | (B)(4). CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. | ARTIFICIAL PANCREAS DEVICE SYSTEM, THRESHOLD SUSPEND | MEDTRONIC PUERTO RICO OPERATIONS CO. | MMT-751NAS | A4751NASJ | OZO |  |  | No | LAY USER/PATIENT | Automated Insulin Dosing , Threshold Suspend
Malfunction | 20200306 | ['Battery Problem'] | IT WAS REPORTED THAT THE PUMP BATTERY INTERMITTENTLY DEPLETED QUICKLY WHEN BELOW A CERTAIN BATTERY PERCENTAGE. BLOOD GLUCOSE WAS NOT IMPACTED. THE CUSTOMER CONTINUED TO USE THE PUMP FOR INSULIN THERAPY. |  | NO PRODUCT WAS RETURNED FOR EVALUATION. SHOULD NEW RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. | CONTINUOUS GLUCOSE MONITOR | TANDEM DIABETES CARE | 1000096 |  | OYC |  |  | No | LAY USER/PATIENT | Pump, Infusion, Insulin, To Be Used With Invasive Glucose Sensor
Injury | 20200423 | ['Battery Problem'] | FOLLOWING THE BATTERY PERFORMANCE ALERT (BPA) ADVISORY, A BPA WAS RECEIVED BY THE CLINICIAN AND THE DEVICE WAS EXPLANTED. THE PATIENT WAS IN STABLE CONDITION. |  | THE DEVICE IS INCLUDED IN THE BATTERY PERFORMANCE ALERT ADVISORY ISSUED BY ABBOTT ON 28 AUGUST 2017. THE RESULTS, METHOD, AND CONCLUSION CODES ALONG WITH INVESTIGATION RESULTS WILL BE PROVIDED IN THE FINAL REPORT. PREMATURE BATTERY DEPLETION WAS CONFIRMED BY ANALYSIS. A DEVICE EVALUATION WAS PERFORMED, AND NO SOURCES OF HIGH CURRENT WERE NOTED. THE CAUSE OF THE PREMATURE BATTERY DEPLETION WAS CONSISTENT WITH LITHIUM (LI) CLUSTER FORMATION. FROM THESE ANALYSES, IN THE ABSENCE OF HIGH CURRENT DRAW, IT IS PROBABLE THAT THE PREMATURE BATTERY DEPLETION WAS CAUSED BY A LI CLUSTER INDUCED SHORT CIRCUIT. LI CLUSTERS ARE A KNOWN DEPLETION MECHANISM FOR THESE ADVISORY PRODUCTS THAT HAS BEEN INVESTIGATED AND ASSOCIATED WITH A FIELD ACTION IN OCTOBER 2016. | IMPLANTABLE CARDIOVERTER DEFIBRILLATOR | ABBOTT | CD3371-40Q | 4137015 | NIK | 20150630 | 20200528 | Device was returned to manufacturer | HEALTH PROFESSIONAL | Defibrillator, Automatic Implantable Cardioverter, With Cardiac Resynchronization (Crt-D)
Malfunction | 20200407 | ['Filling Problem', 'Failure to Sense', 'Battery Problem'] | IT WAS REPORTED THAT THE PUMP BATTERY WAS DEPLETING QUICKLY AND THAT THE PUMP WAS DELAYED IN RESPONDING TO PRESSES. ADDITIONALLY, IT WAS REPORTED THAT "PUMP FILLED TUBING SLOWER." THERE WAS NO REPORTED IMPACT TO THE CUSTOMER¿S BLOOD GLUCOSE LEVEL. CUSTOMER DECLINED TROUBLESHOOTING WITH TANDEM TECHNICAL SUPPORT AND DECLINED TO PROVIDE FURTHER INFORMATION. |  | NO PRODUCT WAS RETURNED FOR EVALUATION. SHOULD NEW RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. | INSULIN PUMP | TANDEM DIABETES CARE | 4628-003 |  | OYC |  |  | No | LAY USER/PATIENT | Pump, Infusion, Insulin, To Be Used With Invasive Glucose Sensor
Malfunction | 20200101 | ['Battery Problem'] | IT WAS REPORTED THAT THE PUMP BATTERY INTERMITTENTLY DEPLETED QUICKLY. THERE WAS NO REPORTED IMPACT TO THE CUSTOMER¿S BLOOD GLUCOSE LEVEL. THE CUSTOMER REVERTED TO AN ALTERNATE METHOD OF INSULIN THERAPY. |  | THE DEVICE IS EXPECTED TO BE RETURNED; HOWEVER, THE DEVICE HAS NOT YET BEEN RECEIVED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED IF THE DEVICE IS RECEIVED. | INSULIN PUMP | TANDEM DIABETES CARE | 1000096 |  | LZG |  | 20200527 | Device was returned to manufacturer | LAY USER/PATIENT | Pump, Infusion, Insulin
Malfunction | 20200306 | ['Battery Problem'] | IT WAS REPORTED THAT THE PUMP BATTERY INTERMITTENTLY DEPLETED QUICKLY. THE CUSTOMER¿S BLOOD GLUCOSE RANGED FROM 80-110 MG/DL. MULTIPLE ATTEMPTS WERE MADE TO FOLLOW UP WITH THE CUSTOMER REGARDING THE REPORTED ISSUE; HOWEVER, NO RESPONSE FROM THE CUSTOMER WAS RECEIVED. |  | NO PRODUCT WAS RETURNED FOR EVALUATION. SHOULD NEW RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. | INSULIN PUMP | TANDEM DIABETES CARE | 1000096 |  | LZG |  |  | No | LAY USER/PATIENT | Pump, Infusion, Insulin
Malfunction | 20200427 | ['Failure to Charge', 'Battery Problem'] | A VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR SERVICE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, AN ERROR CODE RELATED TO THE CHARGING OF THE INTERNAL BATTERY WAS OBSERVED. THE DEVICE'S INTERNAL BATTERY WAS REPLACED TO ADDRESS THE ISSUE. |  |  | VENTILATOR, CONTINUOUS, FACILITY USE | RESPIRONICS, INC | 1054260 |  | CBK |  | 20200305 | Device was returned to manufacturer | HEALTH PROFESSIONAL | Ventilator, Continuous, Facility Use
Malfunction | 20200406 | ['Battery Problem'] | IT WAS REPORTED THAT THE PUMP BATTERY INTERMITTENTLY DEPLETED QUICKLY. THE CUSTOMER'S BLOOD GLUCOSE WAS 80 MG/DL. THE CUSTOMER CONTINUED TO USE THE PUMP FOR INSULIN THERAPY. |  | THE DEVICE IS EXPECTED TO BE RETURNED; HOWEVER, THE DEVICE HAS NOT YET BEEN RECEIVED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED IF THE DEVICE IS RECEIVED. | AUTOMATED INSULIN DOSING, THRESHOLD SUSPEND | TANDEM DIABETES CARE | 1000096 |  | OZO |  | 20200608 | Device was returned to manufacturer | LAY USER/PATIENT | Automated Insulin Dosing , Threshold Suspend
Malfunction | 20200406 | ['Battery Problem'] | IT WAS REPORTED THAT THE PUMP BATTERY WAS DEPLETING QUICKLY. THERE WAS NO REPORTED IMPACT TO THE CUSTOMER¿S BLOOD GLUCOSE LEVEL. THE CUSTOMER DECLINED FOLLOW UP FROM TANDEM TECHNICAL SUPPORT. |  | NO PRODUCT WAS RETURNED FOR EVALUATION. SHOULD NEW RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. | INSULIN PUMP | TANDEM DIABETES CARE | 4628-003 |  | OYC |  |  | No | LAY USER/PATIENT | Pump, Infusion, Insulin, To Be Used With Invasive Glucose Sensor
Malfunction | 20200406 | ['Battery Problem'] | IT WAS REPORTED THAT THE PUMP BATTERY WAS DEPLETING QUICKLY. THERE WAS NO REPORTED IMPACT TO THE CUSTOMER¿S BLOOD GLUCOSE LEVEL. THE CUSTOMER DECLINED FOLLOW UP FROM TANDEM TECHNICAL SUPPORT. |  | NO PRODUCT WAS RETURNED FOR EVALUATION. SHOULD NEW RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. | INSULIN PUMP | TANDEM DIABETES CARE | 4628-003 |  | OYC |  |  | No | LAY USER/PATIENT | Pump, Infusion, Insulin, To Be Used With Invasive Glucose Sensor
Malfunction | 20200406 | ['Battery Problem'] | IT WAS REPORTED THAT THE PUMP BATTERY WAS DEPLETING QUICKLY. THERE WAS NO REPORTED IMPACT TO THE CUSTOMER¿S BLOOD GLUCOSE LEVEL. THE CUSTOMER DECLINED FOLLOW UP FROM TANDEM TECHNICAL SUPPORT. |  | NO PRODUCT WAS RETURNED FOR EVALUATION. SHOULD NEW RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. | CONTINUOUS GLUCOSE MONITOR | TANDEM DIABETES CARE | 1000096 |  | OYC |  |  | No | LAY USER/PATIENT | Pump, Infusion, Insulin, To Be Used With Invasive Glucose Sensor
Malfunction | 20200406 | ['Battery Problem'] | IT WAS REPORTED THAT THE PUMP BATTERY WAS DEPLETING QUICKLY. THE CUSTOMER¿S BLOOD GLUCOSE WAS 94 MG/DL. THE CUSTOMER REVERTED TO AN ALTERNATE METHOD OF INSULIN THERAPY. |  |  | AUTOMATED INSULIN DOSING, THRESHOLD SUSPEND | TANDEM DIABETES CARE | 1000096 |  | OZO |  | 20200415 | Device was returned to manufacturer | LAY USER/PATIENT | Automated Insulin Dosing , Threshold Suspend
Malfunction | 20200406 | ['Battery Problem'] | IT WAS REPORTED THAT THE PUMP BATTERY WAS DEPLETING QUICKLY. THERE WAS NO REPORTED IMPACT TO THE CUSTOMER¿S BLOOD GLUCOSE LEVEL. THE CUSTOMER DECLINED TO PROVIDE ADDITIONAL INFORMATION AND DECLINED FOLLOW UP FROM TANDEM TECHNICAL SUPPORT. |  | NO PRODUCT WAS RETURNED FOR EVALUATION. SHOULD NEW RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. | INSULIN PUMP | TANDEM DIABETES CARE | 4628-003 |  | OYC |  |  | No | LAY USER/PATIENT | Pump, Infusion, Insulin, To Be Used With Invasive Glucose Sensor
Malfunction | 20200401 | ['Battery Problem'] | IT WAS REPORTED THAT THE PUMP BATTERY WAS DEPLETING QUICKLY. THERE WAS NO REPORTED ADVERSE IMPACT TO THE CUSTOMER¿S BLOOD GLUCOSE. THE CUSTOMER REVERTED TO AN ALTERNATE METHOD OF INSULIN THERAPY. |  | THE DEVICE IS EXPECTED TO BE RETURNED; HOWEVER, THE DEVICE HAS NOT YET BEEN RECEIVED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED IF THE DEVICE IS RECEIVED AND EVALUATION IS PERFORMED. | PUMP, INFUSION, INSULIN, TO BE USED WITH INVASIVE GLUCOSE SENSOR | TANDEM DIABETES CARE | 1000096 |  | OYC |  | 20200520 | Device was returned to manufacturer | LAY USER/PATIENT | Pump, Infusion, Insulin, To Be Used With Invasive Glucose Sensor
Injury | 20200108 | ['Intermittent Continuity', 'Disconnection', 'Failure to Deliver Energy', 'High impedance', 'Therapy Delivered to Incorrect Body Area', 'Battery Problem', 'Therapeutic or Diagnostic Output Failure', 'Insufficient Information'] | ADDITIONAL INFORMATION WAS RECEIVED FROM A HEALTHCARE PROVIDER (HCP) VIA A MANUFACTURER REPRESENTATIVE REGARDING A PATIENT WHO WAS IMPLANTED WITH A NEUROSTIMULATOR (INS). IT WAS REPORTED THAT THE PATIENT HAD AN INFECTION OF THE SURGICAL SITE FROM THE REVISION SURGERY. THE ENTIRE SYSTEM WAS EXPLANTED ON (B)(6) AND ALL PRODUCTS WERE DISCARDED BY THE CUSTOMER. THE PRODUCT WILL NOT BE REPLACED IN THE FUTURE. THE ISSUE RESOLVED. NO FURTHER COMPLICATIONS WERE REPORTED OR ANTICIPATED. | ADDITIONAL INFORMATION WAS RECEIVED FROM THE REP. IT WAS REPORTED THAT PATIENT WENT HOME AFTER REVISION SURGERY AND BATTERY WAS DEPLETING WITHOUT IT BEING TURNED ON AND WAS DOWN TO 40%. PATIENT CHARGED TO 100% JUST BEFORE COMING INTO CLINIC AND IT IS ALREADY DOWN TO 90%. IMPEDANCES OUT OF RANGE AND HE IS GETTING OOR. WHILE ON THE CALL, CALLER WAS ABLE TO SWITCH ELECTRODES AND GET COVERAGE ONE SIDE. PATIENT ALSO NOTED THAT HIS BATTERY HAS BEEN DRAINING EVERY DAY. IT WAS ADVISED TO THE CALLER THAT THEY SHOULD REPROGRAM AROUND BEST PAIRS AND PATIENT SHOULD MONITOR GOING FORWARD. IF ISSUES DO NOT RESOLVE THEN, LAST STEP WOULD BE TO DISCUSS WITH HCP. NO FURTHER COMPLICATIONS WERE REPORTED. INFORMATION WAS RECEIVED FROM A PATIENT REGARDING THEIR IMPLANTABLE NEUROSTIMULATOR (INS). IT WAS REPORTED THAT EVER SINCE THE IMPLANT, HE HAS NOT HAD RELIABLE THERAPEUTIC BENEFIT. PATIENT HAS 15 ELECTRODES ON HIS LEADS AND ONLY 5 WORK. PATIENT IS NOT GETTING THE STIMULATION THAT HE WAS GETTING BEFORE. IT SEEMED LIKE IT WAS WORKING GOOD BUT THEN WOULDN'T COVER RIGHT SIDE OF HIS BACK. PATIENT WOULD GO FOR REPROGRAMMING AND GET IT A LITTLE BETTER BUT THEN THE LEGS WOULD QUIT WORKING. PATIENT WOULD GET A SHARP STIMULATION SENSATION IN HIS LEFT SIDE OF RIBS SO THEN HE WOULD GO AGAIN FOR REPROGRAMMING AND GOT RID OF THAT. THE REPS WERE NEVER REALLY ABLE TO GET THE STIMULATION SETTINGS TO WHERE HE NEEDS THEM AND HE WOULD LIKE TO MEET WITH SOMEONE AGAIN TO DISCUSS THIS. NOW HE IS HAVING A PROBLEM WHERE WHEN HE BREATHS IN HE LOSES STIMULATION SENSATION ON THE LEFT SIDE AND WHEN HE EXHALES THE SENSATION COMES BACK. SO HE THINKS HE HAS A BROKEN LEAD. PATIENT WAS REDIRECTED TO THEIR HCP. NO FURTHER COMPLICATIONS WERE REPORTED. ADDITIONAL INFORMATION WAS RECEIVED FROM THE CONSUMER VIA MANUFACTURER'S REPRESENTATIVE. IT WAS REPORTED THAT THE PATIENT EXPERIENCED A LOSS OF PARESTHESIA FROM TONIC SCS STIMULATION EVEN THOUGH THE IPG WAS FUNCTIONAL AND CHARGED. PATIENT WAS GETTING OOR ON PATIENT CONTROLLER. NO KNOWN CONTRIBUTING FACTORS. PATIENT WILL HAVE IMAGING AND OTHER STUDIES PERFORMED. PATIENT TO FOLLOW-UP WITH SURGEON TO INVESTIGATE POSSIBLE SURGICAL INTERVENTION IF IMAGING SHOWS DAMAGE TO LEADS OR EXTENSIONS. ISSUE NOT RESOLVED AT THIS TIME. INTERVENTION PLANNED. NO FURTHER COMPLICATIONS WERE REPORTED/ANTICIPATED. ADDITIONAL INFORMATION WAS RECEIVED FROM THE REP. IT WAS REPORTED THAT THE REP SAW THE PATIENT FOR HIS POST-OP APPOINTMENT ON (B)(6) 2020. ACCORDING TO THE NOTES, HIGH IMPEDANCES WERE OBSERVED IN THE OR AND IN THE RECOVERY ROOM AT THE TIME OF BATTERY REPLACEMENT. THE CAUSE WAS UNKNOWN. TO REP'S KNOWLEDGE, IT HAS NOT BEEN CONFIRMED IF THE LEAD WAS BROKEN. ON (B)(6) 2020, AT THE POST-OP APPOINTMENT, REP REPROGRAMMED AROUND THE BAD CONTACTS TO PROVIDE BETTER STIMULATION COVERAGE. IT WAS UNKNOWN WHAT ACTIONS/INTERVENTIONS TOOK PLACE, IF THE ISSUE WAS RESOLVED. REP BELIEVED THAT THE HCP WAS AWARE. NO FURTHER COMPLICATIONS WERE REPORTED. ADDITIONAL INFORMATION WAS RECEIVED. IT WAS REPORTED THAT THE PATIENT HAD A REVISION ON (B)(6) 2020 DUE TO HIGH IMPEDANCES. X-RAYS WERE TAKEN PRIOR TO THE SURGICAL INTERVENTION, AND THE IMAGING RESULTS INDICATED THAT ONE EXTENSION HAD COME OUT OF THE INS HEADER BLOCK. THE PHYSICIAN REPLACED THE EXTENSION WITH THE 60 CM EXTENSION AS THE SHORTER 40 CM EXTENSION SEEMED A BIT TIGHT. NO FURTHER COMPLICATIONS WERE REPORTED OR ANTICIPATED. ADDITIONAL INFORMATION WAS RECEIVED FROM THE REP. IT WAS REPORTED THAT AT THE PATIENT APPOINTMENT ON (B)(6), THE PATIENT STILL HAD 5 ELECTRODES WITH BAD IMPEDANCES. REP WAS ABLE TO PROGRAM AROUND THOSE TO GET THE PATIENT THE APPROPRIATE COVERAGE. THE PATIENT SAID THE COVERAGE WAS STILL NOT AS GOOD AS HE HAD BEFORE THE JANUARY BATTERY REPLACEMENT. IT WAS UNKNOWN WHY THE EXTENSION CAME OUT OF THE HEADER BLOCK. THE 40 CM EXTENSIONS WERE NOT SAVED OR RETURNED TO MANUFACTURER AS THE HCP DID NOT BELIEVE THEY CAUSED THE MALFUNCTION. HCP JUST THOUGHT THE 60 CM EXTENSIONS WOULD PROVIDE A LITTLE MORE SLACK. AFTER REPLACING THE EXTENSIONS, SINCE WE STILL HAVE BAD IMPEDANCES, HE IS CONVINCED THE PROBLEM MUST BE IN THE LEAD. HOWEVER, HE DID NOT WANT TO REPLACE THE LEAD AT THIS TIME. HE WANTS TO SEE IF THE PATIENT IS GETTING SATISFACTORY COVERAGE BEFORE DOING ANYTHING ELSE. NO FURTHER COMPLICATIONS WERE REPORTED. | CONTINUATION OF D11: PRODUCT ID 3708160 LOT# SERIAL# (B)(6) IMPLANTED: (B)(6) EXPLANTED: (B)(6) PRODUCT TYPE EXTENSION PRODUCT ID 3708160 LOT# SERIAL# (B)(6) IMPLANTED: (B)(6) EXPLANTED: (B)(6) PRODUCT TYPE EXTENSION PRODUCT ID 39565-30 LOT# SERIAL# V483023010 IMPLANTED: (B)(6) EXPLANTED: (B)(6)PRODUCT TYPE LEAD PRODUCT ID 3708140 LOT# SERIAL# NJB082914V IMPLANTED: (B)(6) EXPLANTED: 2020-05-08 PRODUCT TYPE EXTENSION PRODUCT ID 3708140 LOT# SERIAL# (B)(6) IMPLANTED: (B)(6) EXPLANTED: (B)(6) PRODUCT TYPE EXTENSION MEDTRONIC IS SUBMITTING THIS REPORT TO COMPLY WITH FDA REPORTING REGULATIONS UNDER 21 CFR PARTS 4 AND 803. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY MEDTRONIC, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. MEDTRONIC HAS MADE REASONABLE EFFORTS TO OBTAIN MORE COMPLETE INFORMATION AND HAS PROVIDED AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE AN ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY ¿DEFECTS¿ OR HAS ¿MALFUNCTIONED¿. THESE WORDS ARE INCLUDED IN THE FDA 3500A FORM AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REGULATORY REPORTING. MEDTRONIC OBJECTS TO THE USE OF THESE WORDS AND OTHERS LIKE THEM BECAUSE OF THE LACK OF DEFINITION AND THE CONNOTATIONS IMPLIED BY THESE TERMS. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. A GOOD FAITH EFFORT WILL BE MADE TO OBTAIN THE APPLICABLE INFORMATION RELEVANT TO THE REPORT. IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. B5: ADDED TEXT FOR ADDITIONAL INFORMATION, WAS LEFT OUT IN THE LAST SUBMITTED REGULATORY REPORT. G4: CORRECTED NOTIFIED DATE OF THE NEW INFORMATION TO 2020-MAY-08 FROM 2020-MAY-22. MEDTRONIC IS SUBMITTING THIS REPORT TO COMPLY WITH FDA REPORTING REGULATIONS UNDER 21 CFR PARTS 4 AND 803. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY MEDTRONIC, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. MEDTRONIC HAS MADE REASONABLE EFFORTS TO OBTAIN MORE COMPLETE INFORMATION AND HAS PROVIDED AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE AN ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY ¿DEFECTS¿ OR HAS ¿MALFUNCTIONED¿. THESE WORDS ARE INCLUDED IN THE FDA 3500A FORM AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REGULATORY REPORTING. MEDTRONIC OBJECTS TO THE USE OF THESE WORDS AND OTHERS LIKE THEM BECAUSE OF THE LACK OF DEFINITION AND THE CONNOTATIONS IMPLIED BY THESE TERMS. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. A GOOD FAITH EFFORT WILL BE MADE TO OBTAIN THE APPLICABLE INFORMATION RELEVANT TO THE REPORT. IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. CONCOMITANT MEDICAL PRODUCTS: PRODUCT ID: 39565-30, SERIAL#: (B)(4), IMPLANTED: (B)(6) 2010, PRODUCT TYPE: LEAD. OTHER RELEVANT DEVICE(S) ARE: PRODUCT ID: 39565-30, SERIAL/LOT #: (B)(4), UBD: 27-MAY-2014, UDI#: (B)(4). IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. MEDTRONIC IS SUBMITTING THIS REPORT TO COMPLY WITH FDA REPORTING REGULATIONS UNDER 21 CFR PARTS 4 AND 803. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY MEDTRONIC, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. MEDTRONIC HAS MADE REASONABLE EFFORTS TO OBTAIN MORE COMPLETE INFORMATION AND HAS PROVIDED AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE AN ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY ¿DEFECTS¿ OR HAS ¿MALFUNCTIONED¿. THESE WORDS ARE INCLUDED IN THE FDA 3500A FORM AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REGULATORY REPORTING. MEDTRONIC OBJECTS TO THE USE OF THESE WORDS AND OTHERS LIKE THEM BECAUSE OF THE LACK OF DEFINITION AND THE CONNOTATIONS IMPLIED BY THESE TERMS. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. A GOOD FAITH EFFORT WILL BE MADE TO OBTAIN THE APPLICABLE INFORMATION RELEVANT TO THE REPORT. IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. THE CORRECT AWARE DATE FOR THE LAST SUBMITTED ADDITIONAL INFORMATION IS 05/22/20, NOT 01/22/20. MEDTRONIC IS SUBMITTING THIS REPORT TO COMPLY WITH FDA REPORTING REGULATIONS UNDER 21 CFR PARTS 4 AND 803. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY MEDTRONIC, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. MEDTRONIC HAS MADE REASONABLE EFFORTS TO OBTAIN MORE COMPLETE INFORMATION AND HAS PROVIDED AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE AN ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY ¿DEFECTS¿ OR HAS ¿MALFUNCTIONED¿. THESE WORDS ARE INCLUDED IN THE FDA 3500A FORM AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REGULATORY REPORTING. MEDTRONIC OBJECTS TO THE USE OF THESE WORDS AND OTHERS LIKE THEM BECAUSE OF THE LACK OF DEFINITION AND THE CONNOTATIONS IMPLIED BY THESE TERMS. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. A GOOD FAITH EFFORT WILL BE MADE TO OBTAIN THE APPLICABLE INFORMATION RELEVANT TO THE REPORT. IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. PRODUCT ID: 39565-30, SERIAL# (B)(4), IMPLANTED: (B)(6) 2010, PRODUCT TYPE: LEAD; PRODUCT ID: 3708140, SERIAL# (B)(4), IMPLANTED: (B)(6) 2010, EXPLANTED: (B)(6) 2020, PRODUCT TYPE: EXTENSION; PRODUCT ID: 3708140, SERIAL# (B)(4), IMPLANTED: (B)(6) 2010, EXPLANTED: (B)(6) 2020, PRODUCT TYPE: EXTENSION. IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. | STIMULATOR, SPINAL-CORD, TOTALLY IMPLANTED FOR PAIN RELIEF | MEDTRONIC PUERTO RICO OPERATIONS CO. | 97715 |  | LGW | 20201128 |  | No | HEALTH PROFESSIONAL | Stimulator, Spinal-Cord, Totally Implanted For Pain Relief
Malfunction | 20200406 | ['Battery Problem'] | IT WAS REPORTED THAT THE PUMP BATTERY WAS DEPLETING QUICKLY. THE CUSTOMER'S BLOOD GLUCOSE WAS 107 MG/DL. MULTIPLE ATTEMPTS WERE MADE TO FOLLOW UP WITH THE CUSTOMER REGARDING THE REPORTED ISSUE; HOWEVER, NO RESPONSE FROM THE CUSTOMER WAS RECEIVED. |  | NO PRODUCT WAS RETURNED FOR EVALUATION. SHOULD NEW RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. | AUTOMATED INSULIN DOSING, THRESHOLD SUSPEND | TANDEM DIABETES CARE | 1000096 |  | OZO |  |  | No | LAY USER/PATIENT | Automated Insulin Dosing , Threshold Suspend
Malfunction | 20200406 | ['Battery Problem'] | IT WAS REPORTED THAT THE PUMP BATTERY WAS DEPLETING QUICKLY. THERE WAS NO REPORTED IMPACT TO THE CUSTOMER¿S BLOOD GLUCOSE LEVEL. IN A FOLLOW-UP EMAIL, THE CONTACT REPORTED THAT THE CUSTOMER'S ISSUE HAD BEEN RESOLVED. |  | NO PRODUCT WAS RETURNED FOR EVALUATION. SHOULD NEW RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. | AUTOMATED INSULIN DOSING, THRESHOLD SUSPEND | TANDEM DIABETES CARE | 1000096 |  | OZO |  |  | No | LAY USER/PATIENT | Automated Insulin Dosing , Threshold Suspend
Malfunction |  | ['Battery Problem'] | PER BIOMED, THE BATTERY WILL NOT HOLD THE PROMISED RUN TIME. ON (B)(6) 2020 - SAMPLE EVALUATION CONFIRMED BATTERY ABRUPT SHUTDOWN. |  | THE FOLLOWING WERE REVIEWED AS PART OF THIS INVESTIGATION: PATIENT SEVERITY, APPLICABLE PREVIOUS INVESTIGATION(S), SAMPLE (IF AVAILABLE), APPLICABLE FMEA DOCUMENTS, LABELING, AND APPLICABLE MANUFACTURE RECORDS. BASED ON A REVIEW OF THIS INFORMATION, THE FOLLOWING WAS CONCLUDED: THE DEVICE WAS RETURNED TO THE SERVICE FACILITY FOR EVALUATION. DURING EVALUATION, THE REPORTED ISSUE OF THE BATTERY WILL NOT HOLD THE PROMISED RUN TIME WAS CONFIRMED. THE SR8 WAS RECEIVED IN GOOD PHYSICAL CONDITION. THE SR8 WAS POWERED ON AT 96% BATTERY LIFE AND RAN FOR AROUND THIRTY MINUTES BEFORE SHUTTING DOWN WITH 80% BATTERY LIFE STILL ON THE SCANNER. THE ROOT CAUSE OF THE REPORTED ISSUE WAS IDENTIFIED AS AN INTERNAL BATTERY FAILURE. THE DEVICE WAS SERVICED, TESTED, AND RETURNED TO THE CUSTOMER. A HISTORY REVIEW OF SERIAL NUMBER (B)(4) SHOWED NO OTHER SIMILAR PRODUCT COMPLAINT(S) FROM THIS. | SYSTEM, IMAGING, PULSED ECHO, ULTRASONIC | BARD ACCESS SYSTEMS | 9770501 | N/A | IYO |  | 20200109 | Device was returned to manufacturer | HEALTH PROFESSIONAL | System, Imaging, Pulsed Echo, Ultrasonic
Malfunction |  | ['Battery Problem'] | PER BIOMED, BATTERY WILL NOT HOLD A CHARGE. ON (B)(6) 2020 - SAMPLE EVALUATION CONFIRMED BATTERY ABRUPT SHUTDOWN. |  | THE FOLLOWING WERE REVIEWED AS PART OF THIS INVESTIGATION: PATIENT SEVERITY, APPLICABLE PREVIOUS INVESTIGATION(S), SAMPLE (IF AVAILABLE), APPLICABLE FMEA DOCUMENTS, LABELING, AND APPLICABLE MANUFACTURE RECORDS. BASED ON A REVIEW OF THIS INFORMATION, THE FOLLOWING WAS CONCLUDED: THE DEVICE WAS RETURNED TO THE SERVICE FACILITY FOR EVALUATION. DURING EVALUATION, THE REPORTED ISSUE OF BATTERY WILL NOT HOLD A CHARGE WAS CONFIRMED. THE SITE RITE 8 WAS VISUALLY INSPECTED UPON RECEIPT AND WAS FOUND TO BE IN GOOD PHYSICAL CONDITION. THE ROOT CAUSE OF THE REPORTED ISSUE WAS IDENTIFIED AS AN INTERNAL LI-ION BATTERY FAILURE. THE DEVICE WAS SERVICED, TESTED, AND RETURNED TO THE CUSTOMER. A HISTORY REVIEW OF SERIAL NUMBER (B)(4) SHOWED NO OTHER SIMILAR PRODUCT COMPLAINT(S) FROM THIS SERIAL NUMBER. | SYSTEM, IMAGING, PULSED ECHO, ULTRASONIC | BARD ACCESS SYSTEMS | 9770600Y | N/A | IYO |  | 20200204 | Device was returned to manufacturer | HEALTH PROFESSIONAL | System, Imaging, Pulsed Echo, Ultrasonic
Malfunction |  | ['Battery Problem'] | PER TM - THE SCANNER WON'T HOLD A CHARGE. ON (B)(6) 2020 - SAMPLE EVALUATION CONFIRMED BATTERY ABRUPT SHUTDOWN. |  | THE FOLLOWING WERE REVIEWED AS PART OF THIS INVESTIGATION: PATIENT SEVERITY, APPLICABLE PREVIOUS INVESTIGATION(S), SAMPLE (IF AVAILABLE), APPLICABLE FMEA DOCUMENTS, LABELING, AND APPLICABLE MANUFACTURE RECORDS. BASED ON A REVIEW OF THIS INFORMATION, THE FOLLOWING WAS CONCLUDED: THE DEVICE WAS RETURNED TO THE SERVICE FACILITY FOR EVALUATION. DURING EVALUATION, THE REPORTED ISSUE OF THE SCANNER WON'T HOLD A CHARGE WAS CONFIRMED. THE SR8 WAS RECEIVED IN GOOD PHYSICAL CONDITION. THE SR8 WAS POWERED ON AT 94% BATTERY LIFE AND RAN ON BATTERY POWER FOR AROUND 15 MINUTES BEFORE SHUTTING DOWN WITH 83% BATTERY LIFE STILL ON THE SCANNER. THE ROOT CAUSE OF THE REPORTED ISSUE WAS IDENTIFIED AS AN INTERNAL BATTERY FAILURE. THE DEVICE WAS SERVICED, TESTED, AND RETURNED TO THE CUSTOMER. A HISTORY REVIEW OF SERIAL NUMBER (B)(4) SHOWED NO OTHER SIMILAR PRODUCT COMPLAINT(S) FROM THIS SERIAL NUMBER. | SYSTEM, IMAGING, PULSED ECHO, ULTRASONIC | BARD ACCESS SYSTEMS | 9770500 | N/A | IYO |  | 20200109 | Device was returned to manufacturer | HEALTH PROFESSIONAL | System, Imaging, Pulsed Echo, Ultrasonic
Malfunction | 20200428 | ['Battery Problem'] | IT WAS REPORTED THAT THE BATTERY WAS NOT RECOGNIZED BY THE CONTROLLER WHEN CONNECTED, AS THE CONTROLLER LIGHT EMITTING DIODES (L EDS) WERE STILL OFF. THE BATTERY WAS REPLACED. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. |  | ### A SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR DEVICE EVALUATION. PRODUCT EVENT SUMMARY: THE BATTERY WAS RETURNED FOR EVALUATION. VARIOUS ANALYSES WERE CONDUCTED AND REVIEWED IN ORDER TO EVALUATE THE PERFORMANCE OF THE DEVICE IN RELATION TO THE REPORTED EVENT. FAILURE ANALYSIS OF THE RETURNED DEVICE REVEALED THAT THE UNIT PASSED VISUAL INSPECTION. FUNCTIONAL TESTING REVEALED THAT THE BATTERY WAS UNABLE TO CHARGE OR PROVIDE POWER DUE TO AN ENABLED ZERO-VOLT CHARGE (ZVCHG) FIELD-EFFECT TRANSISTOR (FET). THE GAS GAUGE IS NOT PROGRAMMED TO USE THE ZVCHG FET, WHICH INDICATES THAT THE BATTERY UNINTENTIONALLY ENABLED THE FET. FURTHER INVESTIGATION USING A REPRESENTATIVE SAMPLE REVEALED THAT THE REPORTED EVENT WAS REPLICATED WHILE THE BATTERY WAS CONNECTED TO A BATTERY CHARGER AND EXPOSED TO ELECTROSTATIC DISCHARGE (ESD). THE ZVCHG FET DISABLES THE CHARGE AND DISCHARGE FET, RENDERING THE BATTERY INOPERABLE. THE INABILITY OF THE BATTERY TO CHARGE WILL RESULT IN A FLASHING RED LIGHT WHEN CONNECTED TO THE BATTERY CHARGER. AS A RESULT, THE REPORTED EVENT WAS CONFIRMED. THE MOST LIKELY ROOT CAUSE OF THE REPORTED EVENT CAN BE ATTRIBUTED TO AN ESD EVENT WHILE CONNECTED TO A BATTERY CHARGER. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. ### MEDTRONIC IS SUBMITTING THIS REPORT TO COMPLY WITH FDA REPORTING REGULATIONS UNDER 21 CFR PARTS 4 AND 803. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY MEDTRONIC, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. MEDTRONIC HAS MADE REASONABLE EFFORTS TO OBTAIN MORE COMPLETE INFORMATION AND HAS PROVIDED AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE AN ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY ¿DEFECTS¿ OR HAS ¿MALFUNCTIONED¿. THESE WORDS ARE INCLUDED IN THE FDA 3500A FORM AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REGULATORY REPORTING. MEDTRONIC OBJECTS TO THE USE OF THESE WORDS AND OTHERS LIKE THEM BECAUSE OF THE LACK OF DEFINITION AND THE CONNOTATIONS IMPLIED BY THESE TERMS. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. A GOOD FAITH EFFORT WILL BE MADE TO OBTAIN THE APPLICABLE INFORMATION RELEVANT TO THE REPORT. IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANS PLANTATION AND DESTINATION THERAPY IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY END-STAGE LEFT VENTRICULAR HEART FAILURE. THE SYSTEM IS DESIGNED FOR IN-HOSPITAL AND OUT-OF-HOSPITAL SETTINGS. INVESTIGATION OF THIS EVENT IS PENDING AND A SUPPLEMENTAL REPORT WILL BE SENT UPON ITS COMPLETION. IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. A SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR AN UPDATE TO THE PRODUCT EVENT SUMMARY AND ANNEX C, D AND G CODES. PRODUCT EVENT SUMMARY: THE ZVCHG FET DISABLES THE CHARGE AND DISCHARGE FETS, RENDERING THE BATTERY INOPERABLE. BASED ON AN INVESTIGATION CONDUCTED UNDER CAPA: PR00519785, THE MOST LIKELY ROOT CAUSE OF THE REPORTED EVENT HAS BEEN ATTRIBUTED TO THE BATTERY CONNECTOR INTERFACE DESIGN WHICH DOES NOT GUARANTEE THE DESIRED CONNECTION ANGLE IN SOME USE CONDITIONS. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC IS SUBMITTING THIS REPORT TO COMPLY WITH FDA REPORTING REGULATIONS UNDER 21 CFR PARTS 4 AND 803. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY MEDTRONIC, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. MEDTRONIC HAS MADE REASONABLE EFFORTS TO OBTAIN MORE COMPLETE INFORMATION AND HAS PROVIDED AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE AN ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY ¿DEFECTS¿ OR HAS ¿MALFUNCTIONED¿. THESE WORDS ARE INCLUDED IN THE FDA 3500A FORM AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REGULATORY REPORTING. MEDTRONIC OBJECTS TO THE USE OF THESE WORDS AND OTHERS LIKE THEM BECAUSE OF THE LACK OF DEFINITION AND THE CONNOTATIONS IMPLIED BY THESE TERMS. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT. | VENTRICULAR (ASSISST) BYPASS | HEARTWARE, INC. | 1650DE |  | DSQ | 20190930 | 20200515 | Device was returned to manufacturer | LAY USER/PATIENT | Ventricular (Assist) Bypass
Malfunction | 20180205 | ['Failure to Deliver Energy', 'High impedance', 'Improper or Incorrect Procedure or Method', 'Battery Problem', 'Electromagnetic Compatibility Problem'] | ADDITIONAL INFORMATION WAS RECEIVED FROM THE REP. IT WAS REPORTED THAT NO ACTIONS/ INTERVENTIONS HAVE BEEN TAKING YET. PATIENT GOT A REFERRAL BACK TO THEIR HCP TO PURSUE INTERVENTION HOWEVER NO CONFIRMED APPOINTMENT. | INFORMATION WAS RECEIVED FROM A MANUFACTURING REPRESENTATIVE (REP) REGARDING A PATIENT. THE REP REPORTED THAT AN IMPEDANCE CHECK WAS RAN AND ALL ELECTRODES WERE OUT OF RANGE. THEY STATED THAT IMPEDANCES WERE GREATER THAN 10,000, 20,000, AND 40,000 OHMS. THE REP NOTED THAT THE PATIENT HAD LUMBAR SPINAL SURGERY ABOUT 2 YEARS AGO AND INDICATED THAT THE PATIENT THINKS THAT THE ISSUE OCCURRED AROUND THAT TIME. THEY STATED THAT THE PATIENT STOPPED USING THE DEVICE BECAUSE THEY HAD SURGERY AND COULD NOT FEEL STIMULATION ANYMORE. THE REP ADDED THAT THE PATIENT¿S IMPLANTABLE NEUROSTIMULATOR (INS) APPEARED TO BE IN OVERDISCHARGE. THEY STATED THAT A JUMP START WAS PERFORMED, AND IMPEDANCES WERE STILL OUT OF RANGE. THE REP NOTED THAT NO INTERVENTION HAS OCCURRED AT THIS TIME, AND THAT THE PATIENT WILL NEED TO BE REFERRED BACK TO THEIR HEALTHCARE PROVIDER. NO FURTHER COMPLICATIONS WERE REPORTED OR ANTICIPATED. | IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. | STIMULATOR, SPINAL-CORD, TOTALLY IMPLANTED FOR PAIN RELIEF | MEDTRONIC PUERTO RICO OPERATIONS CO. | 97713 |  | LGW | 20150228 |  | No | HEALTH PROFESSIONAL | Stimulator, Spinal-Cord, Totally Implanted For Pain Relief
Malfunction | 20200425 | ['Mechanical Problem', 'Battery Problem', 'Unintended Movement'] | INFORMATION RECEIVED BY MEDTRONIC INDICATED THAT THE INSULIN PUMP HAD MULTIPLE PUMP ERROR ALARMS. THE CUSTOMER STATED INSULIN PUMP DID NOT ACCEPT NEW BATTERY. NO HARM REQUIRING MEDICAL INTERVENTION WAS REPORTED. THE INSULIN PUMP WILL BE RETURNED FOR ANALYSIS. |  | (B)(4). CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. THE DEVICE WILL BE RETURNED FOR ANALYSIS AND FURTHER INFORMATION WILL FOLLOW ONCE THE ANALYSIS HAS BEEN COMPLETED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. THE INSULIN PUMP INVOLVED IN THIS EVENT IS THE 640G INSULIN INFUSION PUMP, WHICH IS NOT MARKETED IN THE UNITED STATES. HOWEVER, THE DEVICE IS SIMILAR TO THE PARADIGM REAL-TIME INSULIN INFUSION PUMP, WHICH IS MARKETED IN THE UNITED STATES. DEVICE PASSED THE DISPLACEMENT TEST, SELF TEST, ACTIVE CURRENT MEASUREMENT, REWIND TEST, PRIME/SEATING TEST, BASIC OCCLUSION TEST, FORCE SENSOR TEST AND OCCLUSION TEST. NO UNEXPECTED INSULIN FLOW BLOCKED ALARM OR NO DELIVERY ALARM NOTED DURING TESTING. NO UNEXPECTED PUMP ERROR 43 OR PUMP ERROR 130 NOTED. MOTOR WAS TESTED OUTSIDE DEVICE AND PASSED. HOWEVER, DEVICE RECEIVED WITH UNEXPECTED BATTERY POWER LOSS SHOWN IN POWER GRAPH AND HAD HIGH SLEEP CURRENT, PROBLEM ISOLATED CORRODED ELECTRONIC ASSEMBLIES. | PUMP, INFUSION, INSULIN, TO BE USED WITH INVASIVE GLUCOSE SENSOR | MEDTRONIC PUERTO RICO OPERATIONS CO. | MMT-1711K | HG3Y2LP | OYC |  | 20200610 | Device was returned to manufacturer | LAY USER/PATIENT | Pump, Infusion, Insulin, To Be Used With Invasive Glucose Sensor
Malfunction | 20200427 | ['Device Alarm System', 'Battery Problem', 'Power Problem'] | INFORMATION RECEIVED BY MEDTRONIC INDICATED THAT INSULIN PUMP HAD BEEP ANOMALY BUT NO ALERT ON SCREEN. CUSTOMER REPORTED THAT INSULIN PUMP REJECTED NEW BATTERIES AND BATTERIES LAST LESS THAN ONE WEEK. NO HARM REQUIRING MEDICAL INTERVENTION WAS REPORTED. THE INSULIN PUMP WILL NOT BE RETURNED FOR ANALYSIS. |  | (B)(4). CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. | ARTIFICIAL PANCREAS DEVICE SYSTEM, THRESHOLD SUSPEND | MEDTRONIC PUERTO RICO OPERATIONS CO. | MMT-751NAL | A4751NALJ | OZO |  |  | No | LAY USER/PATIENT | Automated Insulin Dosing , Threshold Suspend
Injury | 20200414 | ['Nonstandard Device', 'Device Displays Incorrect Message', 'Battery Problem'] | FOLLOWING THE BATTERY PERFORMANCE ALERT (BPA) ADVISORY, A BPA WAS RECEIVED BY THE CLINICIAN AND THE DEVICE WAS EXPLANTED. THE PATIENT WAS IN STABLE CONDITION. |  | PREMATURE BATTERY DEPLETION WAS CONFIRMED BY ANALYSIS. A DEVICE EVALUATION WAS PERFORMED, AND NO SOURCES OF HIGH CURRENT WERE NOTED. THE CAUSE OF THE PREMATURE BATTERY DEPLETION WAS CONSISTENT WITH LI CLUSTER FORMATION. FROM THESE ANALYSES, IN THE ABSENCE OF HIGH CURRENT DRAW, IT IS PROBABLE THAT THE PREMATURE BATTERY DEPLETION WAS CAUSED BY A LITHIUM CLUSTER INDUCED SHORT CIRCUIT. LI CLUSTERS ARE A KNOWN DEPLETION MECHANISM FOR THESE ADVISORY PRODUCTS THAT HAS BEEN INVESTIGATED AND ASSOCIATED WITH A FIELD ACTION IN OCTOBER 2016. THE DEVICE IS INCLUDED IN THE BATTERY PERFORMANCE ALERT ADVISORY ISSUED BY ABBOTT ON 28 AUGUST 2017. THE RESULTS, METHOD, AND CONCLUSION CODES ALONG WITH INVESTIGATION RESULTS WILL BE PROVIDED IN THE FINAL REPORT. | IMPLANTABLE CARDIOVERTER DEFIBRILLATOR | ABBOTT | CD1233-40 | 4037483 | LWS | 20141031 | 20200505 | Device was returned to manufacturer | HEALTH PROFESSIONAL | Implantable Cardioverter Defibrillator (Non-Crt)
Malfunction | 20200408 | ['Battery Problem'] | IT WAS REPORTED THAT THE PUMP BATTERY WAS DEPLETING QUICKLY. THERE WAS NO REPORTED IMPACT TO THE CUSTOMER¿S BLOOD GLUCOSE LEVEL. CUSTOMER DECLINED TROUBLESHOOTING WITH TANDEM TECHNICAL SUPPORT AND DECLINED TO PROVIDE FURTHER INFORMATION. |  | NO PRODUCT WAS RETURNED FOR EVALUATION. SHOULD NEW RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. | INSULIN PUMP | TANDEM DIABETES CARE | 4628-003 |  | OYC |  |  | No | LAY USER/PATIENT | Pump, Infusion, Insulin, To Be Used With Invasive Glucose Sensor
Malfunction |  | ['Battery Problem'] | THE CUSTOMER IDENTIFIED THAT THE BUILT-IN ACCUMULATOR BATTERY IS OUT OF ORDER. THE FIELD SERVICE ENGINEER (FSE) CONFIRMED THE FAILURE. THE FSE REMOVED THE CABLE FROM BATTERY AND THE ISSUE WAS RESOLVED. THE UNIT HAS RETURNED TO SERVICE MODE. THERE WAS NO PATIENT INVOLVEMENT. |  | G4: 14MAY2020. B4: 15MAY2020. H11: H6: PATIENT INVOLVEMENT:: THE DEVICE WAS IN CLINICAL USE AT THE TIME THE REPORTED ISSUE WAS DISCOVERED; HOWEVER, THERE WAS NO REPORTED PATIENT OR USER HARM. SUBMISSION OF A REPORT DOES NOT CONSTITUTE AN ADMISSION THAT MEDICAL PERSONNEL, USER FACILITY, IMPORTER, DISTRIBUTOR, MANUFACTURER, OR PRODUCT CAUSED OR CONTRIBUTED TO THE EVENT. DATE OF EVENT: (B)(6) 2020. DATE OF REPORT: 29APR2020. | VENTILATOR, CONTINUOUS, MINIMAL VENTILATORY SUPPORT,FACILITY USE | RESPIRONICS CALIFORNIA, INC | V60 |  | MNT |  |  | Yes | I | Ventilator, Continuous, Minimal Ventilatory Support, Facility Use
Malfunction | 20200426 | ['Battery Problem', 'Power Problem'] | INFORMATION RECEIVED BY MEDTRONIC INDICATED THAT INSULIN PUMP RECEIVED REPLACE BATTERY AND POWER LOSS ALARM. CUSTOMER WAS CHANGING BATTERY AND RECEIVED POWER LOSS ALARM. NO HARM REQUIRING MEDICAL INTERVENTION WAS REPORTED. THE INSULIN PUMP WILL NOT BE RETURNED FOR ANALYSIS. |  | (B)(4). CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. THE INSULIN PUMP INVOLVED IN THIS EVENT IS THE 640G INSULIN INFUSION PUMP, WHICH IS NOT MARKETED IN THE UNITED STATES. HOWEVER, THE DEVICE IS SIMILAR TO THE PARADIGM REAL-TIME INSULIN INFUSION PUMP, WHICH IS MARKETED IN THE UNITED STATES. | PUMP, INFUSION, INSULIN, TO BE USED WITH INVASIVE GLUCOSE SENSOR | MEDTRONIC PUERTO RICO OPERATIONS CO. | MMT-1711K | HG10T5Y | OYC |  |  | No | LAY USER/PATIENT | Pump, Infusion, Insulin, To Be Used With Invasive Glucose Sensor
Injury | 20200415 | ['Electrical /Electronic Property Problem', 'Battery Problem', 'Power Problem'] | IT WAS REPORTED THAT THE PATIENT WAS HOSPITALIZED. THE CONTROLLER WAS DEFECTED. THE CONTROLLER EXHIBITED SEVERAL UNEXPECTED CONTROLLER POWER LOSSES, A CONTROLLER FAULT ALARM, AND POWER SWITCHING. THE BATTERY EXHIBITED COMMUNICATION ERROR WITH SEVERAL ASSOCIATED POWER DISCONNECTS ALARMS. THE BATTERY REMAINS IN USE AND THE CONTROLLER WAS EXCHANGED. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. | IT WAS REPORTED THAT THE CONTROLLER WAS DEFECTED. THE CONTROLLER EXHIBITED SEVERAL UNEXPECTED CONTROLLER POWER LOSSES, A CONTROLLER FAULT ALARM, POWER SWITCHING AND A VENTRICULAR ASSIST DEVICE (VAD) DISCONNECT ALARM. THE BATTERY EXHIBITED COMMUNICATION ERROR WITH SEVERAL ASSOCIATED POWER DISCONNECTS ALARMS. THE BATTERY REMAINS IN USE AND THE CONTROLLER WAS EXCHANGED. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. | AS A RESULT OF REVIEW OF THE COMPLAINT FILE, THIS REPORT CONTAINS UPDATES TO THE EVALUATION RESULT AND CONCLUSION CODES AND THE PRODUCT EVENT SUMMARY TO MORE ACCURATELY REFLECT WHAT IS IN THE COMPLAINT FILE. PRODUCT EVENT SUMMARY: THE CONTROLLER WAS RETURNED FOR EVALUATION. ONE BATTERY WAS NOT RETURNED FOR EVALUATION. A REVIEW OF THE MANUFACTURING DOCUMENTATION CONFIRMED THAT THE ASSOCIATED BATTERY MET ALL REQUIREMENTS FOR RELEASE. VARIOUS ANALYSES WERE CONDUCTED AND REVIEWED IN ORDER TO EVALUATE THE PERFORMANCE OF THE RETURNED DEVICE IN RELATION TO THE REPORTED EVENT. A VISUAL INSPECTION OF CONTROLLER UNDER 10X MAGNIFICATION REVEALED HAIRLINE CRACKS AROUND POWER PORT ONE (1). AN INTERNAL INSPECTION DID NOT REVEAL EVIDENCE OF FLUID INGRESS. THE OBSERVED HAIRLINE CRACK IS NOT RELATED TO THE REPORTED EVENT. BASED ON AN INVESTIGATION CONDUCTED UNDER AN INTERNAL INVESTIGATION, THE ROOT CAUSE OF THE HAIRLINE CRACK WAS DETERMINED TO BE DUE TO CHEMICAL ADDITIVES APPLIED TO THE POWER PORT GASKETS DURING THE MANUFACTURING PROCESS. THE CHEMICAL ADDITIVES CONTRIBUTED TO ENVIRONMENTAL STRESS CRACKING. EVEN THOUGH THIS INVESTIGATION IS CLOSED, (B)(6) FALLS WITHIN THE BOUNDS OF THIS INVESTIGATION. ADDITIONALLY, VISUAL INSPECTION REVEALED CONTAMINATION ON BOTH THE POWER PORTS ONE AND TWO (1 <(>&<)> 2), LIKELY DUE TO HANDLING OF THE DEVICE. FUNCTIONAL TESTING OF THE CONTROLLER REVEALED THAT "NO POWER" ALARM HAS SOUNDED FOR ONLY TWO MINUTES WHEN BOTH POWER SOURCES WERE DISCONNECTED. EVALUATION OF THE INTERNAL NICKEL-METAL HYDRIDE (NIMH) BATTERY REVEALED THAT ONE OF THE CELL'S VOLTAGE LEVEL WAS BELOW WHAT WAS EXPECTED. ADDITIONALLY, INTERNAL INSPECTION REVEALED THAT THE INTERNAL BATTERY CELLS WERE SWOLLEN. AFTER REPLACING THE BATTERY, THE CONTROLLER PERFORMED AS INTENDED. LOG FILE ANALYSIS REVEALED THAT THE CONTROLLER IN USE DURING THE REPORTED EVENT, THE CONTROLLER, CONTAINS A FEATURE THAT RECORDS WHETHER A POWER SOURCE EXPERIENCED A COMMUNICATION ERROR OR A DISCONNECTION WITHIN EACH 15-MINUTE INTERVAL. ANALYSIS OF DATA LOG FILE REVEALED MULTIPLE PREMATURE POWER SWITCHING EVENTS DUE TO MOMENTARY DISCONNECTIONS INVOLVING THE BATTERY WITHIN THE ANALYZED PERIOD. ADDITIONALLY, LOG FILES ALSO REVEALED THREE (3) CONTROLLER POWER UP EVENTS ON (B)(6) 2020 AT 15:25:56 AND ON (B)(6) 2020 AT 13:46:51 AND 14:46:04. THE CONTROLLER WAS WITHOUT POWER 11 SECONDS, 16 SECONDS AND 24 SECONDS RESPECTIVELY. ANALYSIS OF ALARM LOG FILES REVEALED ONE (1) CONTROLLER FAULT ALARM ON (B)(6) 2020 AT 14:30:23, DUE TO INTERNAL BATTERY FAILURE AND SEVERAL POWER DISCONNECT ALARMS INVOLVING THE BATTERY WERE LOGGED SINCE (B)(6) 2020, DUE TO A CELL UNDER VOLTAGE. DURING THE POWER DISCONNECT ALARM A SAFETY ALERT WORD (SAW) VALUE WAS RECORDED INDICATING THAT A CELL PAIR DEPLETED BELOW A VOLTAGE THRESHOLD. ADDITIONALLY, LOG FILES REVEALED THAT THE CONTROLLER WAS IN USE FOR MORE THAN TWO (2) YEARS. AS A RESULT, THE REPORTED UNEXPECTED LOSS OF POWER, PREMATURE POWER SWITCHING, CONTROLLER FAULT ALARM, AND POWER DISCONNECT ALARM EVENTS WERE CONFIRMED; HOWEVER, THE REPORTED COMMUNICATION ERROR EVENT COULD NOT BE CONFIRMED. THERE IS NO EVIDENCE OF POWER SOURCE LUBRICATION PROCEDURE PERFORMED ON THE ASSOCIATED BATTERIES. A POSSIBLE ROOT CAUSE OF THE LOSS OF POWER CAN BE ATTRIBUTED TO A DISCONNECTION OF BOTH POWER SOURCES AND/OR TO AN INTERMITTENT DISCONNECTION ON ONE OR BOTH POWER SOURCE. THE MOST LIKELY ROOT CAUSE OF THE REPORTED PREMATURE POWER SWITCHING EVENT CAN BE ATTRIBUTED TO MOMENTARY DISCONNECTIONS BETWEEN THE CONTROLLER AND BATTERIES. AN INTERNAL INVESTIGATION EVALUATED MOMENTARY DISCONNECTIONS. THE MOST LIKELY ROOT CAUSE OF THE REPORTED CONTROLLER FAULT ALARM EVENT CAN BE ATTRIBUTED TO A FAULTY INTERNAL NIMH BATTERY. AN INTERNAL INVESTIGATION WAS INITIATED TO INVESTIGATE THIS ISSUE. BASED ON THE AVAILABLE INFORMATION, A POSSIBLE ROOT CAUSE OF THE REPORTED POWER DISCONNECT ALARM EVENT COULD BE ATTRIBUTED, BUT NOT LIMITED, TO A BATTERY MALFUNCTION. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC IS SUBMITTING THIS REPORT TO COMPLY WITH FDA REPORTING REGULATIONS UNDER 21 CFR PARTS 4 AND 803. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY MEDTRONIC, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. MEDTRONIC HAS MADE REASONABLE EFFORTS TO OBTAIN MORE COMPLETE INFORMATION AND HAS PROVIDED AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE AN ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY ¿DEFECTS¿ OR HAS ¿MALFUNCTIONED¿. THESE WORDS ARE INCLUDED IN THE FDA 3500A FORM AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REGULATORY REPORTING. MEDTRONIC OBJECTS TO THE USE OF THESE WORDS AND OTHERS LIKE THEM BECAUSE OF THE LACK OF DEFINITION AND THE CONNOTATIONS IMPLIED BY THESE TERMS. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. A GOOD FAITH EFFORT WILL BE MADE TO OBTAIN THE APPLICABLE INFORMATION RELEVANT TO THE REPORT. IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. A SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR DEVICE EVALUATION AND INVESTIGATION COMPLETION. PRODUCT EVENT SUMMARY: THE CONTROLLER WAS RETURNED FOR EVALUATION. THE BATTERY WAS NOT RETURNED FOR EVALUATION. A REVIEW OF THE MANUFACTURING DOCUMENTATION CONFIRMED THAT THE ASSOCIATED BATTERY MET ALL REQUIREMENTS FOR RELEASE. VARIOUS ANALYSES WERE CONDUCTED AND REVIEWED IN ORDER TO EVALUATE THE PERFORMANCE OF THE RETURNED DEVICE IN RELATION TO THE REPORTED EVENT. A VISUAL INSPECTION OF CONTROLLER UNDER 10X MAGNIFICATION REVEALED HAIRLINE CRACKS AROUND POWER PORT ONE. AN INTERNAL INSPECTION DID NOT REVEAL EVIDENCE OF FLUID INGRESS. THE OBSERVED HAIRLINE CRACK IS NOT RELATED TO THE REPORTED EVENT. BASED ON AN INVESTIGATION CONDUCTED UNDER CAPA PR00381374, THE ROOT CAUSE OF THE HAIRLINE CRACK WAS DETERMINED TO BE DUE TO CHEMICAL ADDITIVES APPLIED TO THE POWER PORT GASKETS DURING THE MANUFACTURING PROCESS. THE CHEMICAL ADDITIVES CONTRIBUTED TO ENVIRONMENTAL STRESS CRACKING. EVEN THOUGH THIS CAPA IS CLOSED, THE CONTROLLER FALLS WITHIN THE BOUNDS OF THIS CAPA. ADDITIONALLY, VISUAL INSPECTION REVEALED CONTAMINATION ON BOTH THE POWER PORTS ONE AND TWO, LIKELY DUE TO HANDLING OF THE DEVICE. FUNCTIONAL TESTING OF THE CONTROLLER REVEALED THAT "NO POWER" ALARM HAS SOUNDED FOR ONLY TWO MINUTES WHEN BOTH POWER SOURCES WERE DISCONNECTED. EVALUATION OF THE INTERNAL NICKEL-METAL HYDRIDE (NIMH) BATTERY REVEALED THAT ONE OF THE CELL'S VOLTAGE LEVEL WAS BELOW WHAT WAS EXPECTED. ADDITIONALLY, INTERNAL INSPECTION REVEALED THAT THE INTERNAL BATTERY CELLS WERE SWOLLEN. AFTER REPLACING THE BATTERY, THE CONTROLLER PERFORMED AS INTENDED. LOG FILE ANALYSIS REVEALED THAT THE CONTROLLER IN USE DURING THE REPORTED EVENT, CONTAINS A FEATURE THAT RECORDS WHETHER A POWER SOURCE EXPERIENCED A COMMUNICATION ERROR OR A DISCONNECTION WITHIN EACH 15-MINUTE INTERVAL. ANALYSIS OF DATA LOG FILE REVEALED MULTIPLE PREMATURE POWER SWITCHING EVENTS DUE TO MOMENTARY DISCONNECTIONS INVOLVING THE BATTERY WITHIN THE ANALYZED PERIOD. ADDITIONALLY, LOG FILES ALSO REVEALED THREE CONTROLLER POWER UP EVENTS ON (B)(6) 2020 AT 15:25:56 AND ON (B)(6) 2020 AT 13:46:51 AND 14:46:04. THE CONTROLLER WAS WITHOUT POWER 11 SECONDS, 16 SECONDS AND 24 SECONDS RESPECTIVELY. ANALYSIS OF ALARM LOG FILES REVEALED ONE CONTROLLER FAULT ALARM ON (B)(6) 2020 AT 14:30:23, DUE TO INTERNAL BATTERY FAILURE AND SEVERAL POWER DISCONNECT ALARMS INVOLVING BATTERY WERE LOGGED SINCE (B)(6) 2020, DUE TO A CELL UNDER VOLTAGE. DURING THE POWER DISCONNECT ALARM A SAFETY ALERT WORD (SAW) VALUE WAS RECORDED INDICATING THAT A CELL PAIR DEPLETED BELOW A VOLTAGE THRESHOLD. AS A RESULT, THE REPORTED UNEXPECTED LOSS OF POWER, PREMATURE POWER SWITCHING, CONTROLLER FAULT ALARM, AND POWER DISCONNECT ALARM EVENTS WERE CONFIRMED; HOWEVER, THE REPORTED COMMUNICATION ERROR EVENT COULD NOT BE CONFIRMED. THERE IS NO EVIDENCE OF POWER SOURCE LUBRICATION PROCEDURE PERFORMED ON THE ASSOCIATED BATTERIES. A POSSIBLE ROOT CAUSE OF THE LOSS OF POWER CAN BE ATTRIBUTED TO A DISCONNECTION OF BOTH POWER SOURCES AND/OR TO AN INTERMITTENT DISCONNECTION ON ONE OR BOTH POWER SOURCE. THE MOST LIKELY ROOT CAUSE OF THE REPORTED PREMATURE POWER SWITCHING EVENT CAN BE ATTRIBUTED TO MOMENTARY DISCONNECTIONS BETWEEN THE CONTROLLER AND BATTERIES. CAPA PR00389403 INVESTIGATED MOMENTARY DISCONNECTIONS. THE MOST LIKELY ROOT CAUSE OF THE REPORTED CONTROLLER FAULT ALARM EVENT CAN BE ATTRIBUTED TO A FAULTY INTERNAL NIMH BATTERY. BASED ON THE AVAILABLE INFORMATION, A POSSIBLE ROOT CAUSE OF THE REPORTED POWER DISCONNECT ALARM EVENT COULD BE ATTRIBUTED, BUT NOT LIMITED, TO A BATTERY MALFUNCTION. ADDITIONAL PRODUCTS: D4: SERIAL #: (B)(6), H3: YES. H6: FDA METHOD CODE(S): 3331, 4112, 4114. H6: FDA RESULTS CODE(S): 131, 3213. H6: FDA CONCLUSION CODE(S): 12, 4307. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC IS SUBMITTING THIS REPORT TO COMPLY WITH FDA REPORTING REGULATIONS UNDER 21 CFR PARTS 4 AND 803. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY MEDTRONIC, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. MEDTRONIC HAS MADE REASONABLE EFFORTS TO OBTAIN MORE COMPLETE INFORMATION AND HAS PROVIDED AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE AN ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY ¿DEFECTS¿ OR HAS ¿MALFUNCTIONED¿. THESE WORDS ARE INCLUDED IN THE FDA 3500A FORM AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REGULATORY REPORTING. MEDTRONIC OBJECTS TO THE USE OF THESE WORDS AND OTHERS LIKE THEM BECAUSE OF THE LACK OF DEFINITION AND THE CONNOTATIONS IMPLIED BY THESE TERMS. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT. A SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR A CORRECTION. CORRECTED FIELDS: H1 TYPE OF REPORT: UPDATED TO SERIOUS INJURY. B5 EVENT DESCRIPTION WAS UPDATED ADDITIONAL CODES: IMF CODE WAS UPDATED TO F08 B1 ADV EVENT/PRODUCT PROBLEM UPDATED THIS REGULATORY REPORT IS BEING SUBMITTED AS PART OF A RETROSPECTIVE REVIEW AND REMEDIATION PER D00595827 DUE TO AN FDA AUDIT OBSERVATION. MEDTRONIC IS SUBMITTING THIS REPORT TO COMPLY WITH FDA REPORTING REGULATIONS UNDER 21 CFR PARTS 4 AND 803. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY MEDTRONIC, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. MEDTRONIC HAS MADE REASONABLE EFFORTS TO OBTAIN MORE COMPLETE INFORMATION AND HAS PROVIDED AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE AN ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY ¿DEFECTS¿ OR HAS ¿MALFUNCTIONED¿. THESE WORDS ARE INCLUDED IN THE FDA 3500A FORM AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REGULATORY REPORTING. MEDTRONIC OBJECTS TO THE USE OF THESE WORDS AND OTHERS LIKE THEM BECAUSE OF THE LACK OF DEFINITION AND THE CONNOTATIONS IMPLIED BY THESE TERMS. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. A GOOD FAITH EFFORT WILL BE MADE TO OBTAIN THE APPLICABLE INFORMATION RELEVANT TO THE REPORT. IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. THIS DEVICE IS USED FOR TREATMENT NOT DIAGNOSIS. THE VENTRICULAR ASSIST SYSTEM IS INDICATED FOR USE AS A BRIDGE TO CARDIAC TRANS PLANTATION AND DESTINATION THERAPY IN PATIENTS WHO ARE AT RISK OF DEATH FROM REFRACTORY END-STAGE LEFT VENTRICULAR HEART FAILURE. THE SYSTEM IS DESIGNED FOR IN-HOSPITAL AND OUT-OF-HOSPITAL SETTINGS. ADDITIONAL PRODUCTS: HEARTWARE VENTRICULAR ASSIST SYSTEM ¿ BATTERY, MODEL #: 1650DE / CATALOG #: 1650DE / EXPIRATION DATE: 28-FEB-2017 / SERIAL #: (B)(4), UDI #: (B)(4), DEVICE AVAILABLE FOR EVALUATION: NO, MFG DATE: 29-FEB-2016, LABELED FOR SINGLE USE: NO, (B)(4). INVESTIGATION OF THIS EVENT IS PENDING AND A SUPPLEMENTAL REPORT WILL BE SENT UPON ITS COMPLETION. IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. | VENTRICULAR (ASSISST) BYPASS | HEARTWARE, INC. | 1420 |  | DSQ | 20180331 | 20200508 | Device was returned to manufacturer | LAY USER/PATIENT | Ventricular (Assist) Bypass
Malfunction | 20200427 | ['Display or Visual Feedback Problem', 'Obstruction of Flow', 'Battery Problem'] | INFORMATION RECEIVED BY MEDTRONIC INDICATED THAT THE INSULIN PUMP SCREEN WAS FADING. INSULIN PUMP HAD UNEXPLAINED NO DELIVERY ALARMS, CHANGING BATTERIES MORE FREQUENTLY. NO HARM REQUIRING MEDICAL INTERVENTION WAS REPORTED. THE INSULIN PUMP WILL NOT BE RETURNED FOR ANALYSIS. |  | (B)(4). CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. | PUMP, INFUSION, INSULIN, TO BE USED WITH INVASIVE GLUCOSE SENSOR | MEDTRONIC PUERTO RICO OPERATIONS CO. | MMT-723LNAS | A6723LNASJ | OYC |  |  | No | LAY USER/PATIENT | Pump, Infusion, Insulin, To Be Used With Invasive Glucose Sensor
Injury | 20200416 | ['Battery Problem'] | FOLLOWING THE BATTERY PERFORMANCE ALERT (BPA) ADVISORY, A BPA WAS RECEIVED BY THE CLINICIAN AND THE DEVICE WAS EXPLANTED. THE PATIENT WAS IN STABLE CONDITION. |  | PRE-MATURE BATTERY DEPLETION WAS CONFIRMED BY ANALYSIS. A DEVICE EVALUATION WAS PERFORMED, AND NO SOURCES OF HIGH CURRENT WERE NOTED. THE CAUSE OF THE PREMATURE BATTERY DEPLETION WAS CONSISTENT WITH LI CLUSTER FORMATION. FROM THESE ANALYSES, IN THE ABSENCE OF HIGH CURRENT DRAW, IT IS PROBABLE THAT THE PREMATURE BATTERY DEPLETION WAS CAUSED BY A LITHIUM CLUSTER INDUCED SHORT CIRCUIT. LI CLUSTERS ARE A KNOWN DEPLETION MECHANISM FOR THESE ADVISORY PRODUCTS THAT HAS BEEN INVESTIGATED AND ASSOCIATED WITH A FIELD ACTION IN OCTOBER 2016. THE DEVICE IS INCLUDED IN THE BATTERY PERFORMANCE ALERT ADVISORY ISSUED BY ABBOTT ON 28 AUGUST 2017. THE RESULTS, METHOD, AND CONCLUSION CODES ALONG WITH INVESTIGATION RESULTS WILL BE PROVIDED IN THE FINAL REPORT. | IMPLANTABLE CARDIOVERTER DEFIBRILLATOR | ABBOTT | CD1233-40Q | 3499646 | LWS | 20130331 | 20200520 | Device was returned to manufacturer | HEALTH PROFESSIONAL | Implantable Cardioverter Defibrillator (Non-Crt)
Malfunction | 20200310 | ['No Display/Image', 'Battery Problem'] | INFORMATION RECEIVED BY MEDTRONIC INDICATED THAT INSULIN PUMP HAD BLANK SCREEN AND HAD SHORT BATTERY. NO HARM REQUIRING MEDICAL INTERVENTION WAS REPORTED. THE INSULIN PUMP WAS RETURNED FOR ANALYSIS. |  | THE INSULIN PUMP PASSED THE DISPLACEMENT TEST AND SELF TEST. HOWEVER, THE PUMP FAILED THE SLEEP CURRENT MEASUREMENT AND ACTIVE CURRENT MEASUREMENT DUE HIGH CURRENT, PROBLEM ISOLATED TO CONNECTOR. (B)(4). | ARTIFICIAL PANCREAS DEVICE SYSTEM, SINGLE HORMONAL CONTROL | MEDTRONIC PUERTO RICO OPERATIONS CO. | MMT-1780KPK | HG409MT | OZP |  | 20200318 | Device was returned to manufacturer | LAY USER/PATIENT | Automated Insulin Dosing Device System, Single Hormonal Control
Malfunction | 20200427 | ['Device Difficult to Program or Calibrate', 'Obstruction of Flow', 'Battery Problem'] | INFORMATION RECEIVED BY MEDTRONIC INDICATED THAT BUTTONS WERE HARD TO PUSH. THE CUSTOMER STATED THAT THE INSULIN PUMP HAD UNEXPLAINED NO DELIVERY ALERTS, DIMINISHED BATTERY LIFE. NO HARM REQUIRING MEDICAL INTERVENTION WAS REPORTED. THE INSULIN PUMP WILL NOT BE RETURNED FOR ANALYSIS. |  | (B)(4). CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. DEVICE RECEIVED WITH INTERMITTENT BUTTON RESPONSE DUE TO FLATTENED ACT BUTTON DOME SWITCH. NO BUTTON ERROR ALARM DURING TESTING. NO UNLOCKED LCD KEYPAD CONNECTOR. DEVICE PASSED SELF TEST. DEVICE RECEIVED WITH ALL OPERATING CURRENTS WITHIN SPECIFICATION AND PASSED FUNCTIONAL TESTING INCLUDING THE REWIND, BASIC OCCLUSION TEST, OCCLUSION TEST, PRIME/A33 TEST, EXCESSIVE NO DELIVERY TEST AND DISPLACEMENT TEST. NO EXCESSIVE NO DELIVERY OR OCCLUSION ALARM NOTED. NO UNEXPECTED LOW BATTERY ALARM ANOMALIES NOTED. (B)(4). | ARTIFICIAL PANCREAS DEVICE SYSTEM, THRESHOLD SUSPEND | MEDTRONIC PUERTO RICO OPERATIONS CO. | MMT-751NAB | A4751NABJ | OZO |  | 20200608 | Device was returned to manufacturer | LAY USER/PATIENT | Automated Insulin Dosing , Threshold Suspend
Injury | 20200409 | ['Battery Problem'] | FOLLOWING THE BATTERY PERFORMANCE ALERT (BPA) ADVISORY, A BPA WAS RECEIVED BY THE CLINICIAN AND THE DEVICE WAS EXPLANTED. PATIENT CONDITION WAS STABLE AFTER THE PROCEDURE. |  | THE RESULTS/METHOD AND CONCLUSION CODES ALONG WITH INVESTIGATION RESULTS WILL BE PROVIDED IN THE FINAL REPORT. PREMATURE BATTERY DEPLETION WAS CONFIRMED BY ANALYSIS. NO SOURCES OF HIGH CURRENT WERE NOTED. THE CAUSE OF THE PREMATURE BATTERY DEPLETION WAS CONSISTENT WITH LI CLUSTER FORMATION. FROM THESE ANALYSES, IN THE ABSENCE OF HIGH CURRENT DRAW, IT IS PROBABLE THAT THE PREMATURE BATTERY DEPLETION WAS CAUSED BY A LITHIUM CLUSTER INDUCED SHORT CIRCUIT. LI CLUSTERS ARE A KNOWN DEPLETION MECHANISM FOR THESE ADVISORY PRODUCTS THAT HAS BEEN INVESTIGATED AND ASSOCIATED WITH A FIELD ACTION IN OCTOBER 2016. | IMPLANTABLE CARDIOVERTER DEFIBRILLATOR | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | CD3365-40Q | 4597679 | NIK | 20160430 | 20200421 | Device was returned to manufacturer | HEALTH PROFESSIONAL | Defibrillator, Automatic Implantable Cardioverter, With Cardiac Resynchronization (Crt-D)
Malfunction | 20200427 | ['Mechanical Problem', 'Battery Problem', 'Power Problem'] | INFORMATION RECEIVED BY MEDTRONIC INDICATED THAT THE INSULIN PUMP HAD MOTOR ERROR ALARMS AND HAD REWIND ERRORS. THE CUSTOMER STATED THAT THE INSULIN PUMP HAD REJECTED NEW BATTERIES AND THE BATTERIES WERE LASTING 2-3 WEEKS ONLY. NO HARM REQUIRING MEDICAL INTERVENTION WAS REPORTED. TROUBLESHOOTING WAS DECLINED. THE INSULIN PUMP WILL NOT BE RETURNED FOR ANALYSIS. |  | (B)(4). CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. | PUMP, INFUSION, INSULIN, TO BE USED WITH INVASIVE GLUCOSE SENSOR | MEDTRONIC PUERTO RICO OPERATIONS CO. | MMT-723LNAS | A7723LNASJ | OYC |  |  | No | LAY USER/PATIENT | Pump, Infusion, Insulin, To Be Used With Invasive Glucose Sensor
Malfunction |  | ['Battery Problem'] | THE CUSTOMER REPORTED BATTERY FAULT. THERE WAS NO PATIENT INVOLVEMENT. |  | DATE OF EVENT: (B)(6) 2020. DATE OF REPORT: 29APR2020. G4: 10AUG2020 B4: (B)(6) 2020 THE MANUFACTURER'S FIELD SERVICE ENGINEER (FSE) CONFIRMED THE REPORTED PROBLEM. THE DEVICE WAS ALSO OBSERVED TO POWER ON AUTOMATICALLY. THE FSE REPLACED THE INTERNAL BATTERY TO ADDRESS AND CORRECT THIS REPORTED EVENT. THE USER INTERFACE PRINTED CIRCUIT BOARD ASSEMBLY WAS REPLACED TO ADDRESS THE POWER-ON ISSUE. SUBMISSION OF A REPORT DOES NOT CONSTITUTE AN ADMISSION THAT MEDICAL PERSONNEL, USER FACILITY, IMPORTER, DISTRIBUTOR, MANUFACTURER, OR PRODUCT CAUSED OR CONTRIBUTED TO THE EVENT. | VENTILATOR, CONTINUOUS, MINIMAL VENTILATORY SUPPORT,FACILITY USE | RESPIRONICS CALIFORNIA, INC | V60 |  | MNT |  |  | Yes | HEALTH PROFESSIONAL | Ventilator, Continuous, Minimal Ventilatory Support, Facility Use
Malfunction |  | ['Battery Problem'] | PER BIOMED: BATTERY DRAINS FAST. |  | THE FOLLOWING WERE REVIEWED AS PART OF THIS INVESTIGATION: PATIENT SEVERITY, APPLICABLE PREVIOUS INVESTIGATION(S), SAMPLE (IF AVAILABLE), APPLICABLE FMEA DOCUMENTS, LABELING, AND APPLICABLE MANUFACTURE RECORDS. BASED ON A REVIEW OF THIS INFORMATION, THE FOLLOWING WAS CONCLUDED: THE DEVICE WAS RETURNED TO THE SERVICE FACILITY FOR EVALUATION. DURING EVALUATION, THE REPORTED ISSUE OF REDUCED BATTERY RUN TIME WAS CONFIRMED. THE SCANNER SHUTS DOWN PREMATURELY WHEN AC IS UNPLUGGED. THE ROOT CAUSE OF THE REPORTED ISSUE IS AN INTERNAL LI-ION BATTERY FAILURE. THE DEVICE WAS SERVICED, TESTED, AND RETURNED TO THE CUSTOMER. A HISTORY REVIEW OF SERIAL NUMBER (B)(4) SHOWED NO OTHER SIMILAR PRODUCT COMPLAINT(S) FROM THIS SERIAL NUMBER. | SYSTEM, IMAGING, PULSED ECHO, ULTRASONIC | BARD ACCESS SYSTEMS | 9770600 | N/A | IYO |  | 20191018 | Device was returned to manufacturer | HEALTH PROFESSIONAL | System, Imaging, Pulsed Echo, Ultrasonic
Malfunction | 20200328 | ['Device Alarm System', 'Battery Problem'] | INFORMATION WAS RECEIVED INDICATING THAT SMITHS MEDICAL MEDFUSION 4000 PUMP HAS MALFUNCTIONED. THE PUMP DISPLAYED A SYSTEM FAILURE, A PRIMARY AUDIBLE ALARM BGND TEST ERROR AND A BIOMED PASSCODE REQUESTED MESSAGE. THE PUMP'S HISTORY SHOWED AN INTERMITTENT VOLUME OVER TIME INFUSION OF 3 ML SYRINGE GAVE AN EMPTY ALARM, THEN SYSTEM FAILURE PRIMARY AUDIBLE ALARM BGND, THEN SYSTEM FAILURE ACU WATCHDOG FAILURE. THERE WERE NO ADVERSE EVENTS REPORTED. |  | OTHER TEXT: ADDITIONAL INFORMATION: H6 DEVICE EVALUATION: THE PUMP WAS NOT RETURNED TO THE MANUFACTURER BUT WAS INVESTIGATED AND REPAIRED AT THE CUSTOMER FACILITY BY A FIELD SERVICE TECHNICIAN. A VISUAL INSPECTION OF THE PUMP FOUND THAT THE FACTORY SEAL IS MISSING. THE J9 CONNECTOR WAS IN GOOD CONDITION. THE PUMP WAS TESTED AFTER REASSEMBLY OF THE RIBBON CABLE CONNECTOR AND J9 CONNECTOR OF THE MAIN PRINTED CIRCUIT BOARD ASSEMBLY (PCBA) AND FOUND TO BE IN GOOD CONDITION; GLUE WAS INSTALLED. DURING FUNCTIONAL TESTING, NO ALARMS WERE PRESENT. THE REPORTED FAILURE WAS NOT CONFIRMED. THE ROOT CAUSE COULD NOT BE ESTABLISHED. A DEVICE HISTORY RECORD (DHR) REVIEW WAS CONDUCTED WHICH INDICATED ALL INSPECTIONS WERE COMPLETED AND NO ISSUES WERE NOTED DURING MANUFACTURING. THIS REMEDIATION MDR WAS GENERATED UNDER PROTOCOL (B)(4), AS A RESULT OF WARNING LETTER CMS# (B)(4). OTHER, OTHER TEXT: ADDITIONAL INFORMATION: H6 DEVICE EVALUATION: THE PUMP WAS NOT RETURNED TO THE MANUFACTURER BUT WAS INVESTIGATED AND REPAIRED AT THE CUSTOMER FACILITY BY A FIELD SERVICE TECHNICIAN. A VISUAL INSPECTION OF THE PUMP FOUND THAT THE FACTORY SEAL IS MISSING. THE J9 CONNECTOR WAS IN GOOD CONDITION. THE PUMP WAS TESTED AFTER REASSEMBLY OF THE RIBBON CABLE CONNECTOR AND J9 CONNECTOR OF THE MAIN PRINTED CIRCUIT BOARD ASSEMBLY (PCBA) AND FOUND TO BE IN GOOD CONDITION; GLUE WAS INSTALLED. DURING FUNCTIONAL TESTING, NO ALARMS WERE PRESENT. THE REPORTED FAILURE WAS NOT CONFIRMED. THE ROOT CAUSE COULD NOT BE ESTABLISHED. A DEVICE HISTORY RECORD (DHR) REVIEW WAS CONDUCTED WHICH INDICATED ALL INSPECTIONS WERE COMPLETED AND NO ISSUES WERE NOTED DURING MANUFACTURING. THIS REMEDIATION MDR WAS GENERATED UNDER PROTOCOL (B)(4), AS A RESULT OF WARNING LETTER CMS# (B)(4). | PUMP, INFUSION | SMITHS MEDICAL ASD, INC | 4000 |  | FRN |  |  | No | HEALTH PROFESSIONAL | Pump, Infusion
Malfunction | 20191220 | ['Obstruction of Flow', 'Battery Problem', 'Power Problem'] | INFORMATION RECEIVED BY MEDTRONIC INDICATED THAT THE INSULIN PUMP HAD REJECTED NEW BATTERIES, BATTERY LIFE WAS DIMINISHED FROM FIRST DAY, HAD UNEXPLAINED RANDOM NO DELIVERY ALARM AND IT WAS SLOWER AND LOUDER WHEN BATTERY WAS LOW. NO HARM REQUIRING MEDICAL INTERVENTION WAS REPORTED. THE INSULIN PUMP WILL BE RETURNED FOR ANALYSIS. |  | (B)(4). CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. | PUMP, INFUSION, INSULIN, TO BE USED WITH INVASIVE GLUCOSE SENSOR | MEDTRONIC PUERTO RICO OPERATIONS CO. | MMT-722NAL | A5722NALJ | OYC |  | 20200407 | Device was returned to manufacturer | LAY USER/PATIENT | Pump, Infusion, Insulin, To Be Used With Invasive Glucose Sensor
Malfunction | 20200101 | ['Failure to Interrogate', 'Battery Problem', 'Charging Problem', 'Communication or Transmission Problem', 'Adverse Event Without Identified Device or Use Problem'] | ADDITIONAL INFORMATION WAS RECEIVED FROM A MANUFACTURER¿S REPRESENTATIVE (REP). THE REP REPORTED THAT THEY WERE MEETING WITH THE PATIENT ON (B)(6) 2020 AND THE PATIENT WAS STILL UNABLE TO COMMUNICATE WITH THE INS. THE REP INDICATED THAT THE LAST TIME HE MET WITH THE PATIENT ON (B)(6) 2020-04-13 HE RAN THE PASSIVE RECHARGE FUNCTION 5 TIMES AND DISABLED AND REENABLED THE INS 2 TIMES. NO FURTHER COMPLICATIONS WERE REPORTED. | ADDITIONAL INFORMATION RECEIVED FROM A MANUFACTURER REPRESENTATIVE (REP). IT WAS REPORTED THAT THE INS WAS REPLACED AND THE RECHARGING ISSUE WAS RESOLVED. INFORMATION WAS RECEIVED FROM A CONSUMER AND A MANUFACTURER REPRESENTATIVE REGARDING A PATIENT WHO IS IMPLANTED WITH A NEUROSTIMULAT OR. INFORMATION WAS REPORTED THAT THE PATIENT HAS BEEN GETTING NO DEVICE FOUND FOR A FEW WEEKS. THE PATIENT CONFIRMED IN 2020. THE PATIENT CALLED BACK WHEN THEY HAD THEIR EQUIPMENT. WHILE ON THE CALL, THE PATIENT WITH USING THEIR CONTROLLER AND RECHARGER THERAPY MONITOR (RTM) TO CONNECT TO THE IMPLANT. THE PATIENT STATED SHE IS GETTING THE PASSIVE RECHARGE SCREEN AND THAT'S WHAT SHE WAS GETTING A WEEK AGO. THE MEANING OF THE PASSIVE RECHARGE MODE SCREEN WAS REVIEWED AND IT WAS RECOMMENDED THE PATIENT KEEP THE RTM OVER THE IMPLANT TO COMPLETE THE PASSIVE RECHARGE PROCESS UNTIL THE NORMAL RECHARGING SCREEN COMES UP AND TO CALL IF THERE IS A QUESTION. THEY CALLED BACK ON (B)(6) STATING THAT SHE WAS HAVING THE SAME ISSUES WHERE SHE CANNOT CHARGE HER INS AND HER INS HAS BEEN DEPLETED"FOR A WHILE". SHE WOULD LIKE TO HAVE A REP BE AT THE UPCOMING DOCTOR'S APPOINTMENT IN ORDER TO ASSIST HER BECAUSE EVEN AFTER TROUBLESHOOTING DONE IN THIS CASE, SHE WAS STILL HAVING THE SAME ISSUE. ADDITIONAL INFORMATION RECEIVED FROM A MANUFACTURER REPRESENTATIVE AND A CONSUMER ON (B)(6) 2020 STATED THAT THE INS COULD STILL NOT BE CHARGED. THEY HAD GONE THROUGH PASSIVE RECHARGE MODE THREE TIMES AND THE IMPLANT DID NOT GO BACK TO HER NORMAL CHARGING SCREEN. THE ANTENNA LOCATE WAS TRIED, AND THE VALUES WERE : 73-73-73 (IN AIR) 73 -100-100 (ON IMPLANT). SOFTWARE PROBLEM 1. INTELLIS 2. 3.6 3. 58 4. QF_NEW WAS SEEN. IT WAS REPORTED THE CONTROLLER HAD BEEN RESET, AND THE INS HAD BEEN DEAD FOR A WHILE. THE INS WAS NOT ABLE TO BE CHARGED. THE PATIENT REPORTED SHE HAD AN APPOINTMENT WITH HER HCP. ADDITIONAL INFORMATION RECEIVED VIA A REP STATED THAT THE PATIENT HAD AN APPOINTMENT WITH HER HEALTHCARE PROVIDER (HCP) AND THE REP REQUESTED ADDITIONAL INFORMATION ABOUT RESTARTING THE INS. IT WAS REPORTED THAT THE PATIENT HAD BEEN UNABLE TO CHARGE IN TWO MONTHS, THEY WERE STUCK IN PMR AND THE DEVICE NOT FOUND AND SOFTWARE PROBLEM SCREENS WOULD BE SEEN. ADDITIONAL INFORMATION RECEIVED STATED THAT 3 PASSIVE RECHARGE SESSIONS WERE DONE, BUT CHARGING WAS NOT POSSIBLE. THE REP PLANNED TO DISABLE AND REENABLE THE DEVICE BUT THE ISSUE WAS NOT RESOLVED. THE SOFTWARE PROBLEM MESSAGE WAS STILL SEEN. IT WAS REPORTED THE RECHARGE TELEMETRY MODULE WAS NOT CHARGING THE INS. PATIENT WAS ISSUED A NEW RECHARGE TELEMETRY MODULE. THEY CALLED BACK ON (B)(6) AND STATED THAT THEY SAW NO DEVICE FOUND WITH THE RECHARGER. THEY WERE SENT A PATIENT CONTROLLER. ADDITIONAL INFORMATION WAS RECEIVED FROM A CONSUMER AND A MANUFACTURER REPRESENTATIVE REGARDING THE PATIENT ON(B)(6) 2020. IT WAS REPORTED THAT THE PATIENT RECEIVED THE REPLACEMENT PATIENT CONTROLLER AND SHE WAS STILL GETTING THE NO DEVICE FOUND SCREEN, SHE WAS UNABLE TO RECHARGE, AND THEN SHE NEEDED TO FINISH SETTING UP THE PATIENT CONTROLLER. IT WAS NOTED THAT ALL OF THE EXTERNAL EQUIPMENT HAD BEEN REPLACED. THE PATIENT WAS GOING TO CONTACT HER MANUFACTURER REPRESENTATIVE AND HEALTH CARE PROFESSIONAL TO DISCUSS NEXT STEPS. THE PATIENT NOTED THAT THEY HAVE NOT BEEN ABLE TO RECOVER THE DEVICE AND THAT THERE ARE NO FURTHER ATTEMPTS ARE PLANNED TO TRY AND RESOLVE THE ISSUE. SURGICAL INTERVENTION HAS NOT YET BEEN PLANNED TO RESOLVE THE ISSUE. THERE WERE NO FURTHER COMPLICATIONS REPORTED. ADDITIONAL INFORMATION WAS RECEIVED FROM A MANUFACTURER REPRESENTATIVE (REP). IT WAS REPORTED THAT THE TECHNICAL SERVICE SPECIALIST (TSS) AND ENGINEERING TEAM SPOKE WITH REP TODAY PRIOR TO MEETING WITH PATIENT. TSS REVIEWED FACTORS AROUND THE SITUATION: SO FAR, MULTIPLE PASSIVE CHARGING SESSIONS HAD BEEN DONE OVER THE PAST MONTH AND MULTIPLE DISABLE/RE-ENABLE ATTEMPTS HAD BEEN MADE WITHOUT RESOLUTION. THE RTM WAS REPLACED EARLIER IN APRIL BUT THAT HAD NOT RESOLVED THE TELEMETRY ISSUE EITHER. TSS COACHED THE REP TO TRY PLACING A NEW CONTROLLER CLOSE TO THE INS WHILE WORKING WITH THE PATIENT. TSS SPOKE AGAIN LATER WITH THE REP WHILE HE WAS WITH THE PATIENT AND THEY ATTEMPTED 3 PASSIVE RECHARGE SESSIONS (THEY WERE GETTING NUMBERS AROUND 98 AND 100) BUT THIS DIDN'T RESOLVE THE TELEMETRY ISSUE. 3 DISABLE AND RE-ENABLE COMMANDS WERE TRIED BUT THIS DIDN'T RESOLVE THE ISSUE EITHER. THIS WAS ALL DONE USING THE PATIENT'S NEW RTM, A NEW CONTROLLER AND THE PATIENT'S EXISTING LI BATTERY PACK. NO FURTHER COMPLICATIONS WERE REPORTED. ADDITIONAL INFORMATION WAS RECEIVED FROM REP WHO STATED THAT A BOX WAS ORDERED TO RETURN IPG TODAY AND THAT REP WILL UPDATE ONCE IN THE MAIL. ADDITIONAL INFORMATION WAS RECEIVED FROM THE REP ON (B)(6) 2020. THEY REPORTED INFORMATION THAT HAS ALREADY BEEN DOCUMENTED IN THIS FILE REGARDING THE INS NOT BEING ABLE TO BE PULLED OUT OF A PASSIVE MODE RECHARGE (PMR) STATE. THEY CONFIRMED THAT THERE CURRENTLY WERE NO PLANS TO EXPLANT THE INS. SINCE THE PATIENT IS ABLE TO GET PAIN RELIEF FROM THEIR PUMP DEVICE, THEY HAVE DECIDED TO KEEP THE INS IMPLANTED BUT JUST NOT USE IT, SINCE THEY WERE UNABLE TO GET IT TO CHARGE. THE REP STATED THAT IF SURGICAL INTERVENTION TO REMOVE THE INS IS PLANNED OR SCHEDULED, THEY WILL PROVIDE AN UPDATE. NO FURTHER COMPLICATIONS WERE REPORTED OR ANTICIPATED. ADDITIONAL INFORMATION WAS RECEIVED FROM THE PATIENT. IT WAS REPORTED THEY HAD THEIR INS REPLACED IN 2019 THE (CURRENT) INS ONLY WORKED FOR "MAYBE 4-5 MONTHS". PATIENT DIDN'T CHARGE THE INS FOR OVER A WEEK, BUT WHEN SHE ATTEMPTED TO RECHARGE THE INS, THE INS WOULDN'T CHARGE. SHE THEN MET WITH MANUFACTURER REPRESENTATIVES SEVERAL TIMES FOR TROUBLESHOOTING. THROUGH THIS TROUBLESHOOTING, IT WAS DETERMINED THEY HAD A "DEFECTIVE BATTERY" SO THE PATIENT IS SCHEDULED FOR ANOTHER SURGERY TO ADDRESS THE ISSUES, TWO WEEKS FROM (B)(6) 2020. NO FURTHER COMPLICATIONS WERE REPORTED. | PRODUCT ID 97745,SERIAL# (B)(4). PRODUCT TYPE PROGRAMMER, PATIENT PRODUCT ID 97755, SERIAL# (B)(4). PRODUCT TYPE RECHARGER. IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. MEDTRONIC IS SUBMITTING THIS REPORT TO COMPLY WITH FDA REPORTING REGULATIONS UNDER 21 CFR PARTS 4 AND 803. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY MEDTRONIC, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. MEDTRONIC HAS MADE REASONABLE EFFORTS TO OBTAIN MORE COMPLETE INFORMATION AND HAS PROVIDED AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE AN ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY ¿DEFECTS¿ OR HAS ¿MALFUNCTIONED¿. THESE WORDS ARE INCLUDED IN THE FDA 3500A FORM AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REGULATORY REPORTING. MEDTRONIC OBJECTS TO THE USE OF THESE WORDS AND OTHERS LIKE THEM BECAUSE OF THE LACK OF DEFINITION AND THE CONNOTATIONS IMPLIED BY THESE TERMS. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. A GOOD FAITH EFFORT WILL BE MADE TO OBTAIN THE APPLICABLE INFORMATION RELEVANT TO THE REPORT. IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. MEDTRONIC IS SUBMITTING THIS REPORT TO COMPLY WITH FDA REPORTING REGULATIONS UNDER 21 CFR PARTS 4 AND 803. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY MEDTRONIC, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. MEDTRONIC HAS MADE REASONABLE EFFORTS TO OBTAIN MORE COMPLETE INFORMATION AND HAS PROVIDED AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE AN ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY ¿DEFECTS¿ OR HAS ¿MALFUNCTIONED¿. THESE WORDS ARE INCLUDED IN THE FDA 3500A FORM AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REGULATORY REPORTING. MEDTRONIC OBJECTS TO THE USE OF THESE WORDS AND OTHERS LIKE THEM BECAUSE OF THE LACK OF DEFINITION AND THE CONNOTATIONS IMPLIED BY THESE TERMS. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. A GOOD FAITH EFFORT WILL BE MADE TO OBTAIN THE APPLICABLE INFORMATION RELEVANT TO THE REPORT. IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. H3: ANALYSIS OF THE IMPLANTABLE NEUROSTIMULATOR FOUND AN INS HYBRID ANOMALY WITH EOS/ESD INDUCED DAMAGE. MEDTRONIC IS SUBMITTING THIS REPORT TO COMPLY WITH FDA REPORTING REGULATIONS UNDER 21 CFR PARTS 4 AND 803. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY MEDTRONIC, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. MEDTRONIC HAS MADE REASONABLE EFFORTS TO OBTAIN MORE COMPLETE INFORMATION AND HAS PROVIDED AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE AN ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY ¿DEFECTS¿ OR HAS ¿MALFUNCTIONED¿. THESE WORDS ARE INCLUDED IN THE FDA 3500A FORM AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REGULATORY REPORTING. MEDTRONIC OBJECTS TO THE USE OF THESE WORDS AND OTHERS LIKE THEM BECAUSE OF THE LACK OF DEFINITION AND THE CONNOTATIONS IMPLIED BY THESE TERMS. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. A GOOD FAITH EFFORT WILL BE MADE TO OBTAIN THE APPLICABLE INFORMATION RELEVANT TO THE REPORT. IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. MEDTRONIC IS SUBMITTING THIS REPORT TO COMPLY WITH FDA REPORTING REGULATIONS UNDER 21 CFR PARTS 4 AND 803. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY MEDTRONIC, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. MEDTRONIC HAS MADE REASONABLE EFFORTS TO OBTAIN MORE COMPLETE INFORMATION AND HAS PROVIDED AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE AN ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY ¿DEFECTS¿ OR HAS ¿MALFUNCTIONED¿. THESE WORDS ARE INCLUDED IN THE FDA 3500A FORM AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REGULATORY REPORTING. MEDTRONIC OBJECTS TO THE USE OF THESE WORDS AND OTHERS LIKE THEM BECAUSE OF THE LACK OF DEFINITION AND THE CONNOTATIONS IMPLIED BY THESE TERMS. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. A GOOD FAITH EFFORT WILL BE MADE TO OBTAIN THE APPLICABLE INFORMATION RELEVANT TO THE REPORT. IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. MEDTRONIC IS SUBMITTING THIS REPORT TO COMPLY WITH FDA REPORTING REGULATIONS UNDER 21 CFR PARTS 4 AND 803. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY MEDTRONIC, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. MEDTRONIC HAS MADE REASONABLE EFFORTS TO OBTAIN MORE COMPLETE INFORMATION AND HAS PROVIDED AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE AN ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY ¿DEFECTS¿ OR HAS ¿MALFUNCTIONED¿. THESE WORDS ARE INCLUDED IN THE FDA 3500A FORM AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REGULATORY REPORTING. MEDTRONIC OBJECTS TO THE USE OF THESE WORDS AND OTHERS LIKE THEM BECAUSE OF THE LACK OF DEFINITION AND THE CONNOTATIONS IMPLIED BY THESE TERMS. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. A GOOD FAITH EFFORT WILL BE MADE TO OBTAIN THE APPLICABLE INFORMATION RELEVANT TO THE REPORT. IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. | STIMULATOR, SPINAL-CORD, TOTALLY IMPLANTED FOR PAIN RELIEF | MEDTRONIC PUERTO RICO OPERATIONS CO. | 97715 |  | LGW | 20200528 | 20200909 | Device was returned to manufacturer | HEALTH PROFESSIONAL | Stimulator, Spinal-Cord, Totally Implanted For Pain Relief
Malfunction | 20200427 | ['Device Difficult to Program or Calibrate', 'Obstruction of Flow', 'Battery Problem'] | INFORMATION RECEIVED BY MEDTRONIC INDICATED THAT INSULIN PUMP HAD BUTTONS LESS RESPONSIVE HIT MULTIPLE TIMES AND RANDOM NO DELIVERY. NO HARM REQUIRING MEDICAL INTERVENTION WAS REPORTED. THE CUSTOMER STATED THAT INSULIN PUMP HAD BATTERY DRAINS DAILY AND REJECT BATTERY AND LOST SETTINGS. THE DEVICE WILL NOT BE RETURNED FOR ANALYSIS. |  | (B)(4). CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. | ARTIFICIAL PANCREAS DEVICE SYSTEM, THRESHOLD SUSPEND | MEDTRONIC PUERTO RICO OPERATIONS CO. | MMT-751NAB | A4751NABJ | OZO |  |  | No | LAY USER/PATIENT | Automated Insulin Dosing , Threshold Suspend
Malfunction | 20200426 | ['Battery Problem', 'Power Problem'] | INFORMATION RECEIVED BY MEDTRONIC INDICATED THAT THE INSULIN PUMP RECEIVED REPLACE BATTERY ALERTS. NO HARM REQUIRING MEDICAL INTERVENTION WAS REPORTED. THE INSULIN PUMP WILL NOT BE RETURNED FOR ANALYSIS. |  | (B)(4). CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. | ARTIFICIAL PANCREAS DEVICE SYSTEM, SINGLE HORMONAL CONTROL | MEDTRONIC PUERTO RICO OPERATIONS CO. | MMT-1780KPK | HG3F2K5 | OZP |  |  | No | LAY USER/PATIENT | Automated Insulin Dosing Device System, Single Hormonal Control
Malfunction | 20200427 | ['Mechanical Problem', 'Battery Problem', 'Power Problem'] | INFORMATION RECEIVED BY MEDTRONIC INDICATED THAT THE INSULIN PUMP HAD A MOTOR ERROR CODE DURING PRIMING. THE CUSTOMER STATED THAT THE INSULIN PUMP REJECTED NEW BATTERIES AND THE SOME BATTERIES ONLY LAST A COUPLE DAYS. THE CUSTOMER STATED THAT THE INSULIN PUMP WAS LABORED DURING REWINDING AND SOMETIMES THE TIME HAD RESET AFTER CHANGING THE BATTERY. NO HARM REQUIRING MEDICAL INTERVENTION WAS REPORTED. THE INSULIN PUMP WILL NOT BE RETURNED FOR ANALYSIS. |  | (B)(4). CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. | PUMP, INFUSION, INSULIN, TO BE USED WITH INVASIVE GLUCOSE SENSOR | MEDTRONIC PUERTO RICO OPERATIONS CO. | MMT-723NAS | A7723NASJ | OYC |  |  | No | LAY USER/PATIENT | Pump, Infusion, Insulin, To Be Used With Invasive Glucose Sensor
Injury | 20200407 | ['Battery Problem', 'Missing Test Results'] | IT WAS REPORTED THAT THE PATIENT EXPERIENCED BEING PHYSICALLY ILL AND THE IMPLANTABLE PULSE GENERATOR (IPG) EXHIBITED BATTERY DEPLETION. THE IPG WAS REPROGRAMMED THEN EXPLANTED AND REPLACED. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. |  | PRODUCT EVENT SUMMARY: THE DEVICE WAS RETURNED AND ANALYZED. RETURNED PRODUCT ANALYSIS WAS PERFORMED AND NO ANOMALIES WERE FOUND. IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. CONCOMITANT MEDICAL PRODUCTS: 5076-45, LEAD, IMPLANT DATE: (B)(6) 2010. IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. | PULSE GENERATOR, PERMANENT, IMPLANTABLE | MEDTRONIC PUERTO RICO OPERATIONS CO. | ADDR01 |  | NVZ | 20181114 | 20200413 | Device was returned to manufacturer | LAY USER/PATIENT | Pulse Generator, Permanent, Implantable
Malfunction | 20200126 | ['Battery Problem'] | INFORMATION RECEIVED BY MEDTRONIC INDICATED THAT, THE INSULIN PUMP HAD BATTERIES END AFTER 1-3 DAYS AND THE BATTERY CAP WAS OKAY. NO HARM REQUIRING MEDICAL INTERVENTION WAS REPORTED. THE INSULIN PUMP WILL BE RETURNED FOR ANALYSIS. |  | INSULIN PUMP WAS RECEIVED WITH A PARTIALLY BROKEN RETAINER, MISSING RESERVOIR TUBE O-RING AND BROKEN RESERVOIR TUBE LIP. THE TEST RESERVOIR DOES NOT LOCK IN PLACE AND UNABLE TO PERFORM THE DISPLACEMENT TEST DUE TO PARTIALLY BROKEN RETAINER, MISSING RESERVOIR TUBE O-RING AND BROKEN RESERVOIR TUBE LIP. INSULIN PUMP PASSED SLEEP CURRENT MEASUREMENT, ACTIVE CURRENT MEASUREMENT AND SELF TEST. THE POWER MANAGEMENT TOOL CONFIRMED THE UNLOADED VOLTAGE (UL VLITH) AND LOADED VOLTAGE (LOADED VLITH) WAS WITHIN SPEC RANGE. NO UNEXPECTED LOW BATTERY LIFE OR LOW BATTERY ALERT NOTED DURING TESTING OR IN THE PUMP TRACE DOWNLOAD. (B)(4). THE INSULIN PUMP INVOLVED IN THIS EVENT IS THE 640G INSULIN INFUSION PUMP, WHICH IS NOT MARKETED IN THE UNITED STATES. HOWEVER, THE DEVICE IS SIMILAR TO THE PARADIGM REAL-TIME INSULIN INFUSION PUMP, WHICH IS MARKETED IN THE UNITED STATES. THIS REPORT IS PART OF A RETROSPECTIVE REVIEW AND REMEDIATION EFFORTS IN RESPONSE TO A WARNING LETTER. UPDATED H9: Z-0958-2020. MEDTRONIC, INC. (MEDTRONIC) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY MEDTRONIC, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. MEDTRONIC HAS MADE REASONABLE EFFORTS TO OBTAIN MORE COMPLETE INFORMATION IN THE TIME ALLOTTED AND HAS PROVIDED AS MUCH INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE AN ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY "DEFECTS" OR HAS "MALFUNCTIONED". THESE WORDS ARE INCLUDED IN THE FDA 3500A FORM AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. MEDTRONIC OBJECTS TO THE USE OF THESE WORDS AND OTHERS LIKE IT BECAUSE OF THE LACK OF DEFINITION AND THE CONNOTATIONS IMPLIED BY THESE TERMS. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT. | PUMP, INFUSION, INSULIN, TO BE USED WITH INVASIVE GLUCOSE SENSOR | MEDTRONIC PUERTO RICO OPERATIONS CO. | MMT-1712K | HG3D1C6 | OYC |  | 20200423 | Device was returned to manufacturer | LAY USER/PATIENT | Pump, Infusion, Insulin, To Be Used With Invasive Glucose Sensor
Malfunction | 20200425 | ['Battery Problem'] | INFORMATION RECEIVED BY MEDTRONIC INDICATED THAT THE INSULIN PUMP HAD POWER ERROR DETECTED ALARM. NO HARM REQUIRING MEDICAL INTERVENTION WAS REPORTED. THE INSULIN PUMP WILL NOT BE RETURNED FOR ANALYSIS. |  | (B)(4). CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. THE INSULIN PUMP INVOLVED IN THIS EVENT IS THE 640G INSULIN INFUSION PUMP, WHICH IS NOT MARKETED IN THE UNITED STATES. HOWEVER, THE DEVICE IS SIMILAR TO THE PARADIGM REAL-TIME INSULIN INFUSION PUMP, WHICH IS MARKETED IN THE UNITED STATES. | PUMP, INFUSION, INSULIN, TO BE USED WITH INVASIVE GLUCOSE SENSOR | MEDTRONIC PUERTO RICO OPERATIONS CO. | MMT-1711K | HG0M6D9 | OYC |  |  | No | LAY USER/PATIENT | Pump, Infusion, Insulin, To Be Used With Invasive Glucose Sensor
Malfunction | 20200131 | ['Battery Problem'] | INFORMATION RECEIVED BY MEDTRONIC INDICATED THAT THE INSULIN PUMP HAD A POWER ERROR DETECTED ALARM. POWER ERROR DETECTED ALARM RECURRED WITHIN A WEEK OF TROUBLESHOOTING THE PREVIOUS OCCURRENCE. NO HARM REQUIRING MEDICAL INTERVENTION WAS REPORTED. THE INSULIN PUMP WILL BE RETURNED FOR ANALYSIS. |  | INSULIN PUMP TRACE DOWNLOAD ANALYSIS CONFIRMED THE UNIT ALARMED PUMP ERROR 25. THE POWER MANAGEMENT TOOL CONFIRMED PUMP ERROR 25 WAS TRIGGERED WHEN THE BACKUP BATTERY LOADED VOLTAGE (LOADED VLITH) WAS LESS THAN 3.5V FOR 4 CONSECUTIVE HOURS DUE TO CONNECTOR RESISTANCE J6 OR PCB 1. (B)(4). | ARTIFICIAL PANCREAS DEVICE SYSTEM, SINGLE HORMONAL CONTROL | MEDTRONIC PUERTO RICO OPERATIONS CO. | MMT-1780KPK | HG418H4 | OZP |  | 20200309 | Device was returned to manufacturer | LAY USER/PATIENT | Automated Insulin Dosing Device System, Single Hormonal Control
Malfunction | 20200214 | ['Battery Problem', 'Patient Device Interaction Problem'] | INFORMATION RECEIVED BY MEDTRONIC INDICATED THAT RESERVOIR LIP RING WAS MISSING. THE CUSTOMER STATED THAT THE RESERVOIR DID LOCK INTO PLACE. NO HARM REQUIRING MEDICAL INTERVENTION WAS REPORTED. THE INSULIN PUMP WILL BE RETURNED FOR ANALYSIS. |  | UNIT RECEIVED WITH MISSING RETAINER, MISSING RESERVOIR TUBE O-RING AND BROKEN RESERVOIR TUBE LIP. UNABLE TO PERFORM DISPLACEMENT TEST OR LOCK RESERVOIR IN PLACE DUE TO MISSING RETAINER. UNIT PASSED THE SLEEP CURRENT MEASUREMENT AND ACTIVE CURRENT MEASUREMENT TEST. NO LOW BATTERY ALARMS OR UNEXPECTED BATTERY POWER LOSS NOTED DURING TESTING. UNIT FUNCTIONING PROPERLY. THE POWER MANAGEMENT PARAMETERS GRAPH CONFIRMED THE UNLOADED VOLTAGE (UNLOADED VLITH) AND LOADED VOLTAGE (LOADED VLITH) WAS WITHIN SPEC RANGE. PUMP RECEIVED WITH NORMAL OPERATING CURRENTS. (B)(4). THE INSULIN PUMP INVOLVED IN THIS EVENT IS THE 640G INSULIN INFUSION PUMP, WHICH IS NOT MARKETED IN THE UNITED STATES. HOWEVER, THE DEVICE IS SIMILAR TO THE PARADIGM REAL-TIME INSULIN INFUSION PUMP, WHICH IS MARKETED IN THE UNITED STATES. THIS REPORT IS PART OF A RETROSPECTIVE REVIEW AND REMEDIATION EFFORTS IN RESPONSE TO A WARNING LETTER. UPDATED H9: Z-0958-2020. MEDTRONIC, INC. (MEDTRONIC) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY MEDTRONIC, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. MEDTRONIC HAS MADE REASONABLE EFFORTS TO OBTAIN MORE COMPLETE INFORMATION IN THE TIME ALLOTTED AND HAS PROVIDED AS MUCH INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE AN ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY "DEFECTS" OR HAS "MALFUNCTIONED". THESE WORDS ARE INCLUDED IN THE FDA 3500A FORM AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. MEDTRONIC OBJECTS TO THE USE OF THESE WORDS AND OTHERS LIKE IT BECAUSE OF THE LACK OF DEFINITION AND THE CONNOTATIONS IMPLIED BY THESE TERMS. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT. | PUMP, INFUSION, INSULIN, TO BE USED WITH INVASIVE GLUCOSE SENSOR | MEDTRONIC PUERTO RICO OPERATIONS CO. | MMT-1711K | HG1WGU1 | OYC |  | 20200317 | Device was returned to manufacturer | LAY USER/PATIENT | Pump, Infusion, Insulin, To Be Used With Invasive Glucose Sensor
Malfunction | 20200420 | ['Battery Problem'] | INFORMATION WAS RECEIVED FROM A HEALTHCARE PROFESSIONAL (HCP) VIA MANUFACTURING REPRESENTATIVE (REP) REGARDING A PATIENT WHO WAS IMPLANTED WITH AN IMPLANTABLE NEUROSTIMULATOR (INS). IT WAS REPORTED THAT THE PATIENT USED THE PATIENT PROGRAMMER TO CHECK THE INS AND THEY SAW POWER ON RESET (POR) SCREEN. THE PATIENT FOLLOWED THE INSTRUCTIONS TO RESTORE THE THERAPY, AND THERAPY HAS BEEN WORKING PROPERLY SINCE. THE PATIENT DID NOT WANT TO GO TO THE CLINIC DUE TO COVID-19 AND WILL GO AFTER IT IS DONE. THE ISSUE WAS RESOLVED AT THE TIME OF THE REPORT. NO SYMPTOMS WERE REPORTED. NO FURTHER COMPLICATIONS WERE REPORTED/ANTICIPATED. |  | MEDTRONIC IS SUBMITTING THIS REPORT TO COMPLY WITH FDA REPORTING REGULATIONS UNDER 21 CFR PARTS 4 AND 803. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY MEDTRONIC, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. MEDTRONIC HAS MADE REASONABLE EFFORTS TO OBTAIN MORE COMPLETE INFORMATION AND HAS PROVIDED AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE AN ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY ¿DEFECTS¿ OR HAS ¿MALFUNCTIONED¿. THESE WORDS ARE INCLUDED IN THE FDA 3500A FORM AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REGULATORY REPORTING. MEDTRONIC OBJECTS TO THE USE OF THESE WORDS AND OTHERS LIKE THEM BECAUSE OF THE LACK OF DEFINITION AND THE CONNOTATIONS IMPLIED BY THESE TERMS. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT. A GOOD FAITH EFFORT WILL BE MADE TO OBTAIN THE APPLICABLE INFORMATION RELEVANT TO THE REPORT. IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. | STIMULATOR, ELECTRICAL, IMPLANTED, FOR PARKINSONIAN TREMOR | MEDTRONIC PUERTO RICO OPERATIONS CO. | 37602 |  | MRU | 20190728 |  | No | HEALTH PROFESSIONAL | Implanted Subcortical Electrical Stimulator (Motor Disorders)
Malfunction | 20200227 | ['Battery Problem', 'Material Integrity Problem', 'Power Problem', 'Patient Device Interaction Problem'] | INFORMATION RECEIVED BY MEDTRONIC INDICATED THAT THE INSULIN PUMP'S RETAINER RING WAS CRACKED. NO HARM REQUIRING MEDICAL INTERVENTION WAS REPORTED. THE RESERVOIR WAS ABLE TO LOCK IN PLACE WHEN INSERTED INTO THE INSULIN PUMP. CUSTOMER ALSO REPORTED CAP WAS DAMAGED AND FAILED BATTERY ALARM. THE CUSTOMER WAS ADVISED TO DISCONTINUE USE OF THE INSULIN PUMP AND REVERT TO THE BACKUP PLAN. THE INSULIN PUMP WILL BE RETURNED FOR ANALYSIS. |  | THIS REPORT IS PART OF A RETROSPECTIVE REVIEW AND REMEDIATION EFFORTS IN RESPONSE TO A WARNING LETTER. UPDATED H9: Z-0955-2020. MEDTRONIC, INC. (MEDTRONIC) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY MEDTRONIC, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. MEDTRONIC HAS MADE REASONABLE EFFORTS TO OBTAIN MORE COMPLETE INFORMATION IN THE TIME ALLOTTED AND HAS PROVIDED AS MUCH INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE AN ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY "DEFECTS" OR HAS "MALFUNCTIONED". THESE WORDS ARE INCLUDED IN THE FDA 3500A FORM AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. MEDTRONIC OBJECTS TO THE USE OF THESE WORDS AND OTHERS LIKE IT BECAUSE OF THE LACK OF DEFINITION AND THE CONNOTATIONS IMPLIED BY THESE TERMS. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT. UNIT RECEIVED WITH MISSING RETAINER, MISSING RESERVOIR TUBE O-RING AND BROKEN RESERVOIR TUBE LIP. UNABLE TO PERFORM DISPLACEMENT TEST OR LOCK RESERVOIR IN PLACE DUE TO MISSING RETAINER. UNIT PASSED THE SLEEP CURRENT MEASUREMENT AND ACTIVE CURRENT MEASUREMENT TEST. NO LOW BATTERY ALARMS,FAILED BATT OR UNEXPECTED BATTERY POWER LOSS NOTED DURING TESTING. UNIT FUNCTIONING PROPERLY. THE POWER MANAGEMENT PARAMETERS GRAPH CONFIRMED THE UNLOADED VOLTAGE (UNLOADED VLITH) AND LOADED VOLTAGE (LOADED VLITH) WAS WITHIN SPEC RANGE. PUMP RECEIVED WITH NORMAL OPERATING CURRENTS. (B)(4). | ARTIFICIAL PANCREAS DEVICE SYSTEM, THRESHOLD SUSPEND | MEDTRONIC PUERTO RICO OPERATIONS CO. | MMT-1715KL | HG2D3ZP | OZO | 20180907 | 20200311 | Device was returned to manufacturer | LAY USER/PATIENT | Automated Insulin Dosing , Threshold Suspend
Malfunction | 20200303 | ['Battery Problem', 'Power Problem'] | INFORMATION RECEIVED BY MEDTRONIC INDICATED THAT THE INSULIN PUMP HAD REPLACE BATTERY ALARM, POWER LOSS ALARM AND LOW BATTERY ALARM. THE CUSTOMER WAS ABLE TO CLEAR THE ALARM AND REWIND THE INSULIN PUMP. NO HARM REQUIRING MEDICAL INTERVENTION WAS REPORTED. THE INSULIN PUMP WILL BE RETURNED FOR ANALYSIS. |  | THIS REPORT IS PART OF A RETROSPECTIVE REVIEW AND REMEDIATION EFFORTS IN RESPONSE TO A WARNING LETTER. UPDATED H9: Z-0956-2020 MEDTRONIC, INC. (MEDTRONIC) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY MEDTRONIC, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. MEDTRONIC HAS MADE REASONABLE EFFORTS TO OBTAIN MORE COMPLETE INFORMATION IN THE TIME ALLOTTED AND HAS PROVIDED AS MUCH INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE AN ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY "DEFECTS" OR HAS "MALFUNCTIONED". THESE WORDS ARE INCLUDED IN THE FDA 3500A FORM AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. MEDTRONIC OBJECTS TO THE USE OF THESE WORDS AND OTHERS LIKE IT BECAUSE OF THE LACK OF DEFINITION AND THE CONNOTATIONS IMPLIED BY THESE TERMS. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT. UNIT PASSED THE DISPLACEMENT TEST. SLEEP CURRENT MEASUREMENT AND ACTIVE CURRENT MEASUREMENT WITHIN SPECIFICATIONS. THE POWER MANAGEMENT TOOL CONFIRMED POWER ERROR 25, LOW BATTERY ALERT, POWER LOSS ALARM, WAS TRIGGERED WHEN THE BACKUP BATTERY LOADED VOLTAGE (LOADED VLITH) WAS LESS THAN 3.5V FOR 4 CONSECUTIVE HOURS DUE TO CONNECTOR RESISTANCE J6 OR PCB 1. AFTER DISCONNECTING AND RECONNECTING THE INTERNAL BATTERY CONNECTOR ON J6 OR PCB 1, THE PUMP WAS MONITORED AND FUNCTIONED PROPERLY. UNIT RECEIVED WITH CRACKED RETAINER AND PILLOWING KEYPAD OVERLAY. (B)(4). | ARTIFICIAL PANCREAS DEVICE SYSTEM, SINGLE HORMONAL CONTROL | MEDTRONIC PUERTO RICO OPERATIONS CO. | MMT-1780K | HG25WZ7 | OZP | 20180602 | 20200320 | Device was returned to manufacturer | LAY USER/PATIENT | Automated Insulin Dosing Device System, Single Hormonal Control
Malfunction |  | ['Battery Problem'] | PER BIOMED, THE BATTERY WON'T HOLD A CHARGE. ON (B)(6) 2020 - SAMPLE EVALUATION CONFIRMED BATTERY ABRUPT SHUTDOWN. |  | THE FOLLOWING WERE REVIEWED AS PART OF THIS INVESTIGATION: PATIENT SEVERITY, APPLICABLE PREVIOUS INVESTIGATION(S), SAMPLE (IF AVAILABLE), APPLICABLE FMEA DOCUMENTS, LABELING, AND APPLICABLE MANUFACTURE RECORDS. BASED ON A REVIEW OF THIS INFORMATION, THE FOLLOWING WAS CONCLUDED: THE DEVICE WAS RETURNED TO THE SERVICE FACILITY FOR EVALUATION. DURING EVALUATION, THE REPORTED ISSUE BATTERY WON'T HOLD A CHARGE WAS CONFIRMED. THE SCANNER WAS INSPECTED THE UNIT SHUTS DOWN WHEN UNPLUGGED FROM THE AC ADAPTER, THE CAUSE OF THE ISSUE IS DUE TO INTERNAL BATTERY FAILURE. THE DEVICE WAS SERVICED, TESTED AND RETURNED TO THE CUSTOMER. | SYSTEM, IMAGING, PULSED ECHO, ULTRASONIC | BARD ACCESS SYSTEMS | 9770500 | N/A | IYO |  | 20200115 | Device was returned to manufacturer | HEALTH PROFESSIONAL | System, Imaging, Pulsed Echo, Ultrasonic
Malfunction |  | ['Battery Problem'] | PER BIOMED, UNIT WILL NOT TURN ON. GREEN BATTERY LIGHT IS PRESENT BUT WILL NOT BOOT UP. ON (B)(6) 2020 - SAMPLE EVALUATION CONFIRMED BATTERY ABRUPT SHUTDOWN. |  | THE FOLLOWING WERE REVIEWED AS PART OF THIS INVESTIGATION: PATIENT SEVERITY, APPLICABLE PREVIOUS INVESTIGATION(S), SAMPLE (IF AVAILABLE), APPLICABLE FMEA DOCUMENTS, LABELING, AND APPLICABLE MANUFACTURE RECORDS. BASED ON A REVIEW OF THIS INFORMATION, THE FOLLOWING WAS CONCLUDED: THE DEVICE WAS RETURNED TO THE SERVICE FACILITY FOR EVALUATION. DURING EVALUATION, THE REPORTED ISSUE OF UNIT WILL NOT TURN ON. GREEN BATTERY LIGHT IS PRESENT BUT WILL NOT BOOT UP WAS CONFIRMED. THE SITE RITE 8 WAS VISUALLY INSPECTED UPON RECEIPT AND WAS FOUND TO BE IN GOOD PHYSICAL CONDITION. THE ROOT CAUSE OF THE REPORTED ISSUE WAS IDENTIFIED AS AN INTERNAL LI-ION BATTERY FAILURE. THE DEVICE WAS SERVICED, TESTED, AND RETURNED TO THE CUSTOMER. A HISTORY REVIEW OF SERIAL NUMBER (B)(4) SHOWED NO OTHER SIMILAR PRODUCT COMPLAINT(S) FROM THIS SERIAL NUMBER. | SYSTEM, IMAGING, PULSED ECHO, ULTRASONIC | BARD ACCESS SYSTEMS | 9770600 | N/A | IYO |  | 20200113 | Device was returned to manufacturer | HEALTH PROFESSIONAL | System, Imaging, Pulsed Echo, Ultrasonic
Malfunction |  | ['Battery Problem'] | SAMPLE EVALUATION CONFIRMED BATTERY ABRUPT SHUTDOWN. |  | THE FOLLOWING WERE REVIEWED AS PART OF THIS INVESTIGATION: PATIENT SEVERITY, APPLICABLE PREVIOUS INVESTIGATION(S), SAMPLE (IF AVAILABLE), APPLICABLE FMEA DOCUMENTS, LABELING, AND APPLICABLE MANUFACTURE RECORDS. BASED ON A REVIEW OF THIS INFORMATION, THE FOLLOWING WAS CONCLUDED: THE SR8 WAS RECEIVED WITH PHYSICAL DAMAGE NOTED. THE SR8 WAS POWERED ON AT 67% BATTERY LIFE AND RAN FOR AROUND TEN MINUTES ON BATTERY POWER BEFORE SHUTTING DOWN WITH 64% BATTERY LIFE STILL ON THE SCANNER. THE ROOT CAUSE OF THE REPORTED ISSUE IS THE INTERNAL BATTERY FAILED. THE DEVICE WAS SERVICED, TESTED AND RETURNED TO THE CUSTOMER. A HISTORY REVIEW OF SERIAL NUMBER (B)(4) SHOWED NO OTHER SIMILAR PRODUCT COMPLAINT(S) FROM THIS SERIAL NUMBER. | SYSTEM, IMAGING, PULSED ECHO, ULTRASONIC | BARD ACCESS SYSTEMS |  | N/A | IYO |  | 20191220 | Device was returned to manufacturer | HEALTH PROFESSIONAL | System, Imaging, Pulsed Echo, Ultrasonic
Malfunction |  | ['Battery Problem'] | ON (B)(6) 2020 - SAMPLE EVALUATION CONFIRMED BATTERY ABRUPT SHUTDOWN. |  | THE FOLLOWING WERE REVIEWED AS PART OF THIS INVESTIGATION: PATIENT SEVERITY, APPLICABLE PREVIOUS INVESTIGATION(S), SAMPLE (IF AVAILABLE), APPLICABLE FMEA DOCUMENTS, LABELING, AND APPLICABLE MANUFACTURE RECORDS. BASED ON A REVIEW OF THIS INFORMATION, THE FOLLOWING WAS CONCLUDED: THE SITE RITE 8 WAS VISUALLY INSPECTED UPON RECEIPT AND WAS FOUND TO BE IN GOOD PHYSICAL CONDITION. DURING EVALUATION, THE REPORTED ISSUE OF BATTERY WILL NOT HOLD A CHARGE WAS CONFIRMED. THE ROOT CAUSE OF THE REPORTED ISSUE IS AN INTERNAL LI-ION BATTERY FAILURE. THE DEVICE WAS SERVICED, TESTED, AND RETURNED TO THE CUSTOMER. A HISTORY REVIEW OF SERIAL NUMBER (B)(4) SHOWED NO OTHER SIMILAR PRODUCT COMPLAINT(S) FROM THIS SERIAL NUMBER. | SYSTEM, IMAGING, PULSED ECHO, ULTRASONIC | BARD ACCESS SYSTEMS | 9770600Y | N/A | IYO |  | 20191216 | Device was returned to manufacturer | HEALTH PROFESSIONAL | System, Imaging, Pulsed Echo, Ultrasonic
Malfunction |  | ['Battery Problem'] | PER OFFICE MANAGER THE BATTERY WILL NOT HOLD A CHARGE. ON (B)(6) 2020 - SAMPLE EVALUATION CONFIRMED ABRUPT SHUTDOWN. |  | THE FOLLOWING WERE REVIEWED AS PART OF THIS INVESTIGATION: PATIENT SEVERITY, APPLICABLE PREVIOUS INVESTIGATION(S), SAMPLE (IF AVAILABLE), APPLICABLE FMEA DOCUMENTS, LABELING, AND APPLICABLE MANUFACTURE RECORDS. BASED ON A REVIEW OF THIS INFORMATION, THE FOLLOWING WAS CONCLUDED: THE DEVICE WAS RETURNED TO THE SERVICE FACILITY FOR EVALUATION. DURING EVALUATION, THE REPORTED ISSUE OF THE BATTERY WILL NOT HOLD A CHARGE IS CONFIRMED. THE ROOT CAUSE OF THE UNIT NOT HOLDING A CHARGE IS THE INTERNAL BATTERY FAILED. THE DEVICE WAS SERVICED, TESTED AND RETURNED TO THE CUSTOMER. A HISTORY REVIEW OF SERIAL NUMBER (B)(4) SHOWED NO OTHER SIMILAR PRODUCT COMPLAINT(S) FROM THIS SERIAL NUMBER. | SYSTEM, IMAGING, PULSED ECHO, ULTRASONIC | BARD ACCESS SYSTEMS | 9770600Y | N/A | IYO |  | 20200219 | Device was returned to manufacturer | HEALTH PROFESSIONAL | System, Imaging, Pulsed Echo, Ultrasonic
Malfunction |  | ['Battery Problem'] | ON (B)(6) 2020 - DURING SAMPLE EVALUATION BATTERY ABRUPT SHUTDOWN WAS CONFIRMED. |  | THE FOLLOWING WERE REVIEWED AS PART OF THIS INVESTIGATION: PATIENT SEVERITY, APPLICABLE PREVIOUS INVESTIGATION(S), SAMPLE (IF AVAILABLE), APPLICABLE FMEA DOCUMENTS, LABELING, AND APPLICABLE MANUFACTURE RECORDS. BASED ON A REVIEW OF THIS INFORMATION, THE FOLLOWING WAS CONCLUDED: THE SR8 WAS RECEIVED WITH PHYSICAL DAMAGE NOTED. THE SR8 RANDOMLY SHUTS DOWN ON BATTERY POWER WAS CONFIRMED. THE ROOT CAUSE OF THE REPORTED ISSUE IS THE INTERNAL BATTERY FAILED. THE DEVICE WAS SERVICED, TESTED, AND RETURNED TO THE CUSTOMER. A HISTORY REVIEW OF SERIAL NUMBER (B)(4) SHOWED NO OTHER SIMILAR PRODUCT COMPLAINT(S) FROM THIS SERIAL NUMBER. | SYSTEM, IMAGING, PULSED ECHO, ULTRASONIC | BARD ACCESS SYSTEMS | 9770501 | N/A | IYO |  | 20200204 | Device was returned to manufacturer | HEALTH PROFESSIONAL | System, Imaging, Pulsed Echo, Ultrasonic
Malfunction |  | ['Battery Problem'] | PER BIOMED, UNIT WILL NOT CHARGE. ON (B)(6) 2020 - SAMPLE EVALUATION CONFIRMED BATTERY ABRUPT SHUTDOWN. |  | THE FOLLOWING WERE REVIEWED AS PART OF THIS INVESTIGATION: PATIENT SEVERITY, APPLICABLE PREVIOUS INVESTIGATION(S), SAMPLE (IF AVAILABLE), APPLICABLE FMEA DOCUMENTS, LABELING, AND APPLICABLE MANUFACTURE RECORDS. BASED ON A REVIEW OF THIS INFORMATION, THE FOLLOWING WAS CONCLUDED: THE DEVICE WAS RETURNED TO THE SERVICE FACILITY FOR EVALUATION. DURING EVALUATION, THE REPORTED ISSUE OF UNIT WILL NOT CHARGE WAS CONFIRMED. THE SR8 WAS RECEIVED WITH PHYSICAL DAMAGE NOTED. THE SR8 WAS POWERED ON AT 77% BATTERY LIFE AND WHEN UNPLUGGED FROM THE AC ADAPTER THE SR8 SHUTS DOWN. THE ROOT CAUSE OF THE REPORTED ISSUE IS AN INTERNAL BATTERY FAILURE. THE DEVICE WAS SERVICED, TESTED, AND RETURNED TO THE CUSTOMER. A HISTORY REVIEW OF SERIAL NUMBER (B)(4) SHOWED NO OTHER SIMILAR PRODUCT COMPLAINT(S) FROM THIS SERIAL NUMBER. | SYSTEM, IMAGING, PULSED ECHO, ULTRASONIC | BARD ACCESS SYSTEMS | 9770600Y | N/A | IYO |  | 20200214 | Device was returned to manufacturer | HEALTH PROFESSIONAL | System, Imaging, Pulsed Echo, Ultrasonic
Malfunction |  | ['Battery Problem'] | PER BIOMED, BATTERY WILL NOT HOLD A CHARGE. ON (B)(6) 2020 - SAMPLE EVALUATION CONFIRMED BATTERY ABRUPT SHUTDOWN. |  | THE FOLLOWING WERE REVIEWED AS PART OF THIS INVESTIGATION: PATIENT SEVERITY, APPLICABLE PREVIOUS INVESTIGATION(S), SAMPLE (IF AVAILABLE), APPLICABLE FMEA DOCUMENTS, LABELING, AND APPLICABLE MANUFACTURE RECORDS. BASED ON A REVIEW OF THIS INFORMATION, THE FOLLOWING WAS CONCLUDED: THE DEVICE WAS RETURNED TO THE SERVICE FACILITY FOR EVALUATION. DURING EVALUATION, THE REPORTED ISSUE OF BATTERY WILL NOT HOLD CHARGE WAS CONFIRMED. THE ROOT CAUSE OF THE REPORTED ISSUE IS THE INTERNAL BATTERY FAILED. A HISTORY REVIEW OF SERIAL NUMBER (B)(4) SHOWED NO OTHER SIMILAR PRODUCT COMPLAINT(S) FROM THIS SERIAL NUMBER. | SYSTEM, IMAGING, PULSED ECHO, ULTRASONIC | BARD ACCESS SYSTEMS | 9770600 | N/A | IYO |  | 20200219 | Device was returned to manufacturer | HEALTH PROFESSIONAL | System, Imaging, Pulsed Echo, Ultrasonic
Malfunction |  | ['Battery Problem'] | PER BIOMED THE SCANNER IS NOT HOLDING A CHARGE. ON (B)(6) 2020 - SAMPLE EVALUATION CONFIRMED BATTERY ABRUPT SHUTDOWN. |  | THE FOLLOWING WERE REVIEWED AS PART OF THIS INVESTIGATION: PATIENT SEVERITY, APPLICABLE PREVIOUS INVESTIGATION(S), SAMPLE (IF AVAILABLE), APPLICABLE FMEA DOCUMENTS, LABELING, AND APPLICABLE MANUFACTURE RECORDS. BASED ON A REVIEW OF THIS INFORMATION, THE FOLLOWING WAS CONCLUDED: THE DEVICE WAS RETURNED TO THE SERVICE FACILITY FOR EVALUATION. DURING EVALUATION, THE REPORTED ISSUE OF THE SCANNER IS NOT HOLDING A CHARGE WAS CONFIRMED. THE ROOT CAUSE OF THE REPORTED ISSUE WAS IDENTIFIED AS AN INTERNAL LI-ION BATTERY FAILURE. THE SITE RITE 8 WAS VISUALLY INSPECTED UPON RECEIPT AND WAS FOUND TO BE IN GOOD PHYSICAL CONDITION. THE DEVICE WAS SERVICED, TESTED, AND RETURNED TO THE CUSTOMER. A HISTORY REVIEW OF SERIAL NUMBER (B)(4) SHOWED NO OTHER SIMILAR PRODUCT COMPLAINT(S) FROM THIS SERIAL NUMBER. | SYSTEM, IMAGING, PULSED ECHO, ULTRASONIC | BARD ACCESS SYSTEMS | 9770600 | NA | IYO |  | 20200302 | Device was returned to manufacturer | HEALTH PROFESSIONAL | System, Imaging, Pulsed Echo, Ultrasonic
Malfunction | 20200422 | ['Battery Problem'] | ADDITIONAL INFORMATION WAS RECEIVED FROM THE REP: THE REP WASN'T MADE AWARE OF THE POR AND CHECKED WITH THE HCP ACCOUNT AND THEY HAVE THE PATIENT SCHEDULED TO COME IN TO BE SEEN THIS WEEK. THE REP WILL FOLLOW UP ONCE THEY DETERMINE WHAT CAUSED THE POR. THIS PATIENT IS ON THE PRODUCT REPLACEMENT LIST AND HAD BEEN FOR SOME TIME, SO IT¿S LIKELY DUE TO NORMAL BATTERY DEPLETION. PATIENT ALSO HAS HOSPICE SO END OF LIFE IS NEAR, SO THE REP WAS NOT SURE IF PATIENT WILL MAKE IT TO APPOINTMENT TO BE SEEN AT THIS TIME. NO FURTHER COMPLICATIONS WERE REPORTED. | INFORMATION WAS RECEIVED FROM A CONSUMER REGARDING A PATIENT WHO WAS IMPLANTED WITH AN IMPLANTABLE NEUROSTIMULATOR (INS) FOR URINARY DYSFUNCTION/SACRAL NERVE STIMULATION AND GASTROINTESTINAL/PELVIC FLOOR. IT WAS REPORTED THAT THE PATIENT GOT THE POWER ON RESET (POR) MESSAGE ON THE DAY OF THE CALL. WHEN ASKED IF THE PATIENT HAD ANY RECENT MEDICAL PROCEDURES OR EMI, THE CALLER INDICATED THAT THE PATIENT WAS HOME BOUND IN HOSPICE. THE CALLER WAS REDIRECTED TO FOLLOW UP WITH THE PATIENT¿S HEALTHCARE PROVIDER TO GET THE POR CLEARED. LATER THAT DAY, THE CALLER CALLED BACK REPORTING THE SAME ISSUE AND THAT THE HEALTHCARE PROVIDER HAD TOLD THEM TO CALL BACK. AN EMAIL WAS SENT TO FIELD REPRESENTATIVES. NO PATIENT SYMPTOMS WERE REPORTED. NO FURTHER COMPLICATIONS WERE REPORTED. ADDITIONAL INFORMATION WAS RECEIVED FROM A HEALTHCARE PROVIDER (HCP) INDICATING THAT THE PATIENT WAS LAYING DOWN IN BED AT THE TIME OF THE POR MESSAGE BEING SEEN. THE CALLER WAS REDIRECTED TO THE PATIENT'S MANAGING HEALTHCARE PROVIDER. NO FURTHER COMPLICATIONS WERE REPORTED. | IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. | STIMULATOR, ELECTRICAL, IMPLANTABLE, FOR INCONTINENCE | MEDTRONIC MED REL MEDTRONIC PUERTO RICO | 3058 |  | EZW | 20130414 |  | No | HEALTH PROFESSIONAL | Stimulator, Electrical, Implantable, For Incontinence
Malfunction | 20200328 | ['Battery Problem', 'Adverse Event Without Identified Device or Use Problem', 'Power Problem'] | INFORMATION RECEIVED BY MEDTRONIC INDICATED THAT THE INSULIN PUMP HAD LOW BATTERY AND FAILED BATTERY TEST. NO HARM REQUIRING MEDICAL INTERVENTION WAS REPORTED. THE INSULIN PUMP WILL BE RETURNED FOR ANALYSIS. |  | UNIT RECEIVED WITH FAILED BATTERY TEST DUE TO CORRODED BATTERY TUBE. UNABLE TO PERFORM OPERATING CURRENTS, SELF-TEST, A21 ERROR TEST AND DISPLACEMENT TEST OR VERIFY LOW BATT ALARM DUE TO FAILED BATT TEST ALARM. (B)(4). | PUMP, INFUSION, INSULIN, TO BE USED WITH INVASIVE GLUCOSE SENSOR | MEDTRONIC PUERTO RICO OPERATIONS CO. | MMT-723RNAS | A1723RNASJ | OYC |  | 20200408 | Device was returned to manufacturer | LAY USER/PATIENT | Pump, Infusion, Insulin, To Be Used With Invasive Glucose Sensor
Malfunction | 20200323 | ['Battery Problem'] | IT WAS REPORTED THAT ON START-UP, THE PUMP ALARMED "BATTERY COMMUNICATION TIME-OUT." THE PRODUCT PROBLEM WAS OBSERVED DURING TESTING, AND WAS NOTED AS AN OUT-OF-BOX FAILURE. |  | EVALUATION RESULTS: ONE MEDFUSION PUMP WAS RETURNED FOR INVESTIGATION IN GOOD CONDITION. THERE WERE NO SIGNS OF PHYSICAL DAMAGE ON THE DEVICE AND THE TAMPER SEAL WAS INTACT. MULTIPLE ALARMS WERE OBSERVED UPON REVIEW OF THE EVENT HISTORY LOG. TO FUNCTIONALLY TEST THE DEVICE, THE INVESTIGATOR VERIFIED THE BATTERY STATUS IN BIOMED MENU. THERE WAS NO READING ON THE SCREEN. THE INVESTIGATOR THEN REPLACED A KNOWN GOOD RECHARGEABLE BATTERY PACK. THE PUMP SUBSEQUENTLY READ THE BATTERY STATUS. THE REPORTED PRODUCT PROBLEM WAS THEREFORE ATTRIBUTED TO THE RECHARGEABLE BATTERY PACK. THIS COMPONENT WAS REPLACED, AFTER WHICH, THE PUMP PASSED ALL FUNCTIONAL AND DELIVERY TESTS. THE PROBLEM SOURCE OF THE REPORTED PRODUCT PROBLEM WAS UNKNOWN. A ROOT CAUSE WAS NOT ESTABLISHED. | PUMP, INFUSION | ST PAUL | 4000 |  | FRN |  | 20200402 | No | I | Pump, Infusion
Malfunction | 20200420 | ['Ambient Noise Problem', 'Battery Problem'] | THE PATIENT REPORTED THAT THEY TRY TO CHARGE THE IMPLANTABLE NEUROSTIMULATOR (INS), THEY ARE SEEING A WARNING POWER ON RECHARGE (POR) MESSAGE SINCE THE WEEK PRIOR TO THE REPORT. THE PATIENT NOTED THAT THEY HAVE NOT HAD ANY RECENT TRAUMA OR FALLS. THEY STATED THAT THEY HAD A DIAGNOSTIC ULTRASOUND RECENTLY. COMMUNICATION WAS SENT OUT TO THE FIELD. THE PATIENT WAS REDIRECTED TO FOLLOW-UP WITH THEIR HEALTHCARE PROVIDER. NO FURTHER COMPLICATIONS OR PATIENT SYMPTOMS WERE REPORTED OR ANTICIPATED. |  | IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. | STIMULATOR, SPINAL-CORD, TOTALLY IMPLANTED FOR PAIN RELIEF | MEDTRONIC PUERTO RICO OPERATIONS CO. | 97714 |  | LGW | 20160114 |  | No | HEALTH PROFESSIONAL | Stimulator, Spinal-Cord, Totally Implanted For Pain Relief
Malfunction | 20200427 | ['Overheating of Device', 'Battery Problem', 'Charging Problem', 'Insufficient Information'] | ADDITIONAL INFORMATION WAS RECEIVED. CONSUMER REPORTED THEY ASSUMED THE CAUSE OF THE ISSUE WAS A DEVICE MALFUNCTION. NO CHANGES ON THEIR PART. PATIENT REPORTED THAT IT STARTED WITH THE DEVICE NOT BEING ABLE TO FIND THE IMPLANT WHEN THEY TRIED TO RECHARGE. THEY ISSUE WAS RESOLVED WITH THE REPLACEMENT DEVICE. NO FURTHER COMPLICATIONS REPORTED/ANTICIPATED. | SEE H10. INFORMATION WAS RECEIVED FROM THE CONSUMER REGARDING A PATIENT WITH AN IMPLANTABLE NEUROSTIMULATOR (INS). THE PATIENT REPORTED THAT WHEN CHARGING, THE SCREEN WAS SAYING CANNOT FIND DEVICE. THE PATIENT REPORTED THAT THE PADDLE GETS VERY HOT AGAINST SKIN. THE PATIENT REPORTED THAT THEY WERE ABLE TO CHARGE TO 80 YESTERDAY BUT THE DEVICE IS DEAD AND NEEDS TO BE CHARGE. | MEDTRONIC IS SUBMITTING THIS REPORT TO COMPLY WITH FDA REPORTING REGULATIONS UNDER 21 CFR PARTS 4 AND 803. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY MEDTRONIC, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. MEDTRONIC HAS MADE REASONABLE EFFORTS TO OBTAIN MORE COMPLETE INFORMATION AND HAS PROVIDED AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE AN ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY ¿DEFECTS¿ OR HAS ¿MALFUNCTIONED¿. THESE WORDS ARE INCLUDED IN THE FDA 3500A FORM AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REGULATORY REPORTING. MEDTRONIC OBJECTS TO THE USE OF THESE WORDS AND OTHERS LIKE THEM BECAUSE OF THE LACK OF DEFINITION AND THE CONNOTATIONS IMPLIED BY THESE TERMS. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. A GOOD FAITH EFFORT WILL BE MADE TO OBTAIN THE APPLICABLE INFORMATION RELEVANT TO THE REPORT. IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. H6: UPON FURTHER REVIEW, IT WAS DETERMINED THAT THE DEVICE CODE C63030 IS NOT APPLICABLE TO THE REPORTABLE RECHARGER TELEMETRY MODULE (RTM) IN THIS EVENT. H3: ANALYSIS OF THE RETURNED RECHARGER TELEMETRY MODULE (RTM) FOUND THAT THERE WAS A RECHARGE ANTENNA FAILURE, THE 'NO DEVICE FOUND' MESSAGE WAS SEEN AND THUS THE DEVICE WAS SCRAPPED. MEDTRONIC IS SUBMITTING THIS REPORT TO COMPLY WITH FDA REPORTING REGULATIONS UNDER 21 CFR PARTS 4 AND 803. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY MEDTRONIC, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. MEDTRONIC HAS MADE REASONABLE EFFORTS TO OBTAIN MORE COMPLETE INFORMATION AND HAS PROVIDED AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE AN ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY ¿DEFECTS¿ OR HAS ¿MALFUNCTIONED¿. THESE WORDS ARE INCLUDED IN THE FDA 3500A FORM AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REGULATORY REPORTING. MEDTRONIC OBJECTS TO THE USE OF THESE WORDS AND OTHERS LIKE THEM BECAUSE OF THE LACK OF DEFINITION AND THE CONNOTATIONS IMPLIED BY THESE TERMS. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. A GOOD FAITH EFFORT WILL BE MADE TO OBTAIN THE APPLICABLE INFORMATION RELEVANT TO THE REPORT. IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. H6: CONCLUSION CODE UPDATED. MEDTRONIC IS SUBMITTING THIS REPORT TO COMPLY WITH FDA REPORTING REGULATIONS UNDER 21 CFR PARTS 4 AND 803. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY MEDTRONIC, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. MEDTRONIC HAS MADE REASONABLE EFFORTS TO OBTAIN MORE COMPLETE INFORMATION AND HAS PROVIDED AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE AN ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY ¿DEFECTS¿ OR HAS ¿MALFUNCTIONED¿. THESE WORDS ARE INCLUDED IN THE FDA 3500A FORM AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REGULATORY REPORTING. MEDTRONIC OBJECTS TO THE USE OF THESE WORDS AND OTHERS LIKE THEM BECAUSE OF THE LACK OF DEFINITION AND THE CONNOTATIONS IMPLIED BY THESE TERMS. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT. CONCOMITANT MEDICAL PRODUCTS: PRODUCT ID: 97755, SERIAL#: (B)(4), PRODUCT TYPE: RECHARGER. OTHER RELEVANT DEVICE(S) ARE: PRODUCT ID: 97755, SERIAL/LOT #: (B)(4), UDI#: (B)(4). IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. | STIMULATOR, SPINAL-CORD, TOTALLY IMPLANTED FOR PAIN RELIEF | MEDTRONIC PUERTO RICO OPERATIONS CO. | 97715 |  | LGW | 20190528 |  | No | HEALTH PROFESSIONAL | Stimulator, Spinal-Cord, Totally Implanted For Pain Relief
Malfunction |  | ['Battery Problem'] | INFORMATION WAS RECEIVED FROM A CONSUMER REGARDING A PATIENT WHO WAS IMPLANTED WITH AN IMPLANTABLE NEUROSTIMULATOR (INS). THE CONSUMER REPORTED THAT ONE INS WAS DISCHARGING VERY QUICKLY. THE CONSUMER REPORTED THAT THIS HAD BEEN HAPPENING FOR A COUPLE OF MONTHS. THE PATIENT HAD NOT MADE ANY CHANGES TO HER STINGS AND THE PATIENT DIDN¿T KNOW HOW TO DO IT. THE PATIENT GETS GOOD COUPLING WHEN CHARGING. NO FURTHER COMPLICATIONS WERE REPORTED. | ADDITIONAL INFORMATION WAS RECEIVED FROM THE PATIENT. THE PATIENT REPORTED THAT NOTHING UNUSUAL LED TO THE IMPLANT DISCHARGING VERY QUICKLY. THE PATIENT REPORTED THAT BOTH WERE CHARGED FULL, THE LEFT SIDE WAS EMPTY AND THE RIGHT SIDE HAD ¾ CHARGE. THE ISSUE WAS NOT RESOLVED. NO FURTHER COMPLICATIONS WERE REPORTED. | IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. | STIMULATOR, SPINAL-CORD, TOTALLY IMPLANTED FOR PAIN RELIEF | MEDTRONIC PUERTO RICO OPERATIONS CO. | 37714 |  | LGW | 20161028 |  | No | HEALTH PROFESSIONAL | Stimulator, Spinal-Cord, Totally Implanted For Pain Relief
Injury | 20191029 | ['Battery Problem'] | INFORMATION WAS RECEIVED FROM A HEALTHCARE PROFESSIONAL (HCP) REGARDING A PATIENT IN A CLINICAL STUDY (SDY) WHO WAS IMPLANTED WITH A NEUROSTIMULATOR (INS) FOR SPINAL PAIN. IT WAS REPORTED THAT THE PATIENT¿S DEVICE INDICATED THE DEVICE HAD ¿REACHED ESTIMATED REPLACEMENT INTERVAL¿. THIS WAS NOTED TO BE EARLY BATTERY DEPLETION. THE DEVICE WAS REPLACED ON (B)(6) 2019 AND THE EVENT WAS RESOLVED WITHOUT SEQUELAE. THE CAUSE OF THE EVENT WAS NOT PROVIDED. NO FURTHER COMPLICATIONS WERE REPORTED OR ANTICIPATED. |  | PRODUCT ANALYSIS #252747108:ANALYSIS INFORMATION -- 2020(B)(6) 08:00:41 CST PLI# 10 PRODUCT ID# 97702 BELOW IS UNEDITED, SYSTEM GENERATED TEXT BASED ON THE ANALYSIS FINDING CODE(S). THE RETURNED DEVICE WAS SUBJECTED TO A SERIES OF STANDARD TESTS THAT INCLUDE BUT IS NOT LIMITED TO VISUAL INSPECTION, OUTPUT AND TELEMETRY TESTING, AND FUNCTIONAL TESTING. ANALYSIS DETERMINED THAT THE IMPLANTABLE NEUROSTIMULATOR (INS) OUTPUT AND TELEMETRY WERE ACCEPTABLE; HOWEVER, THE BATTERY WAS NEAR NORMAL BATTERY DEPLETION. H3: ANALYSIS OF THE INS ((B)(6) FOUND THAT THE BATTERY REACHED NORMAL END OF LIFE. TELEMETRY AND OUTPUT WERE OKAY. MEDTRONIC IS SUBMITTING THIS REPORT TO COMPLY WITH FDA REPORTING REGULATIONS UNDER 21 CFR PARTS 4 AND 803. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY MEDTRONIC, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. MEDTRONIC HAS MADE REASONABLE EFFORTS TO OBTAIN MORE COMPLETE INFORMATION AND HAS PROVIDED AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE AN ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY ¿DEFECTS¿ OR HAS ¿MALFUNCTIONED¿. THESE WORDS ARE INCLUDED IN THE FDA 3500A FORM AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REGULATORY REPORTING. MEDTRONIC OBJECTS TO THE USE OF THESE WORDS AND OTHERS LIKE THEM BECAUSE OF THE LACK OF DEFINITION AND THE CONNOTATIONS IMPLIED BY THESE TERMS. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. A GOOD FAITH EFFORT WILL BE MADE TO OBTAIN THE APPLICABLE INFORMATION RELEVANT TO THE REPORT. IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. | STIMULATOR, SPINAL-CORD, TOTALLY IMPLANTED FOR PAIN RELIEF | MEDTRONIC PUERTO RICO OPERATIONS CO. | 97702 |  | LGW | 20180828 | 20200417 | Device was returned to manufacturer | HEALTH PROFESSIONAL | Stimulator, Spinal-Cord, Totally Implanted For Pain Relief
Malfunction |  | ['Battery Problem'] | IT WAS REPORTED TO PHILIPS THAT THE DEVICE HAD A FAILED BATTERY. THERE WAS NO PATIENT INVOLVEMENT. |  |  | ALS DEFIBRILLATOR MONITOR | PHILIPS MEDICAL SYSTEMS | 861290 |  | MKJ |  |  | Yes | HEALTH PROFESSIONAL | Automated External Defibrillators (Non-Wearable)
Malfunction | 20200428 | ['Battery Problem'] | ADDITIONAL INFORMATION RECEIVED FROM THE CONSUMER REPORTED THERE WAS NO ISSUE RELATED TO THE IMPLANTABLE NEUROSTIMULATOR (INS) DEPLETING A LITTLE QUICKER. THE CONSUMER STATED THEY ¿MEASURED DEPLETION TWO WEEKS CONSECUTIVELY AND VERIFIED IT AT 25% AND WAS SAME AS BEFORE.¿ THE CONSUMER STATED THE ISSUE OF THE INS DEPLETING A LITTLE QUICKER WAS ¿PROBABLY A FALSE ALARM.¿ NO PATIENT SYMPTOMS OR FURTHER COMPLICATIONS WERE ANTICIPATED. | ADDITIONAL INFORMATION RECEIVED FROM THE CONSUMER REPORTED THE CAUSE OF THE INS DEPLETING QUICKER WAS UNKNOWN. THE CONSUMER CHECKED TWO WEEKS CONSECUTIVELY AND VERIFIED THE BATTERY WAS AT 75% AFTER 7 DAYS OF FULL CHARGE. THE CONSUMER STATED THE DEPLETION AFTER 1 WEEK APPEARED TO BE NORMAL AND THEY WERE GOING TO KEEP MONITORING. NO PATIENT SYMPTOMS OR FURTHER COMPLICATIONS WERE ANTICIPATED. IT WAS REPORTED THE IMPLANTABLE NEUROSTIMULATOR (INS) CHARGE SEEMS TO BE DEPLETING A LITTLE QUICKER BUT THE PATIENT IS STILL OKAY TO CHARGING ONCE A WEEK AS THEY ALWAYS HAVE. THE PATIENT HAS NOT HAD ANY SETTING CHANGES. NO PATIENT SYMPTOMS WERE REPORTED. | MEDTRONIC IS SUBMITTING THIS REPORT TO COMPLY WITH FDA REPORTING REGULATIONS UNDER 21 CFR PARTS 4 AND 803. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY MEDTRONIC, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. MEDTRONIC HAS MADE REASONABLE EFFORTS TO OBTAIN MORE COMPLETE INFORMATION AND HAS PROVIDED AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE AN ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY ¿DEFECTS¿ OR HAS ¿MALFUNCTIONED¿. THESE WORDS ARE INCLUDED IN THE FDA 3500A FORM AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REGULATORY REPORTING. MEDTRONIC OBJECTS TO THE USE OF THESE WORDS AND OTHERS LIKE THEM BECAUSE OF THE LACK OF DEFINITION AND THE CONNOTATIONS IMPLIED BY THESE TERMS. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. A GOOD FAITH EFFORT WILL BE MADE TO OBTAIN THE APPLICABLE INFORMATION RELEVANT TO THE REPORT. IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. | IMPLANTED SUBCORTICAL ELECTRICAL STIMULATOR (MOTOR DISORDERS) | MEDTRONIC PUERTO RICO OPERATIONS CO. | 37612 |  | MRU | 20160528 |  | No | HEALTH PROFESSIONAL | Implanted Subcortical Electrical Stimulator (Motor Disorders)
Injury | 20200228 | ['Delayed Charge Time', 'Battery Problem', 'Charging Problem', 'Communication or Transmission Problem', 'Insufficient Information'] | ADDITIONAL INFORMATION WAS RECEIVED FROM THE PATIENT. IT WAS REPORTED THAT THERE WERE NO CHANGES/ FALLS. PATIENT REPORTED THAT THEIR RECHARGER FAILED TO CHARGE ON MULTIPLE OCCASIONS. PATIENT STATED IT NEEDED IMPROVEMENT SO THAT IT WOULD NOT TURN OFF WHEN REMOVING IT. DUE TO COVID-19 PATIENT HAS OPTED TO WAIT FOR PROCEDURE TO BE DONE IN THE HOSPITAL. PATIENT STATED THEY WANTED INS REMOVED DUE TO TAKING TOO LONG TO CHARGE AND BATTERY DEPLETION. | INFORMATION WAS RECEIVED FROM A CONSUMER REGARDING A PATIENT WHO WAS IMPLANTED WITH AN IMPLANTABLE NEUROSTIMULATOR (INS). IT WAS REPORTED THAT SINCE ABOUT (B)(6) 2020, THE PATIENT NOTICED THAT THE INS WAS DISCHARGING VERY QUICKLY. IT WAS EXPLAINED THAT THEY WOULD WAKE UP IN PAIN AND WOULD CHECK THE CONTROLLER AND FOUND THE INS HAD DEPLETED OVERNIGHT. THEY WERE NOW HAVING TO CHARGE THE INS 2 TIMES PER DAY, WHICH WAS MORE FREQUENT THAN BEFORE. THE PATIENT CONFIRMED THAT THEY HADN'T INCREASED PROGRAMS OR DONE ANYTHING THAT WOULD CAUSE THIS EVENT TO OCCUR. THEY WERE REDIRECTED TO SEE THEIR PHYSICIAN TO HAVE THE DEVICE CHECKED. THE PATIENT COMMENTED THAT AFTER COVID-19 RESTRICTIONS ARE LIFTED, THEY WERE PLANNING ON GETTING THEIR INS REPLACED. IT WAS NOT CLEAR IF THIS MEANT THE PHYSICIAN WAS ALREADY PLANNING ON REPLACING THE INS, SO FOLLOW UP WILL BE DONE TO CONFIRM IF SURGICAL INTERVENTION HAS BEEN PLANNED OR SCHEDULED. NO FURTHER COMPLICATIONS WERE REPORTED OR ANTICIPATED. ADDITIONAL INFORMATION WAS RECEIVED. IT WAS REPORTED AT 10:30PM ON (B)(6) 2020, THE PATIENTS INS WAS FULLY CHARGED AND ON THE MORNING OF (B)(6) 2020.THEY TRIED TO CONNECT THEIR PROGRAMMER TO THEIR INS AND RECEIVED POOR COMMUNICATION. THE PATIENT CONNECTED THEIR RECHARGER AND THEIR INS HAD COMPLETELY DEPLETED OVERNIGHT. THE PATIENT DID NOT CHANGE THEIR PROGRAMMING AND THERE WERE NO FALLS/TRAUMA. FOR 40 MINUTES THEY HAD BEEN ATTEMPTING TO CHARGE THEIR INS WITH THEIR STIMULATION TURNED OFF AND FULL COUPLING BOXES BUT THEIR INS WAS NOT PROGRESSING AT ALL. NO SYMPTOMS WERE REPORTED. ADDITIONAL INFORMATION WAS RECEIVED FROM THE PATIENT. THE PATIENT REPORTED THAT THERE WERE NO CHANGES THAT LED TO THE INS DISCHARGING MORE QUICKLY. THE PATIENT REPORTED THAT HER SURGERY HAD BEEN POSTPONED DUE TO COVID-19 AND UNTIL SHE FELT SAFE GOING TO A HOSPITAL/OR. THE PATIENT REPORTED THAT THE BATTERY DEPLETES TOO QUICKLY AND AT TIMES DIDN¿T ACCEPT ANY CHARGE. NO FURTHER COMPLICATIONS WERE REPORTED. ADDITIONAL INFORMATION WAS RECEIVED FROM THE PATIENT. IT WAS REPORTED THE BATTERY DEPLETED QUICKLY. THE DEVICE WAS TURNED OFF FOR 2 DAYS AND THE PROGRAM WAS CHANGED. THE ISSUE WAS RESOLVED. ADDITIONAL INFORMATION WAS RECEIVED FROM THE PATIENT. IT WAS REPORTED THAT THE PATIENT WAS CHARGING TWICE A DAY AND THE UNIT WAS NOT HOLDING A CHARGE. PATIENT STATED MAYBE IT WAS MONTHS AGO. PATIENT STATED A NEW BATTERY CHANGE AND SINCE THEN BATTERY WAS TAKING TOO LONG TO CHARGE AND BATTERY WAS NOT LASTING. SOMETIMES IT TOOK 5 HOURS AND SOMETIMES 10 HOURS. PATIENT STATED THE TEMPERATURE WHEN CHARGING, SIGNAL CAME ON HOWEVER STATED THEY HAD A 104 DEGREE FEVER WHICH WAS 3 WEEKS AGO AND THINKS THAT WAS WHY IT CAME UP. INFORMATION WAS RECEIVED FROM A PATIENT (PT) REGARDING AN IMPLANTABLE NEUROSTIMULATOR (INS). THE REASON FOR CALL WAS PT REPORTED THAT SINCE IMPLANT THEIR RESTORE DEVICE WAS SO SLOW AND TOOK SO LONG TO RECHARGE. THE DEVICE WAS REPLACED IN 2022. | MEDTRONIC IS SUBMITTING THIS REPORT TO COMPLY WITH FDA REPORTING REGULATIONS UNDER 21 CFR PARTS 4 AND 803. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY MEDTRONIC, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. MEDTRONIC HAS MADE REASONABLE EFFORTS TO OBTAIN MORE COMPLETE INFORMATION AND HAS PROVIDED AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE AN ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY ¿DEFECTS¿ OR HAS ¿MALFUNCTIONED¿. THESE WORDS ARE INCLUDED IN THE FDA 3500A FORM AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REGULATORY REPORTING. MEDTRONIC OBJECTS TO THE USE OF THESE WORDS AND OTHERS LIKE THEM BECAUSE OF THE LACK OF DEFINITION AND THE CONNOTATIONS IMPLIED BY THESE TERMS. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. A GOOD FAITH EFFORT WILL BE MADE TO OBTAIN THE APPLICABLE INFORMATION RELEVANT TO THE REPORT. IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. MEDTRONIC IS SUBMITTING THIS REPORT TO COMPLY WITH FDA REPORTING REGULATIONS UNDER 21 CFR PARTS 4 AND 803. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY MEDTRONIC, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. MEDTRONIC HAS MADE REASONABLE EFFORTS TO OBTAIN MORE COMPLETE INFORMATION AND HAS PROVIDED AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE AN ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY ¿DEFECTS¿ OR HAS ¿MALFUNCTIONED¿. THESE WORDS ARE INCLUDED IN THE FDA 3500A FORM AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REGULATORY REPORTING. MEDTRONIC OBJECTS TO THE USE OF THESE WORDS AND OTHERS LIKE THEM BECAUSE OF THE LACK OF DEFINITION AND THE CONNOTATIONS IMPLIED BY THESE TERMS. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. A GOOD FAITH EFFORT WILL BE MADE TO OBTAIN THE APPLICABLE INFORMATION RELEVANT TO THE REPORT. IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. MEDTRONIC IS SUBMITTING THIS REPORT TO COMPLY WITH FDA REPORTING REGULATIONS UNDER 21 CFR PARTS 4 AND 803. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY MEDTRONIC, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. MEDTRONIC HAS MADE REASONABLE EFFORTS TO OBTAIN MORE COMPLETE INFORMATION AND HAS PROVIDED AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE AN ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY ¿DEFECTS¿ OR HAS ¿MALFUNCTIONED¿. THESE WORDS ARE INCLUDED IN THE FDA 3500A FORM AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REGULATORY REPORTING. MEDTRONIC OBJECTS TO THE USE OF THESE WORDS AND OTHERS LIKE THEM BECAUSE OF THE LACK OF DEFINITION AND THE CONNOTATIONS IMPLIED BY THESE TERMS. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. A GOOD FAITH EFFORT WILL BE MADE TO OBTAIN THE APPLICABLE INFORMATION RELEVANT TO THE REPORT. IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. MEDTRONIC IS SUBMITTING THIS REPORT TO COMPLY WITH FDA REPORTING REGULATIONS UNDER 21 CFR PARTS 4 AND 803. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY MEDTRONIC, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. MEDTRONIC HAS MADE REASONABLE EFFORTS TO OBTAIN MORE COMPLETE INFORMATION AND HAS PROVIDED AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE AN ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY ¿DEFECTS¿ OR HAS ¿MALFUNCTIONED¿. THESE WORDS ARE INCLUDED IN THE FDA 3500A FORM AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REGULATORY REPORTING. MEDTRONIC OBJECTS TO THE USE OF THESE WORDS AND OTHERS LIKE THEM BECAUSE OF THE LACK OF DEFINITION AND THE CONNOTATIONS IMPLIED BY THESE TERMS. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. A GOOD FAITH EFFORT WILL BE MADE TO OBTAIN THE APPLICABLE INFORMATION RELEVANT TO THE REPORT. IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. MEDTRONIC IS SUBMITTING THIS REPORT TO COMPLY WITH FDA REPORTING REGULATIONS UNDER 21 CFR PARTS 4 AND 803. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY MEDTRONIC, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. MEDTRONIC HAS MADE REASONABLE EFFORTS TO OBTAIN MORE COMPLETE INFORMATION AND HAS PROVIDED AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE AN ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY ¿DEFECTS¿ OR HAS ¿MALFUNCTIONED¿. THESE WORDS ARE INCLUDED IN THE FDA 3500A FORM AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REGULATORY REPORTING. MEDTRONIC OBJECTS TO THE USE OF THESE WORDS AND OTHERS LIKE THEM BECAUSE OF THE LACK OF DEFINITION AND THE CONNOTATIONS IMPLIED BY THESE TERMS. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. A GOOD FAITH EFFORT WILL BE MADE TO OBTAIN THE APPLICABLE INFORMATION RELEVANT TO THE REPORT. IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. MEDTRONIC IS SUBMITTING THIS REPORT TO COMPLY WITH FDA REPORTING REGULATIONS UNDER 21 CFR PARTS 4 AND 803. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY MEDTRONIC, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. MEDTRONIC HAS MADE REASONABLE EFFORTS TO OBTAIN MORE COMPLETE INFORMATION AND HAS PROVIDED AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE AN ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY ¿DEFECTS¿ OR HAS ¿MALFUNCTIONED¿. THESE WORDS ARE INCLUDED IN THE FDA 3500A FORM AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REGULATORY REPORTING. MEDTRONIC OBJECTS TO THE USE OF THESE WORDS AND OTHERS LIKE THEM BECAUSE OF THE LACK OF DEFINITION AND THE CONNOTATIONS IMPLIED BY THESE TERMS. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. A GOOD FAITH EFFORT WILL BE MADE TO OBTAIN THE APPLICABLE INFORMATION RELEVANT TO THE REPORT. IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. | STIMULATOR, SPINAL-CORD, TOTALLY IMPLANTED FOR | MEDTRONIC PUERTO RICO OPERATIONS CO. | 97714 |  | LGW | 20170814 |  | No | HEALTH PROFESSIONAL | Stimulator, Spinal-Cord, Totally Implanted For Pain Relief
Malfunction | 20200427 | ['Display or Visual Feedback Problem', 'Battery Problem'] | INFORMATION RECEIVED BY MEDTRONIC INDICATED THAT THE INSULIN PUMP HAD CLOUDY SCREEN. CUSTOMER STATED THE INSULIN PUMP HAD DIMINISHED BATTERY LIFE. NO HARM REQUIRING MEDICAL INTERVENTION WAS REPORTED. THE INSULIN PUMP WILL NOT BE RETURNED FOR ANALYSIS. |  | DEVICE MONITORED FOR SEVERAL DAYS. DEVICE RECEIVED WITH NORMAL OPERATING CURRENT. DEVICE PASSED DISPLACEMENT TEST AND SELF TEST. DEVICE PASSED OFF NO POWER TEST. NO UNEXPECTED LOW BATTERY ALARM ANOMALIES NOTED. NO UNEXPECTED MISSING SEGMENT OR PARTIAL DISPLAY ANOMALIES NOTED. (B)(4). (B)(4). CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. | ARTIFICIAL PANCREAS DEVICE SYSTEM, THRESHOLD SUSPEND | MEDTRONIC PUERTO RICO OPERATIONS CO. | MMT-751NAP | A4751NAPJ | OZO |  | 20200630 | Device was returned to manufacturer | LAY USER/PATIENT | Automated Insulin Dosing , Threshold Suspend
Malfunction | 20200423 | ['Battery Problem'] | IT WAS REPORTED THAT APPROX. 62 MONTHS AFTER THE IMPLANTATION, DURING THE INTERROGATION A BATTERY ERROR APPEARED. |  | THE DEVICE WAS NOT RETURNED FOR ANALYSIS. THE ANALYSIS IS THEREFORE ONLY BASED ON THE RETURNED DEVICE DATA AS WELL AS THE INSPECTION OF THE QUALITY DOCUMENTS ASSOCIATED WITH THE MANUFACTURE OF THIS PARTICULAR DEVICE. THE MANUFACTURING PROCESS FOR THIS DEVICE WAS RE-INVESTIGATED. ALL PRODUCTION STEPS HAD BEEN PERFORMED ACCORDINGLY. THERE WAS NO SIGN OF ANY INCONSISTENCY DURING THE MANUFACTURING. IN A NEXT STEP, THE RETURNED DEVICE DATA WAS INSPECTED. DURING THE INSPECTION THE CLINICAL OBSERVATION COULD BE CONFIRMED. FURTHER ANALYSIS REVEALED AN UNEXPECTED BATTERY BEHAVIOR, WHICH LED TO THE CLINICAL OBSERVATION. BASED ON THE DATA AVAILABLE FOR ANALYSIS THE ROOT CAUSE FOR THAT BEHAVIOR WAS NOT DETERMINABLE AND CAN ONLY BE CLARIFIED BY AN ANALYSIS OF THE DEVICE ITSELF. HOWEVER, IT CANNOT BE EXCLUDED THAT THIS OCCURRENCE RESULTED FROM A DEFECTIVE COMPONENT, WHICH MAY COMPROMISE THE ABILITY OF THE DEVICE TO DELIVER THERAPY. BIOTRONIK HAS INFORMED THE HEALTH CARE PROFESSIONAL ABOUT THIS ANALYSIS RESULT, WHO DECIDES, BASED ON THE PATIENTS INDIVIDUAL CIRCUMSTANCES AND MEDICAL JUDGMENT, IF THE DEVICE WILL BE EXCHANGED. SHOULD FURTHER RELEVANT INFORMATION OR THE DEVICE ITSELF BECOME AVAILABLE, THIS INVESTIGATION WILL BE UPDATED AND YOU WILL BE INFORMED ACCORDINGLY. | PACEMAKER | BIOTRONIK SE & CO. KG | 394973 |  | NKE | 20160131 |  | No | HEALTH PROFESSIONAL | Pulse Generator, Pacemaker, Implantable, With Cardiac Resynchronization (Crt-P)
Malfunction |  | ['Battery Problem', 'Charging Problem', 'Insufficient Information'] | INFORMATION WAS RECEIVED FROM A CONSUMER CONCERNING PATIENT WITH AN IMPLANTABLE NEUROSTIMULATOR (INS) FOR FAILED BACK SURGERY SYNDROME. IT WAS REPORTED THE PATIENT¿S PREVIOUS INS "WAS FOREVER" BUT THE INS "STARTED NOT CHARGING RIGHT". THE PATIENT REPORTED THE INS DEPLETED AND PATIENT WAITED 11 AND A HALF MONTHS, WHEN HCP FINALLY REPLACED THE PATIENT¿S INS. THE PATIENT REPORTED THAT THEY WERE SET BACK AFTER ENDURING OUT OF CONTROL PAIN. INFORMATION PREVIOUSLY RECEIVED INDICATED THE INS WAS AT END OF SERVICE (EOS). THE DEVICE WAS GREATER THAN 9 YEARS OLD AT TIME OF REPLACEMENT. ADDITIONAL INFORMATION WAS REPORTED THAT A DAY AFTER REPLACEMENT THE PATIENT WAS REPORTING A BURNING IN HIS CHEST. WHEN HIS INS WAS REPLACED ONLY ONE LEAD WAS WORKING. THE PATIENT GETS A BURNING SENSATION IN HIS CHEST SHORTLY AFTER TURNING ON HIS THERAPY. THE BURNING GOES AWAY WHEN THE THERAPY IS OFF. THE PATIENT STATED HE SAW HIS CARDIOLOGIST TWO MONTHS AGO AND EVERYTHING WAS CLEAR FOR HIM AND THE PACEMAKER. THE PATIENT'S REPLACEMENT WAS DUE TO NORMAL BATTERY DEPLETION AND HE USED TO BE PROGRAMMED ON A GUARDED CATHODE. THE PATIENT CURRENTLY HAS A DOUBLE BIPOLE. THE PATIENT NEVER HAD ISSUES WITH THE PREVIOUS PROGRAM AND THE CALLER IS NOT SURE WHY IT WAS CHANGED. THE CALLER REPROGRAMMED TO THE PREVIOUS SETTINGS AND TURNED THE THERAPY ON AND THE PATIENT REPORTED THAT HE NO LONGER FEELS A BURNING SENSATION. THE IMPEDANCES LOOKS EXACTLY THE SAME AS WHEN THE PATIENT WAS REPLACED IN (B)(6). ADDITIONAL INFORMATION WAS REPORTED THAT ON SATURDAY ((B)(6) 2020) WAS WHEN IT WAS REPORTED TO THE REP THAT THE PATIENT'S LEAD WAS NOT WORKING AT THE TIME OF THEIR REPLACEMENT DUE TO NORMAL BATTERY DEPLETION. TO THE REP'S UNDERSTANDING THERE WAS ONLY ONE FUNCTIONAL LEAD AT THE TIME OF THE REPLACEMENT. IT'S UNKNOWN WHAT ACTIONS WERE TAKEN TO RESOLVE THE NONFUNCTIONAL LEAD. THE NONFUNCTIONAL LEAD HAS NOT BEEN RESOLVED. THE CAUSE OF THE CHEST BURNING SEEMED TO HAVE BEEN RELATED TO THE PROGRAM SETTING. THERE WAS A RECORD OF HIS PREVIOUS PROGRAM (THE ONE HE HAD WITH HIS ORIGINAL INS AND HAD IT FOR YEARS WITH ADEQUATE PAIN RELIEF AND NO CHEST SYMPTOMS). WHEN THE REP PROGRAMMED THE PATIENT WITH THIS, AND TURNED ON THE STIMULATION, HE STATED HE DIDN'T HAVE ANY OF THE CHEST SYMPTOMS HE HAD COMPLAINED ABOUT PRIOR TO THE ADJUSTMENT. NO FURTHER INFORMATION WAS REPORTED. |  | CONCOMITANT MEDICAL PRODUCTS: PRODUCT ID: 3777-60, SERIAL#: (B)(4), IMPLANTED: (B)(6) 2009, PRODUCT TYPE: LEAD. PRODUCT ID: 3777-60, SERIAL#: (B)(4), IMPLANTED: (B)(6) 2009, PRODUCT TYPE: LEAD. OTHER RELEVANT DEVICE(S) ARE: PRODUCT ID: 3777-60, SERIAL/LOT #: (B)(4), UBD: 02-SEP-2013, UDI#: (B)(4) ; PRODUCT ID: 3777-60, SERIAL/LOT #: (B)(4), UBD: 02-SEP-2013, UDI#: (B)(4). IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. | STIMULATOR, SPINAL-CORD, TOTALLY IMPLANTED FOR PAIN RELIEF | MEDTRONIC MED REL MEDTRONIC PUERTO RICO | 37711 |  | LGW | 20100914 |  | No | HEALTH PROFESSIONAL | Stimulator, Spinal-Cord, Totally Implanted For Pain Relief
Malfunction | 20200405 | ['Battery Problem'] | IT WAS REPORTED THAT THE PUMP BATTERY WAS DEPLETING QUICKLY. BLOOD GLUCOSE WAS NOT IMPACTED. THE CUSTOMER CONTINUED TO USE THE PUMP FOR INSULIN THERAPY. |  | NO PRODUCT WAS RETURNED FOR EVALUATION. SHOULD NEW RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. | ALTERNATE CONTROLLER ENABLED INFUSION PUMP | TANDEM DIABETES CARE | 1002717 |  | QFG |  |  | No | LAY USER/PATIENT | Alternate Controller Enabled Insulin Infusion Pump
Malfunction | 20200407 | ['Battery Problem'] | IT WAS REPORTED THAT THE PUMP BATTERY WAS DEPLETING QUICKLY. THERE WAS NO REPORTED IMPACT TO THE CUSTOMER¿S BLOOD GLUCOSE LEVEL. THE CUSTOMER DECLINED FOLLOW UP FROM TANDEM TECHNICAL SUPPORT. |  | NO PRODUCT WAS RETURNED FOR EVALUATION. SHOULD NEW RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. | AUTOMATED INSULIN DOSING, THRESHOLD SUSPEND | TANDEM DIABETES CARE | 1000096 |  | OZO |  |  | No | LAY USER/PATIENT | Automated Insulin Dosing , Threshold Suspend
Malfunction | 20200407 | ['Battery Problem'] | IT WAS REPORTED THAT THE PUMP BATTERY INTERMITTENTLY DEPLETED QUICKLY. THE CUSTOMER'S BLOOD GLUCOSE WAS 193 MG/DL. THE CUSTOMER CONTINUED TO USE THE PUMP FOR INSULIN THERAPY. |  | THE DEVICE HAS BEEN RECEIVED FOR EVALUATION; HOWEVER, DEVICE EVALUATION IS NOT YET COMPLETE. A SUPPLEMENTAL REPORT WILL BE FILED UPON COMPLETION OF THE EVALUATION. | AUTOMATED INSULIN DOSING, THRESHOLD SUSPEND | TANDEM DIABETES CARE | 1000096 |  | OZO |  | 20200427 | Device was returned to manufacturer | LAY USER/PATIENT | Automated Insulin Dosing , Threshold Suspend
Malfunction | 20200427 | ['Battery Problem', 'Patient Device Interaction Problem'] | INFORMATION RECEIVED BY MEDTRONIC INDICATED THAT INSULIN PUMP HAD SCRATCHES ON THE SCREEN THAT MAKES DIFFICULT TO SEE READINGS AND ALERTS. NO HARM REQUIRING MEDICAL INTERVENTION WAS REPORTED. THE INSULIN PUMP WILL NOT BE RETURNED FOR ANALYSIS. |  | (B)(4). CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. | ARTIFICIAL PANCREAS DEVICE SYSTEM, THRESHOLD SUSPEND | MEDTRONIC PUERTO RICO OPERATIONS CO. | MMT-751NAH | A4751NAHJ | OZO |  |  | No | LAY USER/PATIENT | Automated Insulin Dosing , Threshold Suspend
Injury | 20191231 | ['Battery Problem', 'Charging Problem'] | IT WAS REPORTED INS WAS UNABLE TO CHARGE / DIFFICULTY WITH RECHARGING AND THERE WAS DECREASED THERAPEUTIC RELIEF. | INFORMATION WAS RECEIVED FROM A HEALTHCARE PROFESSIONAL (HCP) REGARDING A PATIENT IN A CLINICAL STUDY (SDY) WHO WAS IMPLANTED WITH A NEUROSTIMULATOR (INS) FOR FAILED BACK SURGERY SYNDROME AND SPINAL PAIN. IT WAS REPORTED THAT THE PATIENT¿S DEVICE HAD BEEN OFF, RESULTING IN INCREASED PAIN, AND THEY WERE UNABLE TO CHARGE THE DEVICE. IT WAS FURTHER NOTED THAT THERE WAS DECREASED THERAPEUTIC RELIEF. THE PATIENT PRESENTED TO THE CLINIC ON (B)(6) 2020 AND NOTED THE DEVICE HAD BEEN OFF FOR ONE MONTH. IT WAS NOTED THAT A MANUFACTURER REPRESENTATIVE (REP) NEEDED TO CHARGE THE DEVICE AND THEY WERE PLANNING TO DO SO AT THE PATIENT¿S NEXT OFFICE VISIT. ON (B)(6) 2020, A REP STATED THAT THE DEVICE WAS NOT ABLE TO HOLD A SUFFICIENT CHARGE. THE DEVICE WAS REPROGRAMMED AND A REPLACEMENT PROCEDURE WAS PLANNED. THE EVENT WAS NOTED TO BE ONGOING AND A CAUSE WAS NOT PROVIDED. NO FURTHER COMPLICATIONS WERE REPORTED OR ANTICIPATED. ADDITIONAL INFORMATION WAS RECEIVED FROM THE HEALTHCARE PROFESSIONAL (HCP). IT WAS REPORTED THAT THE INS WAS REPLACED ON (B)(6) 2020. THE EVENT WAS RESOLVED WITHOUT SEQUELAE. NO PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. ADDITIONAL INFORMATION WAS RECEIVED FROM A HEALTHCARE PROFESSIONAL (HCP) IN A CLINICAL STUDY (SDY). IT WAS REPORTED THAT IT WAS UNKNOWN WHEN THE REPLACEMENT PROCEDURE WAS SCHEDULED FOR. THE EVENT WAS STILL ONGOING AS OF (B)(6) 2020. | MEDTRONIC IS SUBMITTING THIS REPORT TO COMPLY WITH FDA REPORTING REGULATIONS UNDER 21 CFR PARTS 4 AND 803. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY MEDTRONIC, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. MEDTRONIC HAS MADE REASONABLE EFFORTS TO OBTAIN MORE COMPLETE INFORMATION AND HAS PROVIDED AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE AN ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY ¿DEFECTS¿ OR HAS ¿MALFUNCTIONED¿. THESE WORDS ARE INCLUDED IN THE FDA 3500A FORM AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REGULATORY REPORTING. MEDTRONIC OBJECTS TO THE USE OF THESE WORDS AND OTHERS LIKE THEM BECAUSE OF THE LACK OF DEFINITION AND THE CONNOTATIONS IMPLIED BY THESE TERMS. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT. IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. MEDTRONIC IS SUBMITTING THIS REPORT TO COMPLY WITH FDA REPORTING REGULATIONS UNDER 21 CFR PARTS 4 AND 803. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY MEDTRONIC, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. MEDTRONIC HAS MADE REASONABLE EFFORTS TO OBTAIN MORE COMPLETE INFORMATION AND HAS PROVIDED AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE AN ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY ¿DEFECTS¿ OR HAS ¿MALFUNCTIONED¿. THESE WORDS ARE INCLUDED IN THE FDA 3500A FORM AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REGULATORY REPORTING. MEDTRONIC OBJECTS TO THE USE OF THESE WORDS AND OTHERS LIKE THEM BECAUSE OF THE LACK OF DEFINITION AND THE CONNOTATIONS IMPLIED BY THESE TERMS. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. A GOOD FAITH EFFORT WILL BE MADE TO OBTAIN THE APPLICABLE INFORMATION RELEVANT TO THE REPORT. IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. | STIMULATOR, SPINAL-CORD, TOTALLY IMPLANTED FOR PAIN RELIEF | MEDTRONIC PUERTO RICO OPERATIONS CO. | 97714 |  | LGW | 20160514 |  | No | HEALTH PROFESSIONAL | Stimulator, Spinal-Cord, Totally Implanted For Pain Relief
Malfunction | 20200407 | ['Battery Problem'] | IT WAS REPORTED THAT THE PUMP BATTERY WAS DEPLETING QUICKLY. THERE WAS NO REPORTED IMPACT TO THE CUSTOMER¿S BLOOD GLUCOSE LEVEL. THE CUSTOMER DECLINED FOLLOW UP FROM TANDEM TECHNICAL SUPPORT. |  | NO PRODUCT WAS RETURNED FOR EVALUATION. SHOULD NEW RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. | INSULIN PUMP | TANDEM DIABETES CARE | 4628-003 |  | OYC |  |  | No | LAY USER/PATIENT | Pump, Infusion, Insulin, To Be Used With Invasive Glucose Sensor
Malfunction | 20200324 | ['Battery Problem'] | IT WAS REPORTED THAT THE PUMP BATTERY INTERMITTENTLY DEPLETED QUICKLY. THE CUSTOMER¿S BLOOD GLUCOSE RANGED FROM 200-230 MG/DL. MULTIPLE ATTEMPTS WERE MADE TO FOLLOW UP WITH THE CUSTOMER REGARDING THE REPORTED ISSUE; HOWEVER, NO RESPONSE FROM THE CUSTOMER WAS RECEIVED. |  | NO PRODUCT WAS RETURNED FOR EVALUATION. SHOULD NEW RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. | PUMP, INFUSION, INSULIN, TO BE USED WITH INVASIVE GLUCOSE SENSOR | TANDEM DIABETES CARE | 1000096 |  | OYC |  |  | No | LAY USER/PATIENT | Pump, Infusion, Insulin, To Be Used With Invasive Glucose Sensor
Malfunction | 20200407 | ['Break', 'Battery Problem'] | IT WAS REPORTED THAT THE PUMP BATTERY WAS DEPLETING QUICKLY. ADDITIONALLY, THE PUMP USB PORT WAS OBSERVED TO BE DAMAGED. THERE WAS NO REPORTED IMPACT TO THE CUSTOMER¿S BLOOD GLUCOSE LEVEL. THE CUSTOMER DECLINED FOLLOW UP FROM TANDEM TECHNICAL SUPPORT. |  | NO PRODUCT WAS RETURNED FOR EVALUATION. SHOULD NEW RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. | INSULIN PUMP | TANDEM DIABETES CARE | 4628-003 |  | OYC |  |  | No | LAY USER/PATIENT | Pump, Infusion, Insulin, To Be Used With Invasive Glucose Sensor
Malfunction | 20200427 | ['Circuit Failure', 'Battery Problem', 'Material Integrity Problem'] | INFORMATION RECEIVED BY MEDTRONIC INDICATED THAT INSULIN PUMP HAD MULTIPLE PUMP ERROR ALARMS. CUSTOMER WAS ABLE TO SUCCESSFULLY CLEAR THE ALARM AND ABLE TO REWIND THE INSULIN PUMP. CUSTOMER REPORTED THAT BATTERY CAP METAL CONNECTOR WAS ALSO LOOSE. CUSTOMER ALSO REPORTED THAT TROUBLESHOOT WAS PERFORMED FOR PUMP ERROR ALARM. CUSTOMER WAS PERFORMING SELF TEST AND IT DID NOT PASS. NO HARM REQUIRING MEDICAL INTERVENTION WAS REPORTED. THE INSULIN PUMP WILL NOT BE RETURNED FOR ANALYSIS. |  | (B)(4). CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. THE INSULIN PUMP INVOLVED IN THIS EVENT IS THE 640G INSULIN INFUSION PUMP, WHICH IS NOT MARKETED IN THE UNITED STATES. HOWEVER, THE DEVICE IS SIMILAR TO THE PARADIGM REAL-TIME INSULIN INFUSION PUMP, WHICH IS MARKETED IN THE UNITED STATES. UPDATED H9: 2032227-060322-002-C MEDTRONIC, INC. (MEDTRONIC) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY MEDTRONIC, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. MEDTRONIC HAS MADE REASONABLE EFFORTS TO OBTAIN MORE COMPLETE INFORMATION IN THE TIME ALLOTTED AND HAS PROVIDED AS MUCH INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE AN ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY "DEFECTS" OR HAS "MALFUNCTIONED". THESE WORDS ARE INCLUDED IN THE FDA 3500A FORM AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. MEDTRONIC OBJECTS TO THE USE OF THESE WORDS AND OTHERS LIKE IT BECAUSE OF THE LACK OF DEFINITION AND THE CONNOTATIONS IMPLIED BY THESE TERMS. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT. DEVICE PASSED THE DISPLACEMENT TEST. HOWEVER, PUMP ERROR 63 ALARM (VARIABLE INFO=3) DURING SELF TEST AND CONFIRMED IN THE DEVICE HISTORY DUE COLD SOLDER JOINT ON U1 KEYPAD ASSEMBLY. NO ERROR 25 ALARM NOTED IN THE HISTORY FILE. UNABLE TO CONFIRM BATTERY CAP DAMAGE DUE TO INSULIN PUMP RECEIVED WITHOUT THE ORIGINAL BATTERY CAP. | PUMP, INFUSION, INSULIN, TO BE USED WITH INVASIVE GLUCOSE SENSOR | MEDTRONIC PUERTO RICO OPERATIONS CO. | MMT-1711K | HG1WT7E | OYC |  | 20200706 | Device was returned to manufacturer | LAY USER/PATIENT | Pump, Infusion, Insulin, To Be Used With Invasive Glucose Sensor